## Slide 1: Coldbox design update
You can write unit / affiliation. leave blank if not applicable. To change content go to: Menu -> Insert (Mac = Display) -> Header and Footer
19-06-2019

## Slide 2: Peltier element
New Peltier element: HP-127-1.4-2.5-72
It is the optimal element among available elements.
Single 40x40 mm element simplifies construction
We order 4 elements for the prototype. Immediate shipping.
Large order for all institutes (~50) has 16 weeks lead time.
Imax | 3.7 amp(s)
Qmax | 37.4 watt(s)
Vmax | 16.3 volt(s)
DT max | 72 Th=300K
A | 40 mm
B | 40 mm
H | 4.8 mm

## Slide 3: Coldplate
The new Peltier element means that the original coldplate is no longer optimal.
New coldplate: CP10-G18 from AMS Technologies.
Benefits:
Dimensions better suited to single 40x40mm Peltier element
Two vacuum chucks can be mounted on single coldplate
Reduced cost

## Slide 4: Vacuum chuck
Vacuum chuck has been redesigned to reduce thermal mass and make it compatible with the new coldplate and Peltier element.
A slot has been added for a 4-wire PT100 element beneath the top plate.
Vacuum inlet
PT100 slot
Mount holes

## Slide 5: Assembly
Assembly requires 12 holes to be drilled in coldplate.
Not an issue, lots of clearance to tubes.
May be possible to request this modification by manufacturer on large order for all institutes.
Vacuum chuck to coldplate mounting holes
Coldplate to box frame mounting holes

## Slide 6: Coldbox
New coldplates means the coldbox can be further reduced in size.
Current design uses EN4SD362410GY.
914 mm
610 mm
254 mm
The length of this box means plenty of room for routing of services inside shielding enclosure. (30cm)

## Slide 7: Coldbox assembly
Extruded aluminium frame
Manifolds
Bottom insulation
4x coldplate-chuck assemblies
50-pos connector for sensors
Module and Peltier power terminal blocks
DIN rail

## Slide 8: Tube routing
Tubes are routed beneath coldplates.
Consists of four lines.
Each line cools one coldplate
All lines are same length -> flow is evenly distributed.
Bottom view of coldbox assembly

## Slide 9: Coldbox assembly
Two thin (1cm) Armaform insulation plates are added to insulate the vacuum chucks (-25C) from the coldplates (+10C).
Cables are routed through slots in vertical wall.

## Slide 10
-25°C volume
+10°C volume

## Slide 11: Services
To make things simpler, we’ll just route all cables through a hole in the coldbox.
All sensors are connected through a 50-wire flat ribbon cable to a 50-pin DIN-mount terminal block.
Flat ribbon cable out of box to DAQ system
To individual sensors

## Slide 12: Services
Low voltage cables (Peltier and module power) is connected via DIN-mount terminal blocks.
x32

## Slide 13
~25cm
~5cm