## Slide 1: HEPS Injection/extraction hardware development
Jinhui Chen, Zhe Duan, Lei Wang, Lihua Huo , Yuwen Wu, Hua Shi, Guanwen Wang, Xiaolei Shi, Peng Liu, Guanjian Wu, Xinzhe Zhai
Injection group，IHEP
2023-2-14
IAS Program on High Energy Physics 2023
February 12 – 16, 2023, Hongkong

## Slide 2: Outline
1
Overview of HEPS Inj./Ext. systems
2
Lambertson magnets
3
Slotted-pipe kicker system for BST
4
Strip-line kicker system for SR
5
Summary

## Slide 3
Overview of the HEPS injection and extraction systems

## Slide 4: High Energy Photon Source
HEPS is a typical 4th generation light source on green field in Huairou Beijing, China.
Officially approved in Dec. 2017, the construction was scheduled to start at the end of 2018, and completed in 2024. The whole project will be finished in mid-2025 after commissioning.
A R&D program for HEPS called HEPS-TF was started in 2016, and was completed in Sept. 2018 in order to solve some key technical issues, including some injection and extraction technologies.
So far, the Linac has been ready to commission , the booster tunnel installation has been completed, the SR tunnel installation has been started.

## Slide 5: Introduction of HEPS accelerator
The HEPS accelerator complex consists of a Linac, several transport lines, a booster ring and storage rings.
Linac: 0.5Gev/50Hz
Booster: 0.5GeV~6GeV/454.1m/1Hz
Storage ring: 6GeV/ 1360.4m, 34.8pm natural emittance
  48 hybrid-7BA cells, w/ BLGs and ABs
  interleaved high-β and low-β straight sections of 6 m
Linac
Booster
Storage ring

## Slide 6: Top-up injection scheme for HEPS
HEPS SR is a typical Low Emittance Ring (LER) with the MBA lattice. In order to achieve higher brightness, the Dynamic Aperture (DA) large enough for traditional off-axis injection is hard to reach at the same time.
HEPS SR top-up injection scheme:
  On-axis swap-out injection using the booster as an accumulator ring is the baseline scheme
  Reserve possibility for on-axis longitudinal injection with dual active RF system, which frequency is 166MHz and 500MHz.
*Zhe Duan
Color: DA area

## Slide 7: High-energy accumulation in the booster
Major challenge of swap-out injection: a full charge injector
  Full charge bunch prepared in an AR, injected into and accelerated in the booster (APS-U)
  Bunch recycling w/ a dedicated AR between BR and SR (ALS-U)
  Bunch recycling w/ booster serving as a full energy AR (HEPS) [1,2]
500MeV, up to 7nC
FODO cell, 16nm, up to 5nC
[1] Z. Duan, et al., IPAC’18, THPMF052
[2] Y. Jiao et al., J. Synch. Radiat., 25 (2018)1611

## Slide 8: HEPS Injection & Extraction systems
BTS: SR 6GeV injection system
STB: SR 6GeV extraction system
(Bunch by bunch/on-axis swap out injection)
BTS: Booster 6GeV extraction system
(Bunch by bunch/fast extraction assisted by 4 slow bumpers)
STB: Booster 6GeV injection system
(Pulsed local bumped injection by 2 kickers )
LTB: Booster 0.5GeV injection system
(Bunch by bunch/one turn/on-axis injection )
The injection and extraction system of HEPS consists of 5 sub-systems.

## Slide 9: The layout of the injection & extraction sub-systems
6 GeV extraction
6 GeV injection
Stored beam
injected beam
Septum（6mm）

## Slide 10: Hardware challenges of injection system
SR injection and extraction system
  2 straight sections of 6 meter to accommodate injection and extraction system for 6 GeV beam manipulation.
  Fast kicker system: 300mm long strip-line kicker groups and super fast pulser of 10ns bottom width (3%-3%)，±15kV peak into 50ohm.
  Thin septum: Lambertson magnet of 2mm thickness septum
BST injection and extraction system
  kicker system: slotted-pipe kicker and fast pulser of 300ns(at least 600ns) bottom width (10%-10%)，2.8kA peak into 0.7μH to inject 10 (at least 5) bunches into the BST ring in a cycling period.
  Thin septum: Lambertson magnet of 3.5mm thickness septum
  Slow bumper: half-sine wave of 1ms bottom width(only for extraction subsystem)
6ns
10ns
On-axis swap-out injection
6ns
4ns
On-axis long. injection
RF=166MHz
Kicker electrical pulse

## Slide 11
Lambertson magnets

## Slide 12: Typical requirements for HEPS LSM
Parameters2 | Unit | BST1-LSM | BST3/BST4-LSM | R48/R03-LSM
Quantity | - | 1 | 2 | 2
Energy | GeV | 0.5 | 6 | 6
Deflection angle | mrad | 200 | 79.945 | 79.95
Insertion length | m | 0.5 | 1.6 | 1.6
Magnetic field strength for injected/extracted beam | T | 0.7 | 1 | 1
Min. Septum thickness（including septum board, wall of beam pipes, installation gap) | mm | ≤6 | ≤3.5 | ≤2
Field uniformity | - | <±0.05% | <±0.05% | <±0.1%
Leakage field | - | ≤1×10-3 | ≤1×10-3 | ≤0.002T·m
Clearance of stored beam at lambertson（HV） （refer to stored beam orbit） | mm | 1010 | 2228 | 85
Clearance of inj.&ext. beam at lambertson（HV）（refer to inj.&ext. beam orbit） | mm | 2228 | 103 | 61.4
Physical aperture of stored beam vacuum chamber | mm | 2828 | 2828 | 10.5x5
In-air
Half in vacuum

## Slide 13: Lambertson Magnet Design Considerations
In order to decrease the septum thickness, 2 novel structures of magnet were proposed:
  half-in-vacuum (for 2mm septum)
  embedded thin-wall vacuum chamber (for 3.5mm septum)
Injected bunch
Stored bunch
Magnet pole
Installation margin
Yoke
Injected bunch
Magnet pole
Installation margin
Yoke
Stored bunch
Injected bunch
Magnet pole
Installation margin
Installation margin
Yoke
Stored bunch
SS
Iron or
VP(FeCoV，1J22)
Traditional LSM
SS
Half in-vacuum
in-air
Septum
thickness
Wall of stored beam pipe
+
Wall of injected beam pipe
+
Iron septum
+
installation margins
=

## Slide 14: Lambertson Magnet Design Considerations
In order to decrease the absolute values of leakage field with thinner septum, Vanadium Permendur (FeCoV： iron50% cobalt 48% vanadium 2%, domestic brand 1J22 ） is adopted for septum board.
    Higher Bs (Saturation magnetic density)
    Higher μr(Relative permeability)

## Slide 15: Lambertson Magnet Design Considerations
In order to decrease the integrate leakage field, the upstream end of the stored beam chamber is located under the side leg of the yoke to create a leakage field that is opposite in sign.

## Slide 16: Lambertson Magnet Design Considerations
In order to further reduce the integrate leakage field, the transition part and shielding plate design at the end of Lambertson magnet also plays an important role. With the shield plate, the leakage field in the opposite direction to the main field will be generated at the end of the magnet to cancel the leakage field in the body of magnet.

## Slide 17: Half in-vacuum LSM R&D for HEPS SR
Full size prototype engineering design
FeCoV
Cu
2mm
DT4

## Slide 18: Half in-vacuum LSM R&D for HEPS SR
Full size prototype:
  Started in 2021
  All the mechanical component processing was completed before June 2022.
  The biggest challenge is magnetic shielding block machining because the VP is hard and brittle. Although it can be segmented processing by EDM, but annealing deformation is hard to control.
Vacuum chamber
DT4 Yoke
1J22 Magnetic shielding block
Half cu Beam pipe
1J22 with cu coating
DT4 with Ni coating
Upper iorn yoke
Transition part

## Slide 19: Half in-vacuum LSM prototype for HEPS
  Final assembling and vacuum testing was completed in July 2022.
  Vacuum pressure: 5.0×10-8pa（vacuum chamber）, 2.2×10-7pa（Transition section）

## Slide 20: Half in-vacuum LSM prototype for HEPS
  Magnetic field measurement by Hall probe measurement system was finished successfully in Jan. 2023.

## Slide 21: Main field distribution measurement
All magnetic field performance of the prototype can meet the requirements of physics.
Chart: Horizontal integral field distribution in y=0.5mm plane
| -6 | -5 | -4 | -3 | -2 | -1 | 0 | 1 | 2 | 3 | 4 | 5 | 6
X-ERROR | -0.0004932 | -0.00047912 | -0.00050497 | -0.00047063 | -0.00037677 | -0.0001877 | 0 | 7.1494e-05 | 3.2269e-05 | 0.00014347 | 0.00021655 | 0.00018436 | 1.7925e-05
  Horizontal field distribution in range x=±6mm : 7.22E-04 (<±0.1%) @ 175A
Chart: Field in different x position
Chart: Field in different x position

## Slide 22: Main field distribution measurement
  vertical field distribution in range 1.2mm<y<4mm：3.85E-04 (<±0.1%) @ 175A
Chart: Vertical integral field distribution in x=0 plane
| 1.2 | 1.3 | 1.4 | 1.5 | 1.6 | 1.7 | 1.8 | 1.9 | 2 | 2.1 | 2.2000000000000002 | 2.4 | 2.7 | 3.2 | 3.6 | 4
Y-ERROR | 0 | -4.7456e-05 | 8.2464e-05 | -2.223e-06 | 5.3002e-05 | 8.3224e-05 | 1.0219e-05 | 3.1427e-05 | 4.0022e-05 | 4.5867e-05 | 8.7805e-05 | 0.00011159 | 0.00011925 | 0.00021297 | 0.00027063 | 0.00033761
*Y=0 plane is surface of bottom polar
Chart: Field in different y position
Chart: Field in different y position

## Slide 23: Integral leakage field measurement
Vertical integral leakage field
  w/o transition part：integral vertical leakage filed=3.25E-03(53431Gauss·mm )，not meet requirement
  With transition part：integral vertical leakage filed=4.3E-04 (-6997Gauss·mm)， meet requirement
Chart: Vertical leakage field with transition part
Chart: Vertical leakage field w/o transition part
Chart: Vertical leakage field with transition part
Central section
Central section
Outlet section
Entry section

## Slide 24: Integral leakage field measurement
Chart: Horizontal leakage field with transition part
Chart: Horizontal leakage field w/o transition part
Chart: 
Horizontal integral leakage field
  w/o transition part：integral horizontal leakage filed=8.75E-04(14402Gauss·mm )，meet requirement
  With transition part：integral horizontal leakage filed=8.19E-04 (13464Gauss·mm)， meet requirement
Central section
Outlet section
Entry section

## Slide 25: In-air LSM R&D for HEPS BST
Because FeCoV（1J22，Co50）is hard to machine, the magnetic shielding blocks must segmented processing by EDM . And that, the embedded thin wall SST vacuum chamber for stored beam is needed.
Feature：magnet is located in the air; total septum thickness=3.5mm， Length=1.6m

## Slide 26: In-air LSM R&D for HEPS BST
  The process of prototypes being assembled
  The process of prototypes hall probe measurement
(BS1LSM)
(BS3LSM)

## Slide 27: In-air LSM R&D for HEPS BST
Field measurement result shows that The gap at the joint of magnetic poles has a great influence on the leakage field and the main field distribution.
After repair，the BS1LSM and BS3LSM have pass the second time magnetic field measurement
BS1LSM
BS3LSM

## Slide 28: In-air LSM R&D for HEPS BST
BST high energy extraction LSM (BS3LSM)
BS3LSM:horizontal main field distribution is improved from ±0.0732% to ±0.0278%
BS1LSM:vertical main field distribution is improved from ±0.19% to ±0.0468%

## Slide 29: In-air LSM R&D for HEPS BST
BS1LSM has been installed into the booster tunnel.
BS3LSM is ready to been installed.
BS1LSM
BS3LSM
Injected beam from LINA
stored beam

## Slide 30
Slotted-pipe kicker system for the Booster

## Slide 31: Parameters of kicker system in BST
parameter | LE-injection | HE-type I | HE-type II | HE-type III | Unit
name | BS1K | BS2K | BS3K | BS4K | -
Quantity | 1 | 1 | 1 | 1 | -
Magnetic effective length | 0.8000 | 1.4000 | 1.4000 | 1.4000 | m
Deflect direction | Vertical | Vertical | Vertical | Vertical | -
Deflect angle | 9.10 | 1.02 | 1.75 | 1.75 | mrad
Magnetic strength | 0.020 | 0.015 | 0.025 | 0.025 | T
Clearance region(H×V) | 22×28 | 30×28 | 30×28 | 30×28 | mm
Good field region(H×V) | [-6,6]×[-8,8] | [-6,6]×[-4,3] | [-6,6]×[-5,4] | [-6,6]×[-4,3] | mm
Field uniformity in good field region | x=±4,y =±5，better than±1%; x=±6,y =±8，better than±2.5% | ±1% | ±2% | ±1% | -
Repetition rate | 50 | 50 | 50 | 50 | Hz
Amplitude repeatability | ±0.5% | ±0.5% | ±0.5% | ±0.5% | -
Pulse jitter | ≤5 | ≤5 | ≤5 | ≤5 | ns
Bottom width of half sine pulse | < 300（at least 600） | <300（at least 600） | <1000 | <1000 | ns
Exciting current | 2160 | 1630 | 2715 | 2715 | A
Inductance of kicker | 265+200 | 460+200 | 460+200 | 460+200 | nH

## Slide 32: Engineering design of the slotted-pipe kicker completed
Features：the kicker outer body is machined by EDM; One end of the busses are connected to outer body which is on ground potential level. So need a bipolar pulsed power supply. Totally 10 HV RF 50Ω cables are connected to the other end of the busses by the feedthroughs.
H×V=36×36mm
Outer body (316L)
Busses (FOC)
insulation support (ceramic)
HV feedthrough
Cooling water pipe
+
-
To vacuum pump

## Slide 33: Prototype kicker for HEPS BST
  the kicker prototype passed the vacuum assembling test and the magnetic field measurement have completed. The field distribution performance meet the requirements.

## Slide 34: Kicker magnets installation
All 4 kicker magnets for the booster injection & extraction have been installed into the booster tunnel.

## Slide 35: Solid-state fast kicker pulser R&D
Scheme: 20-stage inductive adder based on SiC-MOSFETs.
The co-axial transformer is configured as bipolar output.
The pulser is located outside tunnel and 10 50 Ω cables with length more than 30m are applied to connect with kicker.
Matching terminal resistor is 10Ω.
20-stage inductive adder

## Slide 36: Pulser prototype for HEPS
Progress：Single stage and double stages full power test has been completed（1400V into 0.5Ω）
Final version
Mass production and test

## Slide 37: Pulser prototype for HEPS
Latest Progress：18-stage inductive adder was tested（DC HV=800V, output pulse voltage is 14.2kV into 10Ω）.
Due to the existence of parasitic inductance outside the adder, a reverse voltage spike with high amplitude will be induced at the tail of the pulse when MOSFET is turned off, which must be suppressed by the reverse peak absorption circuit composed of high-speed diodes, Otherwise, MOSFET overvoltage breakdown will be caused.

## Slide 38
Strip-line kicker system for the Storage ring

## Slide 39: Parameters of kicker system for SR
Parameters | Unit | injection |  | extraction | 
Name |  | R03K1~5 |  | R48K1~5 | R48PK
Straight section length | m | 6.086 |  | 6.086 | 
Total kick angle | mrad | 1.6108 |  |  | 0.2~0.075
Kick direction | - | Vertical |  |  | Horizontal
Length of Strip-line kicker electrode | mm | 300 |  |  | 300
Gap between two electrodes | mm | 8 |  |  | 16
Quantity of Strip-line kicker | - | 5 | 5 |  | 1
Longitudinal space between strip-line kicker electrode | mm | 6 |  |  | -
Good field region | mm | x=±1.1，y= (-0.85, 2.1) |  |  | x=±0.5，y±0.4
Integral field uniformity | - | <±1% |  |  | <±1%
Odd mode impedance |  | 50±1 |  |  | 50±1
Even mode impedance |  | <65 |  |  | <65
Amplitude of electrical pulse（into 50Ω） | kV | ±15 |  |  | ±7.5
Bottom width of electrical pulse（3%-3%） | ns | <4 * |  | <10 | <10
FWHM(50%-50%) | ns | >2 * |  | >4 | >4
Electrical Pulse Amplitude Stability | - | <2% (RMS) |  |  | <2% (RMS)
Repetition rate（CW） | Hz | 50 |  |  | 50
Time jitter between electrical pulse and timing clock | ps | <200 |  |  | <200
Time jitter between channels (bipolar) | ps | <100 |  |  | <100

## Slide 40: Strip-line kicker scheme
HEPS strip-line kicker prototype 2D model
The strip-line blades are decided by:
  a, b - axes of the center ellipse
  gap=5mm – ½ distance between blades
  blade=3mm – thickness of blades
The outer body half shell consists of 2 half ellipses that defined by:
  Center half: Xc, a0, b0
  Outer half: Xc, a00, b0
New parameter:
  vane (>=<b)- ½ distance between Vanes
Following APS-U type strip-line kicker
  “D” shaped blades are used to improve field-uniformity in the good field region.
  An ellipse outer body with vanes geometry is adopted to ease common-mode impedance-matching
  Tapered end sections for matching impedance to the feed-throughs.
out body
APS-U
vane
blade
tapered end
HEPS 2D Model

## Slide 41: Prototype I: 750mm-long Strip-line Kicker
750mm
Commercial
Feedthrough
End cover
Beryllium copper braid
Outer body
blade
Silver seal
Coaxial Connector
Holder for assembling
Beryllium copper braid
to Vacuum pump

## Slide 42: Prototype kicker completed in 2018
Pusle voltage=18kV
FWHM=1.92ns
Vacuum =2E-9~2E-8Torr
Kicker HV test bench
HV pulser
Kicker HV test reult：
Kicker electrical pulse：
Rise time（10%-90%）<641ps，
Fall time（90%-10%）<1.5ns。
pluse edge slowing down due to kicker insertion：
rise time (10%-90%)≈86.6ps,
Fall time (90%-10%) ≈ 35ps。
Kicker transmission characteristics test result
Feed through
D shape blades
Outer body
Cooling water tube
End cover
Vacuum leak test
after brazing
750mm
Ion pump
kicker

## Slide 43: Limit longitudinal installation space
The HEPS prefers shorter strip-line kicker(300mm-long) to reduce difficulty of fast pulser, specially for longitudinal injection
5 sets of shorter kickers and a Lambertson magnet must be installed in 6m of straight section
The separated kicker structure is not fit for the shorter kicker for HEPS because of limited longitudinal space.
A new compact design have to be considered.
750mm
3300mm
100mm

## Slide 44: Prototype II: 300mm-long Strip-line Kicker
New kicker structure with 5 cascaded kickers in a single vacuum chamber was proposed.
The longitudinal gap between adjacent kicker blades is <6mm.
The outer body is split into 4ps(1.2m long) which made by CNC machining from FOC.
300mm-long strip-line blades without taper end are also made by CNC machining from FOC.
The coaxial bend transition parts connected blades and HV RF feedthroughs is the key point.

## Slide 45: The strip-line kicker prototype completed in 2019
Feature：5 sets of 300mm strip-line kicker in a single module
300mm long kicker:
Pulse voltage: ±20kV into 50Ω
Tr(10%-90%)=670.7ps
Tf(90%-10%)=1.4ns
FWHM=1.9ns
HV pulse testing at ± 20kV
No.1
No.2
No.3
No.4
No.5
Keysight E5071C VNA
TDR measurement
L. Wang, J.H. Chen, A 300mm long prototype strip-line kicker for the HEPS injection system, IPAC2019

## Slide 46: Progress of final kicker production for HEPS
vacuum chamber
outbody
HV RF feedthrough
Support board
eletrodes
Component made in SST
The all mechanical machining have been completed
first module has been assembled and passed the vacuum commissioning.

## Slide 47: 10ns-fast puler R&D
16 MOSFETs per stage
8 DSRDs(4 parallel, 2 series) switching 600A into 50 Ω;
0.3m long PFL(TL1) for pw=3ns
DSRDs circuit with 6-stage inductive adder driven
K005
Features：
Drift Step Recovery Diode(DSRD) [4]: Opening in sub-ns
Chips from the Chinese vender are available (i p≥300A/ Vp=10kV)
A 6-stage inductive adder designed as DSRD pumper circuit for 15kV:
  avoid isolated switch driver and HV components(capacitors, switches <1kV)
  need much fewer components compared to an inductive adder pulser

## Slide 48: The 10ns fast pulser prototype R&D completed in 2018
4ns
Vpeak=17kV,
FWHM=5ns,
Tr(10-90%)<2.6ns,
Tf(90-10%)<3.2ns,
Bottom width(3%-3%)<10ns
Time jitter<100ps(RMS)
PFL=0.3m
±18kV pulsers installed in 19 inch cabinet
J.H. Chen, Strip-line kicker and fast pulser R&D for the HEPS on-axis injection system, NIMA 920(2019)1-6
DSRD assembling
（2×4=8pcs）
Saturable inductor
（2 DN200L ferrite cores）
DSRD array
To reduce residual voltage：
DSRD array: 4 in parallel and 2 in series
Saturable inductor: 2 NiZn ferrite cores

## Slide 49: New pumper R&D based on commercial SiC MOSFET
    The commercial RF-MOSFET(IXZ631DF12N100) applied in pusler prototype being in “End of Life” status
    Alternative power semiconductor SWITCH：SiC MOSFET C2M0045170D （CREE）with higher Vds and Id, but slower switching speed.
    New prototype based on SiC MOSFET was developed. A Non-linear transmission line is introduced to sharpen the pulse. The best result is the bottom width(3%-3%) of pulse is less than 10ns, the peak large than 15kV into 50ohm.

## Slide 50: Summary
The HEPS accelerators includes 3 injection systems and 2 extraction systems, which were designed to realize on-axis swap-out top-up injection in the storage ring with small DA.
The injection and extraction hardware R&D activities for HEPS was initial from 2016, HEPS-TF project.
2 kinds of Lambertson magnet with thin septum are developed respectively for the storage ring and booster. The half-in-vacuum Lambertson magnet prototype has been completed and the in-air one was installed into the HEPS booster tunnel.
The slotted-pipe kickers were developed successfully and also installed into the into the HEPS booster tunnel. The solid-state pulser with 300ns bottom width are still under R&D.
2 kind of strip-line kicker prototype were developed for the HEPS on-axis swap-out injection successfully. The final devices have been in assembling and vacuum commissioning phase. The fast pulsers based on DSRD with a 10 ns pulse width (3%-3%) was developed successfully. And a new prototype with alternative SiC MOSFET pumping circuit is in commissioning phase.

## Slide 51
Thank you for attentions