## Slide 1: Environmental neutron measurement at the Gran Sasso laboratory in the NEWSdm experiment
T. Shiraishi - Kanagawa Univ.
On the behalf of the NEWSdm collaboration
14 Dec. 2023 CYGNUS 2023 (Sydney)

## Slide 2: For the Verification of DAMA Signal
D. Cintas et al, J. Phys.: Conf. Ser. 2156, 012065 (2021)
DAMA claims signal around 2 – 6 keVee
R. Bernabei et al., Nucl. Phys. At. Energy 19, 307 (2018)
Assuming DAMA signal as neutron background, its energy corresponding to 80 – 250 keV
 Neutron spectrum measurement including sub-MeV region is needed to verify the DAMA signal!
Many experiments agree that quenching factor of Na is ~20%
22 years annual modulation

## Slide 3: Environmental Neutron Measurement @LNGS
Chart: Neutron Energy (MeV)
| 1E-3 | 1.2504999999999999 | 0.1 | 1.75 | 1 | 2 | 2.75 | 3 | 3.75 | 4 | 5 | 5.25 | 6 | 7 | 7.5 | 8 | 9 | 10 | 11 | 12 | 12.5 | 13 | 13.75 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25
⁶Li液体シンチレーター | 0.78
プラスチックシンチレーター | 0.14 | 0.13 | 0.15 | 0.004
BF₃比例計数管 | 0.54 | 0.27 | 0.05 | 0.006 | 5e-06
³He比例計数管 | 2.56
6Li liquid scintillator
(Hall A)
Plastic Scintillator
(Hall C)
BF3 Counter
(Hall A)
3He Counter (Tunnel)
No spectrum and directional data in sub-MeV
Neutron Detector | Energy Range | g-ray rejection power | Energy Resolution | Directionality
Liquid Scintillator | 1MeV − 100MeV | Bad | Good | None
BF₃ ,3He Proportional Counter | Thermal − 20MeV | Good | None | None
Proton-recoil Proportional Counter | 10keV − 2MeV | Bad | Good | None
Nano Imaging Tracker (NIT) | Thermal & 100 keV − | Good | Good | Good
H. Wulandari, et al., Astropart. Phys. 22 (2004) 313.
A. Rindi, et al., Nucl. Inst. Meth. A, 272 (1988) 871.

## Slide 4: Neutron Detection Principle by Nano Imaging Tracker (NIT)
880 keV Neutron
10 mm
θ
Proton range
Ep ~ 600 keV
NIT film
Optical Microscope Image
10 mm
H atomic fraction ~ 40%
Neutron
5 mm
Sub-micrometric
3D tracking
T. Shiraishi, et al., PTEP 2021 4, 043H01 (2021)
T. Shiraishi, et al., Phys. Rev. C 107, 014608 (2023)

## Slide 5: Neutron Detection Methods for Various Energies
n + 10B  α + 7Li + 2.31 MeV
4000 barn @ 26meV
7Li (2.5 µm)
α (4.8 µm)
Boron powder
n + 6Li  α + 3H + 4.78 MeV
950 barn @ 26meV
3H (41 µm)
α (7.3 µm)
Vertex
Neutron Capture
Proton Elastic Scattering
Deep Inelastic Scattering
Neutron Energy
Sub-GeV −
100 keV (~ 1 µm)
meV (thermal neutron)
3 fragments
Nuclear recoil
10 µm
70 µm
880 keV neutron
θ
Recoil proton
(580 keV)
n + 14N  p + 14C + 0.62 MeV
2 barn @ 26meV
p (6.5 µm)
Additional target
Self-contained target
Reconstruct 3D trajectory with sub-micron accuracy
 can obtain position, energy and direction

## Slide 6: High-speed Readout and Image Processes
PTS system @ Toho Univ.
6
Z-drive
XY-drive
Epi-illumination microscope
Light guide
(Blue LED)
High-speed Camera
Objective lens
Achieving 0.5 kg/year/machine with 1 μm range cut
Under constructing an upgraded PTS machine in Kanagawa Univ.
 expected to be 1.5 kg/year/machine until Apr. 2024
T. Shiraishi, et al., PTEP 2021 4, 043H01 (2021)
T. Shiraishi, et al., Phys. Rev. C 107, 014608 (2023)

## Slide 7: Convolutional Neural Network
CNN selection effect
(705 keV neutron sample)
Training Samples
Blue: All data
Red：CNN selected
Proton or alpha
Dust
Multi-prong
CNN cut efficiency
Rp (mm)
Signal probability (%)
Input image size
256 x 256 x 20
*still using 2Dconv for 3D image 

## Slide 8: Background in Neutron Detection
g-ray (b-ray)
Clearly distinguishable
There is no background in sub-MeV region
(2 - 14 mm  0.25 - 1 MeV in proton energy)
MeV region can be analyzed excluding single-a (especially 210Po peak around 24 mm)
Exposed 5x107 g-ray/cm2 from 241Am and 60Co
(5 years accumulation of environmental g-ray)
Topological or Range cut
50mm
a-ray
Th star (228Th  208Pb)
5 mm
5.3 mm
（0.5 MeV）
Recoil proton signal
Environmental g-rays cannot become background
5 mm
210Po (24.5 mm)
50μm
228Th
220Rn
216Po
212Bi
224Ra
Single-a in U series

## Slide 9: Calibration with Monochromatic Sub-MeV Neutron
Monochromatic 880 keV neutron exposure from T(p n)3He reaction at AIST
Exposed 7.9 hours with a stable temperature at -26℃
En cos2θ
Data of recoil proton
Range cut (1 mm)

## Slide 10: Calibration – Comparison with Simulation
GEANT4 simulation
*Color corresponding to neutron energy
Scattering Angle
Neutron Energy
0.88 MeV
Proton Energy of Head-on Collision
Detected recoil protons are almost good agreement with kinematical expectation
Proton Range
Detection efficiency for R < 1.5 μm (< 200 keV) is not 100%

## Slide 11: Calibration – Angular and Range Dependency of Detection Efficiency
Horizontal Vertical
Z
NIT
Base
θz
Almost same efficiency

## Slide 12: Environmental Neutron Measurement by NIT @ LNGS
Without shielding!
because there is no sensitivity for muon and gamma
Refrigerator Box
n-Run1
(Nov. 2021)
Surface ground
Motivation of Surface Run
Demonstration of spectrum measurement for environmental neutron and CR-DM search
There is no detailed data in the sub-MeV region even on the surface ground
Underground
n-Run2
(May 2022)
n-Run3
(Jan. 2023)
n-Run4
(Oct. 2023)
Sample2
Compressor
12cm
Thermal
Insulator
10cm
Sample1

## Slide 13: Event Classification
External a-rays are excluded by fiducial volume cut, then events are topologically classified to Single-prong and Multi-prong
Unfortunately, n-Run1 samples accumulated a lot of Radon, we focused on sub-MeV region (2~14mm  0.25~1MeV) of Single-prong event to analyze with background free

## Slide 14: Data/MC Comparison (n-Run1)
Proton Energy
Plane Angle
Zenith Angle
 Due to low efficiency for vertical
Number of Events
MC : 11.9±0.5 event/g/day
Data : 11.1 ± 0.6(stat.) ± 2.4(sys.) event/g/day
Zenith (= Z)
qZenith
f
X
Horizontal sample
Neutron Flux [0.25 ~ 10 MeV]
PARMA model : 9.0 x 10-3 cm-2 s-1
Data : (8.4±1.8) x 10-3 cm-2 s-1
T. Shiraishi, et al., Phys. Rev. C 107, 014608 (2023)
MC : Geant4 + PARMA model
T. Sato, PLOS ONE 10, e0144679 (2015)
T. Sato, PLOS ONE 11, e0160390 (2016)

## Slide 15: MeV Region (n-Run1)
n-Run1
210Po
214Po
sub-MeV
Sub-MeV
Proton Energy
Track Range
214Po
(short life-time)
210Po
(Long life-time)

## Slide 16: Reduction of 214Po Contamination at Drying
n-Run1
n-Run4
(Radon free room, CR1 @ Hall C)
210Po
214Po
sub-MeV
Hall F (NEWSdm facility)
n-Run1
(granite table)
n-Run2, n-Run3
(climatic chamber)
1 order
more 2 order
Proton Energy

## Slide 17: n-Run1 and n-Run2 Results
210Po
(Long life-time)
214Po
(short life-time)
14N + n  p 584keV + 14C 42keV
track range ~ 6.5mm
As expected
  sub-MeV signal increase
  210Po-alpha increase
  Offset background in MeV decrease
Thermal neutron signal can be seen significantly
Thermalized due to surrounding materials?
or attenuation by water contained in rock was suppressed?
14N(n p)14C
Q=626keV
Sub-MeV region
Track Range
Sub-MeV

## Slide 18: n-Run1 and n-Run2 Results (*after reference subtraction)
Proton energy spectrum (linear scale)
(log scale)
Thanks to reduced 214Po contamination, MeV spectrum close to the simulation
14N(n p)14C
210Po
214Po

## Slide 19: Neutron Run Go to Underground
 | Installed Place | 214Po contamination (/g) | Exposure Time (days) | Experimental Scale (g*month) | Analyzed Scale (g*month) | Proton Energy Threshold (keV)
n-Run1 (Nov. 2021 - ) | Surface ground | O(1000) | 29 | 2 | 1.3 | 250
n-Run2 (May 2022 - ) | Surface ground | O(100) | 58 | 20 | 2.1 | 250
n-Run3 (Jan. 2023 - ) | Underground Hall C & F | O(100) | 120 | 30 | 1.4 | 100
n-Run4 (Nov. 2023 - ) | Underground Hall C | O(1) (using CR1) | 120 | 100 | --- | 100
At least 10 g*month scale is needed for underground neutron measurement
Analysis ongoing
Exposure ongoing

## Slide 20: n-Run3 (Underground) Result
Sub-MeV neutron signal clearly decreased because of underground
There are time-independent signal-like tracks below 300 keV
   Non-physical events
Proton energy
210Po
Track Range
Proton Energy in sub-MeV
Preliminary

## Slide 21: n-Run3 (Underground) Result
Horizontal Vertical
Mask
cosθz < 0.05 && range < 3μm
If we avoid low energy & horizontal angle region, there is no excess in sub-MeV region
210Po
Proton Energy in sub-MeV
Preliminary
Signal-like tracks found below 300 keV are all horizontal!
~ 150 keV
~ 150 keV

## Slide 22: n-Run3 (Underground) Result
Horizontal Vertical
Mis-reconstruction of α-track from 214Po accumulated at wet condition?
 Should be checked by next n-Run4 (low 214Po contamination)
Signal-like tracks found below 300 keV are all horizontal!
210Po
~ 150 keV
~ 150 keV
Wet condition
ρ ~ 1 g/cm3
Dry
～1mm
50mm
ρ = 3.1 g/cm3
Mis-reconstruction

## Slide 23: n-Run3 α-ray Analysis
210Po 5.3 MeV
234U, 230Th, 226Ra
~ 4.7 MeV
238U 4.2 MeV
n-Run3 4 month
Range distribution
α energy

## Slide 24: Intrinsic α Activity
α Multiplicity | Expected # of event by Ge detector (g-1 month-1) | # of event from n-Run1 (g-1 month-1) | # of event from n-Run3 (g-1 month-1)
5 (228Th to 208Pb) | 16±2 (Th) | 15±5 | 15±3
1 (238U) | 2.1±0.5 (U) | --- | 8.4±1.4
1 (234U, 230Th, 226Ra) | 6.3±1.5 (U) | --- | 26±3
1 (210Po) | 2.1±0.5 (U) +222Rn contaminated | 165±16 | 790±23
γ-ray measurement by Ge detector
(228Th: 6.0±0.6 mBq/kg)
(226Ra: 0.8±0.2 mBq/kg)
210Po seems to be increased from n-Run1

## Slide 25: Summary
3-dimensional sub-micrometric tracking technique has been developed for NIT analysis
  Achieved 100 keV threshold analysis for recoil proton with 0.5 kg/year/machine
  Analysis speed will be further upgraded to 1.5 kg/year/machine
Neutron run in Gran Sasso
  Surface run (n-Run1, nRun2)
    Succeeded to measure neutron spectrum and direction
    214Po contamination problem was found
       Solved by using radon free room at the sample preparation in current experimental scale
  Underground run (n-Run3, nRun4) Preliminary
    Aiming 100 g*month scale to measure neutron spectrum
    Unknown horizontal background were found in < 300 keV
      Maybe mis-reconstruction of alpha accumulated at the begging of sample preparation?
      If we avoid this region, there is no signal in sub-MeV region as expected
    n-Run4 with further 2 orders lower 214Po contamination is now ongoing
 T. Shiraishi, et al., Phys. Rev. C 107, 014608 (2023)
 T. Shiraishi, et al., PTEP 2021 4, 043H01 (2021)

## Slide 26: Backup

## Slide 27: α-ray accumulation in drying condition
Track
Silver Grain
Wet condition
After dry
ρ = 3.1 g/cm3
ρ ~ 1 g/cm3
Dry
If α-ray create tracks at wet condition, tracks become longer & darker & horizontal because of
Low mass density
Low crystal density
Shrink less than 1/10 thickness
～1mm
50mm

## Slide 28: Comparison of Nuclear Emulsion
60keV C ion
200nm crystal
500nm
44nm crystal
500nm
OPERA type
Nano Imaging Tracker (NIT) type
AgBr:I crystal (SEM)
AgBr:I crystal (SEM)
Optical microscope image
Optical microscope image
1 µm
100mm
60keV C ion
Range ~ 200nm
1 µm
100mm
500MeV/n Fe ion
500MeV/n Fe ion
Ion
Ion

## Slide 29: High Energy Ion Track in NIT
50 μm
C 290 MeV/n
Ar 500 MeV/n
Fe 500 MeV/n
Xe 150 GeV
Pb 150 GeV
Spallation of 150 GeV Lead Beam
Pb
Multiplicity
Emitting angle
Energy
Brightness (dE/dx)

## Slide 30: Actual Multi-prong Events from n-Run1
Low dE/dx
(> 52mm)
1 short + 3 fragment
Th Star event
(5-prong α-decay from 228Th to 208Pb)
Deep Inelastic Scattering by neutron
50 μm
70 μm
*Projection Image
High dE/dx
(46.7mm)
Short
(4.2mm)
High dE/dx
(16.3mm)
25.0mm
(5.41MeV)
 228Th
27.7mm (5.80MeV)
 224Ra
30.8mm (6.22MeV)
 220Rn
35.2mm (6.80MeV)
 210Po
29.6mm (6.06MeV)
 212Bi

## Slide 31: Analysis flow (n-Run2)
Chain tracking (2um)  CNN (2um)  Manual check
  CNN training sample
    Signal
      AIST 880keV neutron @ -26 deg. (EGS012 wash4)
      n-Run2 2 month signal @ -15 deg. (EGS016)
    Noise
      n-Run2 2 month noise @ -15 deg. (EGS016)
Effect of CNN cut to n-Run2 range distribution
Applied cut
Range (um)
95% noise rejection
with 99% efficiency

## Slide 32: Classification
Signal
(R >= 3.5 mm)
Signal
(1 <= R < 3.5 mm)
Noise
27 mm

## Slide 33: 1st Surface Run (n-Run1) for Environmental Neutron Measurement @ LNGS
Motivation
Demonstration of spectrum measurement for environmental neutron and CR-DM search
There is no detailed data in the sub-MeV region even on the surface
* Recoil protons are accumulated O(2) events /(g month) on the surface
Refrigerator Box
Sample2
Compressor
12cm
Thermal
Insulator
10cm
Sample1
n-Run1 Setup

## Slide 34: Single-prong Analysis Result (n-Run1)
Range spectrum of Single-prong signal
0.25 < En < 10 MeV
9.0 x 10-3 cm-2 s-1 @ LNGS surface
Assumed neutron spectrum expected by PARMA model
T. Sato, PLOS ONE 10, e0144679 (2015)
T. Sato, PLOS ONE 11, e0160390 (2016)
Observed signal increase consistent with environmental neutron signal
Detectable neutron spectrum with 2 < R < 14 mm
(0.25 < Ep < 1 MeV)
14N(n p)14C reaction
(Q = 0.62 MeV)

## Slide 35: Simulation for Surface LNGS-run
Zenith angular dependency of neutron spectra expected by PARMA model
En > a few MeV is anisotropic
En < 1MeV is almost isotropic
Anisotropy
T. Sato, PLOS ONE 10, e0144679 (2015)
T. Sato, PLOS ONE 11, e0160390 (2016)
Conditions for PARMA model
Altitude : 1400 m.s.l. Date : 1 Dec. 2021
Latitude : 42.5 deg Longitude : 13.6 deg
Surrounding environment : Ground
Water fraction in ground : 20%
Zenith Angle

## Slide 36: Expected Result from Next Run
Analysis scale vs Expected neutron events
3g 1month
(Next Run)
Expected Sky Map assuming the flat neutron angle
Surface
Underground
3g 1year

## Slide 37: 2nd Surface Neutron Run (n-Run2)
 | Condition | Exposure Time on Surface (day) | Analyzed Mass (g) | Average Rainfall in Assergi (mm/day)
n-Run1 (Run start from 24 Nov. 2021) | Dried on granite table Run @ -20℃ | 2 | 0.65 | 4 **
 |  | 29 | 1.35 | 4 **
n-Run2 (Run start from 25 May 2022) | Dried in chamber Run @ -15℃ | 0 | 0.95 | ---
 |  | 58 | 1.08 | 1.4
** Estimation from Nov. – Dec. 2022 data

## Slide 38: Neutron Flux
Linear
Log
Angle-integrated neutron flux on LNGS surface with water fraction in rock = 20%
PARMA model take into account the neutron reflection by surface ground
Especially, epi-thermal neutrons flux depends on water fraction in rock
*Expected by PARMA model
fast
epi-thermal
thermal

## Slide 39: Underground Neutron Sources
(α n) reaction in the rock by 238U, 232Th(, 235U)
Spontaneous fission of 238U
m spallation
  Negligible in <10MeV region at LNGS underground
H. Wulandari et al., Astropart. Phys., 22, 313-322 (2004)
ρ = 2.71 g/cm3
Estimate myself using NeuCBOT and GEANT4 simulation

## Slide 40: Neutron Spectrum at LNGS Underground
Neutrons attenuated by water contained concrete
Spectrum spread to thermal region
Original
After attenuation
Attenuation
Tunnel structure by GEANT4
n

## Slide 41: m spallation
J.M. Carmona et al., Astropart.Phys. 21 (2004) 523-533
Underground muon rate is about 6 order less than Surface
Almost negligible

## Slide 42: Calibration with ion implantation system

## Slide 43: Accelerators in our laboratory
Low Velocity Ion Implantation
Acceleration Voltage : 5 ~ 200 kV
Temperature : –196 ~ 1000℃
Ion : H, He, B, C, N, O, F, Si, P, Ar, Ti, Fe, Co, Ge, Kr, Xe, CO, CD4, ...
Valence : 1, 2, 3, (4)
Beam current : 10 pA ~ 100 μA
1MV Tandem Pelletron Ion Accelerator
Acceleration Voltage : 0.5 - 1 MV
Ion : He++, H, Li, B, C, O, Si, Ni, Cu, ...
Valence : 1, 2, ???
Detector : Si semiconductor detector x 2
Beam size : φ 2 mm
Analysis : RBS-channeling, PIXE, Nuclear Reaction, ERDA

## Slide 44
Detecting directly exposed protons
Ion implanter @ Kanagawa Univ.
Recognized tracks
5 µm
↓
↓
↓
↓
↓
↓
150 keV proton
↓
↓
↓
Automated detection
↓
↓

## Slide 45
１
IFM
NIT
5mm
Ion
Beam
X(mm)
Count
X(mm)
Beam
・・・
Detection event vs location(mm)
Ion exposure ➡ Move the plate up ➡ Ion exposure with different energy ・・・
➣One sample is exposured with protons of multiple energies.
・Beam areas and Mask areas are created.
１
IFM
NIT
5mm
Ion
Beam
X(mm)
Count
Detection event vs location(mm)
Detection event vs location(mm)
150keV
80keV
100keV
60keV
50keV
40keV
30keV
Detecting directly exposed protons

## Slide 46
１
IFM
NIT
5mm
Ion
Beam
X(mm)
Count
X(mm)
Beam
・・・
Mask
Ion exposure ➡ Move the plate up ➡ Ion exposure with different energy ・・・
➣One sample is exposured with protons of multiple energies.
・Beam areas and Mask areas are created.
１
IFM
NIT
5mm
Ion
Beam
X(mm)
Count
Detection event vs location(mm)
Detection event vs location(mm)
150keV
80keV
100keV
60keV
50keV
40keV
30keV
Detecting directly exposed protons

## Slide 47
Chart: Proton Energy (keV)
| 195 | 150 | 80 | 100 | 60 | 50 | 40 | 30
0.2umCut | 1 | 0.96465 | 0.34637 | 0.5034 | 0.16359 | 0.1016 | 0.096345 | 0.089151
0.2umcut_3sigmaPhi | 0.70226 | 0.64713 | 0.23479 | 0.33416 | 0.090891 | 0.051441 | 0.047493 | 0.036169
0.7umCut | 1 | 0.74365 | 0.11705 | 0.25283 | 0.071703 | 0.063236 | 0.068487 | 0.079568
phi3sigma_0.7umCut | 0.42662 | 0.47577 | 0.043309 | 0.12962 | 0.018767 | 0.015653 | 0.017404 | 0.024471
×:0.7μmCut
×:0.2μmCut
*With angle information: Gaussian fit results for angle distribution
Number of components with angles within 3σ
Efficiency=
(#of detection@NIT)/(#of deteciotn@IFM)
150keV angle distribtion
For improvement...
Use high-density NIT？
➤increase the silver density
➤Decrease crystal size
1σ=4.5
3σ
With angle information
・The angle distribution is Gaussian-fitted,
and only components with angles within
3σ of it are considered as detection events.
Detecting directly exposed protons

## Slide 48
X(mm)
Ion exposure ➡ Move the plate up ➡ Ion exposure with different energy ・・・
➣One sample is exposured with protons of multiple energies.
・Beam areas and Mask areas are created.
１
IFM
NIT
5mm
Ion
Beam
X(mm)
Count
Exposure Energy(keV)
Measured Energy (keV)
Detection event vs location(mm)
・The correlation between irradiation energy and irradiation position is also visible.
150keV
80keV
100keV
60keV
50keV
40keV
30keV
Detecting directly exposed protons

## Slide 49
φ(°)
ー150keV
ー100keV
ー80keV
Ep150~80くらいのEp分布で
エネルギー分解能を持つことを示す
100KeVがエネルギー分解能を持つ
Ep(keV)
ー150keV
ー100keV
ー80keV
Ep (measured)
Chart: Proton Energy (keV)
| 195 | 150 | 80 | 100 | 60 | 50 | 40 | 30
0.2umCut | 1 | 0.96465 | 0.34637 | 0.5034 | 0.16359 | 0.1016 | 0.096345 | 0.089151
0.2umcut_3sigmaPhi | 0.70226 | 0.64713 | 0.23479 | 0.33416 | 0.090891 | 0.051441 | 0.047493 | 0.036169
0.7umCut | 1 | 0.74365 | 0.11705 | 0.25283 | 0.071703 | 0.063236 | 0.068487 | 0.079568
phi3sigma_0.7umCut | 0.42662 | 0.47577 | 0.043309 | 0.12962 | 0.018767 | 0.015653 | 0.017404 | 0.024471
×:0.7μmCut
×:0.2μmCut
*With angle information: Gaussian fit results for angle distribution
Number of components with angles within 3σ
Angular distribution
Energy resolution
47%(@100keV)
irradiation direction
Efficiency=
(#of detection@NIT)/(#of deteciotn@IFM)
Detecting directly exposed protons

## Slide 50
Implementing CNNs to improve efficiency of analysis
ああ
Noise99.5% cut
Noise99.9% cut
Results the signal classified using
the learning model.
・Neutron exposured sample
・n-Run2 sample
Signal
・γ-ray exposeured sample
・n-Run2 sample
Noise
The selection accuracy is 20-50% at around 1 μm,
and almost 100% at 2 μm and above.
Recoil proton tracks selected
from880 keV neutron beam
samples were automatically
classified using CNN.

## Slide 51: Monochromatic Neutron Calibration

## Slide 52: Calibration with Monochromatic Sub-MeV Neutron
Sample temperature profile
PID control system
Monochromatic 880 keV neutron exposure at AIST

## Slide 53: Analysis of 880keV Sample
Optical Mark (LT)
X projection is uniform
Beam center
Attenuation by NIT or glass

## Slide 54: Simulation
Neutron Scattering Model:
G4HadronElasticPhysicsHP
G4HadronPhysicsShielding
Electromagnetic Model:
G4EmLivermore
*Line color corresponding to neutron energy
Tracking step for recoil proton: 0.1mm
Angular dependency of Energy and Flux in T(p, n)3He reaction is considered

## Slide 55: Simulation of Neutron Exposure
Color corresponding to neutron energy
Energy
Fluence
Considering energy and fluence for each angle

## Slide 56: Correlation of Proton Energy and Range in NIT (GEANT4)
E ≈ 60.68 + 553.16 - 474.59 (keV)
NIT density : 3.2 g/cm3

## Slide 57: Angular dependency
0 < cosθz < 0.1
0.1 < cosθz < 0.2
0.2 < cosθz < 0.3
0.3 < cosθz < 0.4
0.4 < cosθz < 0.5
0.5 < cosθz < 0.6
0.6 < cosθz < 0.7
0.7 < cosθz < 0.8
0.8 < cosθz < 0.9
0.9 < cosθz < 1.0
Data
Simulation
AIST 2019 : 880keV neutron sample

## Slide 58: Neutron Exposure from 7Li(p n)7Be@ AIST 25 Aug. 2022
7Li(p n)7Be
n 100 keV
n 200 keV
Proton-recoil Method
Nuclear-reaction Method
Attenuator
Following nuclear reactions occurred after attenuation
n + 10B  α + 7Li + 2.31MeV
n + 6Li  α + T + 4.78MeV
Behind the attenuator

## Slide 59: Detected event in Boron or Lithium contained sample
n + 10B  α + 7Li + 2.31MeV
7Li (~2.5mm)
α (~4.8mm)
Boron powder
② n + 6Li  α + T + 4.78MeV
T (~41mm)
α (~7.3mm)
Vertex

## Slide 60: Detected event in Boron or Lithium contained sample
③ n + 14N  p + 14C + 0.62MeV
p (~6.5mm)
There are too many events around 6.5 mm…

## Slide 61
Atom | Mass fraction | Mol / Ag mol |  | σ @ 1eV
N | 3.74% | 0.694 | 14N 99.6% | 0.3 barn
B | 0.042% | 0.01 | 10B 20% | 620 barn
Simulation in B contained NIT
Polyetheren
B contained NIT
Expectation of event rate:
event(N/B)

## Slide 62: Calibration of α-ray detection

## Slide 63: Radon Daughter Contamination
Sample | Condition | Analyzed mass (g) | # of internal event (/g) | # of top a (/cm2)
Run16 ID1 Aside | Dry in Rn free room (shielded) | 0.24 | 4 +- 4 | 0.9 +- 0.2
Run16 ID1 Bside |  | 0.47 | 4.3 +- 3.0 | 1.1 +- 0.2
Run16 ID2 Aside | Dry in Rn free room (no-shielded) | 0.50 | 8.0 +- 4.0 | 0.3 +- 0.1
Run16 ID2 Bside |  |  |  | 
Run15 ID3 Aside | Dry in Rn free room + Hall F (35min) | 0.27 | 4 +- 4 | 0.48 +- 0.14
Run15 ID3 Bside |  | 0.38 | 3 +- 3 | 1.1 +- 0.2
Run15 ID5 Aside | Dry with buffer box in Rn free room | 0.58 | 43 +- 9 | 0.40 +- 0.09
Run13 ID11 | N2 purged dry | 0.16 | < 14 (90% C.L.) | 0.1 +- 0.1
Run13 ID8 | Normal dry | 0.08 | 650 +- 90 | 50 +- 3
Run7 | Normal dry in Shield | 0.44 | 220 +- 20 | 11.0 +- 0.5
n-Run1 | Dry outside chamber | 0.65 | 2200 +- 60 | 280 +- 6
(Almost thin tracks)

## Slide 64: Calibration of alpha-ray Energy (Ea) by Th star
228Th (6±1) mBq/kg
(Ge Detector)
Th star event is useful for calibration of run condition, such as brightness or E-R relation
It should be accumulated during run, and 5 prong event can be identified as 228Th to 208Pb
  5.423, 5.685, 6.288, 6.779, [8.785 or 6.051] MeV
10 mm
Th Decay Chain
~64%
~36%
Suggested by Valeri
 15 events/(g 28day)
2002年アイソトープ手帳

## Slide 65: a-energy calibration with Th star
n-Run1 Sample 2 & 3
228Th
224Ra
212Bi (36%)
216Po
212Po (64%)
220Rn
Geant4 Simulation
DEa,FWHM / Ea ~ 5%

## Slide 66: Concerning the excess in MeV region - Understanding for 210Po contamination -
Multiplicity of a in 1event | Expected number of event (31day * 0.71g) | Number of event from Data
5 (228Th to 208Pb via 212Po) | 6.4 (Ge) | 5
5 (228Th to 208Pb via 208Tl) | 3.6 (Ge) | 3
4 (224Ra to 208Pb via 212Po or 208Tl) | ~1 (Ge) | 1
4 (226Ra to 214Bi) | 1.5 (Ge) | 2
3 (222Rn to 214Bi) | ~0.1 (Ge) +222Rn contaminated | 3
1 (238U, 234U, 230Th) | respectively 1.5 (Ge) | Respectively < 3
1 (210Po) | 1.5 (Ge) +222Rn contaminated | 80?
226Ra 0.8 mBq/kg (Ge)
228Th (6±1) mBq/kg
(Ge Detector)
Contaminated from air

## Slide 67: Ea Spectrum
210Po peak (5.304MeV) have time dependence
 Accumulate during the Run
214Po peak (7.687MeV) ???
 214Po should be 3prong!
Can this seem as
internal 1prong???
U decay chain
No time dependence excess
 Short life time alpha decay?
2002年アイソトープ手帳

## Slide 68
Th a level
210Po (dry, -20℃)
214Po (dry, 20℃)
214Po (wet, 20℃)
n-Run1 29day data
218Po or 222Rn (wet)???
214Po (dry)
210Po (dry)
214Po (wet)
218Po or 222Rn (wet)???

## Slide 69
Th a level
210Po (dry, -20℃)
214Po (dry, 20℃)
214Po (wet, 20℃)
LNGS Run2021 28day sample
218Po or 222Rn (wet)???
Brightness > 1500
Brightness < 1500
210Po (dry)
214Po (dry)
214Po (wet)
218Po or 222Rn (wet)???
Horizontal
Vertical

## Slide 70: 2prong Anomaly
Multiplicity | # of event from 2day sample (/g) | # of event from 28day sample (/g)
3 (222Rn to 210Pb) | 3 +- 3 | 6 +- 3
2 | 72 +- 14 | 83 +- 11
1 (>30mm 214Po like) | 1770 +- 70 | 2470 +- 60
3prong is too few if we assume as contamination of 222Rn
2prong/1prong ratio ~ 4%
2prong cannot be explained by 222Rn contamination
Why 2prong detected such too many?

## Slide 71: 2prong Analysis
218Po
(6.003MeV)
214Po (7.687MeV)
Longer due to wet
Horizontal  wet like
(qz1 > 80° && qz2 > 80°)
Non-Horizontal  dry like
(qz1 < 80° && qz2 < 80°)
2prong is likely double a-decay of 218Po and 214Po

## Slide 72: Can these be explained by Radioactive Aerosol?
214Pb
218Po+
Radioactive Aerosol (0.1~1mm diameter)
Rock
226Ra
214Po
214Bi
3prong
222Rn
3.8d
3.1m
27m
20m
a
a
a
b
b
Gas
210Pb
160ms
a
2prong
1prong
 | 2prong/1prong ratio
Rough expectation by contamination of Radioactive Aerosol | 6.6% (from life time ratio)
Data (GS 2day sample) | 4.1%
Data (GS 28day sample) | 3.3%
Contamination
Why is 3prong too few from data???
218Po is known to be injected up to a few mm depth in Silicon or Cupper

## Slide 73: Intrinsic Activity (Chamber dry & Underground)
α source | Energy (MeV) | Event Rate (events/g/4month) | Total Activity (mBq/kg)
218Po, 214Po (Rn short decay) | 6.003, 7.687 | 50 +- 12 | 
210Po (Rn long decay) | 5.304 | 1138 +- 56 | 10.97 +- 0.54
234U, 230Th, 226Ra | 4.775, 4.668, 4.784 | 103 +- 17 | 10.0 +-1.6
238U | 4.198 | 33 +- 10 | 3.2 +- 0.9
These activities are almost 3.3 mBq/kg respectively
Confirmed to be 2 orders less by CR1 dry
According to Fabio’s paper (Astropart. Phys. A 80, 16 (2016)),
226Ra activity is 2.4 mBq/(1kg high deionized gelatin)  0.8 mBq/(1kg NIT)
Because of radiative equilibrium, 234U, 230Th, 226Ra, 238U should be same activity
n-Run3 4 month sample
Contaminated from air
Accumulated during Run

## Slide 74: Multi-prong Neutron Inelastic Scattering

## Slide 75: Multi-prong Analysis
Candidates
μ spallation (photo-Nuclear)
μ- capture (p + μ- → n + vμ : CC weak interaction)
N(Z, A) + μ- → N'(Z-1, A)* + vμ
N‘(Z-1, A)* → N’(Z-??, A-??) + (n + p + a …)
Neutron inelastic scattering
We found 17 events/(0.65g*28day) with multiplicity >= 3 after excluding a-decay
*Projection Image
Count (/0.65g/28day)
~50 event
< 0.01 event
< 0.01 event
More likely neutron inelastic scattering!
Rough estimation of the number of fragments by Geant4
Multi-prong Analysis

## Slide 76: Interesting Multi-prong Event from 28day sample
Grey bottom-out
(> 125.5mm)
Bright bottom-out
(> 35.3mm)
Short
(2.1mm)
Thin top-out
(> 38.9mm)
Projected Image
Grey bottom-out
(> 72.1mm)
What is this ???
1short + 4fragment (Multiplicity 5)

## Slide 77: Kink and 2prong Events are Excluded in Current Analysis
Although I feel there are many kink and 2prong events, they are rejected in current analysis because they might be Rutherford scattering or 2prong of a-ray
They might be included after reduction of MeV excess
Short
(2.3mm)
1short + 1fragment ?
Rutherford scattering ?
Bright
(15.5mm)
2fragment ?
2prong a decay ?
Bright
(53.3mm)
Grey
(50mm)
1short + 1fragment ?
Rutherford scattering ?
Bright
(88.6mm)
Short
(1.4mm)

## Slide 78: Topological Analysis of Neutron Inelastic Scattering
Data is topologically similar to Simulation?
  Geant4 probably has big systematic errors for nFragment because there are no data.
For more detail kinematical analysis, Fragment's Angle, Range (Energy), and Brightness (dE/dx) should be used.
Comparison of Simulation and Data
Preliminary
Event rate (/0.65g/28day)
Excluded from current analysis
Data

## Slide 79: Kinematics of Neutron Inelastic Scattering (Suggested by Gianni and Sato-san)
n
N
f
N
n
Initial State
Final State
If assume , transverse momentum (pT) balance can be calculated
Transverse momentum (pT) should be a good kinematical parameter because it is Lorenz invariant!
: Neutron initial momentum
: Recoil Nuclei momentum
: Fragment momentum
neutron
Nuclei
fragment

## Slide 80
Nuclei
Fragment
Fragment
Fragment
Relative pT to recoil nuclei
Vector sum of relative pT to recoil nuclei
Relative angle to incident neutron
Relative angle to recoil nuclei
Trying to calculate kinematical parameter on the MC-base…
pT Vector Sum cannot be 0 due to:
  Cannot include pT of recoil Nuclei
  Missing pT by neutron
Incident Direction

## Slide 81
Fragment
Recoil Nuclei
Ag
Br
CNO
I

## Slide 82: Muon Simulation in NIT
Simulation m- Spectrum
Em- (MeV)
EXPACS (PARMA model) @ LNGS
Em- (MeV)
Flux (/cm2/s/MeV)
Input to Geant4
Physics Model
Muon Photo-Nuclear Interaction
Muon Capture
Radioactive Decay
Livermore for EM
10keV~500GeV m- flux : 0.0114 /cm2/s

## Slide 83: Neutron Simulation in NIT
Simulation Neutron Spectrum
En (MeV)
EXPACS (PARMA model) @ LNGS
En (MeV)
Flux (/cm2/s/MeV)
Input to Geant4
Neutron Physics Model
Elastic : NeutronHPElastic
Inelastic : NeutronHPInelastic
EM : Livermore
10keV~1GeV n flux : 0.0192 /cm2/s

## Slide 84: Neutron Inelastic Simulation in NIT
Count (/0.65g/28day)
Atomic Mass of Recoil Nuclei
Extract only “Inelastic” of neutron physics process
Atomic Mass of Recoil Nuclei vs Fragment particles
Ag
Br
CNO
I
12C(n a)9Be

## Slide 85: Neutron Inelastic Simulation in NIT
Extract only “Inelastic” of neutron physics process
Range of Recoil Nuclei and Fragment
Fragment
Recoil Nuclei
Count (/0.65g/28day)
Recoil nuclei, almost CNO, is up to 10 mm in maximum
We identified shortest track with less than 10 mm as recoil nuclei, and remained tracks as fragments