## Slide 1: The Deep Space Atomic Clock and Potential Scientific Applications
Eric Burt, John Prestage, Robert Tjoelker, Eric Tardiff, Daphna Enzer, Da Kuang, Dave Murphy, David Robison, Jill Seubert, Rabi Wang, James McKelvey, Vladimir Iltchenko, Andrey Matsko, and Todd Ely
Jet Propulsion Laboratory, California Institute of Technology
Joint funding for the DSAC project provided by the NASA offices of Space Technology Mission Directorate (STMD) and Space Communications and Navigation (SCaN)
9th Symposium on Frequency Standards and Metrology

## Slide 2: JPL Trapped Ion Clock Overview
Key Performance Features:
106-107 199Hg+ trapped ions
  No wall collisions, high Q microwave line
  Buffer gas cooled to ~300K
  Multi-pole ion trap – lower systematics
State selection/detection:
  Optical Pumping from 202Hg+ lamp
  1-2 UV photons per ion per second scattered
High Clock Transition:
  40,507,347,996.8 Hz – low magnetic sensitivity
Adapts to variety of Local Oscillators - flexible
Key Reliability Features: - practical
No Lasers
No Cryogenics
No Microwave cavity
No Light Shift
Low Consumables
Microwave Input
Input light system
Detection light system
Local Oscillator
Controller
QP
MP
shuttle

## Slide 3: 199Hg Ion Clock Development Timeline 40.5 GHz transition in 199Hg+, optically pumped with 202Hg+ discharge lamp, buffer gas cooled
1993 7x10-14/T1/2 , long term drifts < 2x10-16/day.
1996 2 x 10-14/T1/2 advanced LOs: CSO, H-masers
1999 Multi-pole ion trap reduction of 2nd Order Doppler Shift
2006 High temp bake-out -> sealed getter pump
Compensated Multi-pole: UTC quality timekeeping with a single clock
- long term variations below < 3x10-17/day.
GPSIII MAFS: 50W total power demonstrated.
high bake sealed operation demonstrated
2010-16 Ultra-stable multipole ion trap reference clock
2016-present - Miniature clock development: 1 liter (next talk)
2019-2021 DSAC operates in LEO
2022 DSAC-2 – Low 30W, long-life concept for NASA & DoD
low SWaP ion tube
2022, low TRL & SWaP ACES breadboard demo
5th FSM
7th FSM
8th FSM
9th FSM
/PLO
Notes: - DSAC uniqueness: EM trap (no wall collisions) and 40 GHz (lower environmental sensitivies)
- Over 30 years experience with technology at JPL
- Best short term: 2e-14/sqrt(tau)
Best long term: < 3e-17/day drift - UTC quality in a single clock
Trapped Hg+: no wall collisions, highest clock frequency
- 50W, 19 kg DSAC operated in space for 2 years
- 30W, 10 kg DSAC-2 proposed

## Slide 4: Deep Space Atomic Clock
A Technology Demonstration Mission
Successful launch, June 25, 2019
Notes: STMD: Space Technology Mission Directorate
SCaN: Space Communications and Navigation
SST: Surrey Satelite Technologies

## Slide 5: Some Primary Physics Package Subsystems
Multi-pole Trap Electrodes
Quadrupole Trap Electrodes
Titanium Vacuum Tube
Mercury UV Lamp Testing
DSAC Demonstration Unit
Electron Emitter

## Slide 6: DSAC Frequency Stability Measurements vs UTC via GPS
UTC(k)
Adev(y(t))
GPS
Satellites
GPS
Rcv.
DSAC Payload in LEO
Hg+ stability vs UTC (> ~20,000 s) via GPS time transfer:
GPS
Rcv.
Challenges of measurements in space: Noise Sources
Clock noise
GPS noise
Require precise orbit determination to remove relativistic effects – need J2 corrections
OD
Final result: measure all three

## Slide 7: DSAC Results: Calibrating GPS Receiver and Relativistic Effects
GPS phase dynamically corrected for gravitational effects (red shift)
Must include higher order terms (J2) in the gravitational potential
No corrections
Relativity corrections, but no GPSR temperature corrections
Measurement floor after relativity and GPSR temp correction
Estimated clock noise
Notes: Must correct for GRS
Must correct for GPS rcvr tempco
Still just an upper bound for actual clock
Includes residual OD noise, measurement system noise, and clock noise

## Slide 8: DSAC Results
Stability at one-day of 3e-15
significantly better than required 2e-14
Drift of 3.0e-16/day
establishes space clock record
Two subsequent long runs had similar long term stability
GPS + OD + DSAC
DSAC
Optimized DSAC
DSAC measured on the ground &
w/ a Maser LO
DSAC requirement
Run #1 (late 2020):
+3.0(0.6)e-16/day
Notes: 3e-15 at one day
3e-16/day drift – record in space

## Slide 9: DSAC Environmental Sensitivities
Radiation: SAA-induced USO drift variations taken out by clock control loop
Notes: Radiation on USO controlled
Magnetic 100x lab below noise
Ion number understood
Collision and light shifts at noise level
Temperature at 1e-14/C WITH NO THERMAL REGULATION – no other space clock can do this
Path to 2e-15/C known

## Slide 10: DSAC Environmental Sensitivities
Radiation
Magnetic Shifts: below measurement noise floor
250 mG/orbit - 100x lab!
Strength of Hg+ technology
Notes: Radiation on USO controlled
Magnetic 100x lab below noise
Ion number understood
Collision and light shifts at noise level
Temperature at 1e-14/C WITH NO THERMAL REGULATION – no other space clock can do this
Path to 2e-15/C known

## Slide 11: DSAC Environmental Sensitivities
Radiation
Magnetic Shifts
Actual freq offsets
Modeled
Ion number variations:
Second order Doppler shifts
Actual = model: well understood
Notes: Radiation on USO controlled
Magnetic 100x lab below noise
Ion number understood
Collision and light shifts at noise level
Temperature at 1e-14/C WITH NO THERMAL REGULATION – no other space clock can do this
Path to 2e-15/C known

## Slide 12: DSAC Environmental Sensitivities
Radiation
Magnetic Shifts
Actual freq offsets
Modeled
Ion number variations
Collision shifts
Residual frequency offsets with other effects removed
Stable at 4e-16/day level
Notes: Radiation on USO controlled
Magnetic 100x lab below noise
Ion number understood
Collision and light shifts at noise level
Temperature at 1e-14/C WITH NO THERMAL REGULATION – no other space clock can do this
Path to 2e-15/C known

## Slide 13: DSAC Environmental Sensitivities
Fundamental systematic effects are well-understood and below the measurement noise floor
Radiation
Magnetic Shifts
Actual freq offsets
Modeled
Ion number variations (second order Doppler)
Collision shifts (Background gas)
Light shifts: Not measured, but estimated at <3e-15 measurement noise floor
Temperature sensitivity:
Not fundamental
T-sensitivity of these
Overall: 1e-14/C with NO thermal regulation
Path to unregulated 2e-15/C is known
Notes: Radiation on USO controlled
Magnetic 100x lab below noise
Ion number understood
Collision and light shifts at noise level
Temperature at 1e-14/C WITH NO THERMAL REGULATION – no other space clock can do this
Path to 2e-15/C known

## Slide 14: DSAC Clock lifetime: > 7 years
Mercury vapor evolution
Extrapolate current trap load time: 7-year life
Likely explanation: Hg/Au amalgamation – gold will be removed in future versions
Trap load time variations
Vacuum Tube
Notes: All three aging processes extrapolated to > 7 years
Hg and neon understood – path to longer life known
Working on lamp

## Slide 15: DSAC Clock lifetime: > 7 years
Mercury vapor evolution
Extrapolate current trap load time: 7-year life
Neon evolution
Shuttle decay measurements calibrated to neon pressure
Extrapolate to > 8-year life
Method to extend understood
Nov, 2016:
P = 4.7e-6 torr
May, 2020:
P = 2.6e-6 torr
June, 2021:
P = 1.3e-6 torr
Trap load time variations
Trap “shuttling” decay
Vacuum Tube
Notes: All three aging processes extrapolated to > 7 years
Hg and neon understood – path to longer life known
Working on lamp

## Slide 16: DSAC Clock lifetime: > 7 years
Mercury vapor evolution
Extrapolate current trap load time: 7-year life
Clock lifetime > 7 years with known methods to extend this
Neon evolution
Shuttle decay measurements calibrated to neon pressure
Extrapolate to > 8-year life
Nov, 2016:
P = 4.7e-6 torr
May, 2020:
P = 2.6e-6 torr
June, 2021:
P = 1.3e-6 torr
Trap load time variations
Trap “shuttling” decay
Optics aging
13.7% change in 13 months
=> > 7 years total life
Most likely lamp
220169 counts
September 2020
190062 counts
October 2021
Lamp Assy.
Vacuum Tube
Notes: All three aging processes extrapolated to > 7 years
Hg and neon understood – path to longer life known
Working on lamp

## Slide 17: Future Position, Navigation and Timing Clocks

## Slide 18: Position, Navigation and Timing (PNT) Clock Metrics
Takeaway: for PNT, need more than performance: SWaP and operability equally important
*Mean time between physical interaction is only estimated for these standards because this data was not reported

## Slide 19: Technology Maturation Task (TMT) Space and Ground Clocks – common performance properties
Common requirements with broader market and commercial interest:
  Ground Clocks: H-maser (e.g. DSN, VLBI), Cs beam replacement (e.g. telecom)
  Space Clocks: DSAC2, GPS & GNSS
Operability, reliability, and manufacturability paramount.
Instrument Life: 5-10 years
Frequency stability class:
  Local Oscillator = 1E-13 class USO
  1E-13 at 1 second, 2E-13/τ½ , 1E-15 at 1 day. Long term drift <5E-16/day
SWaP class volume driver:
  Space < GPS clock footprint
  Ground < 5071 Cs chassis height (5.25”).
Frequency stability with Quartz USO as the LO
Flight Prototype
Ground Prototype
DSAC (STMD, SCaN)
TMT (SCaN)
DARPA/ONR development
miniature, tactical
(1E-11/root tau)
How much can we reduce SWaP and increase Life while maintaining maser-like performance?
Free running USO
DSAC-1 in space
DSAC-2 expected

## Slide 20: TMT
3D-models
DSAC (+ USO and controller)
TMT (all-inclusive)
TMT approach
Simplified trap architecture: QP-only
Eliminate 1 trap driver
Simplify controller architecture
Integrate electronics
Eliminate empty space
Improve optics efficiency

## Slide 21: TMT Performance Goal
1e-13 at 1 s, <1e-15 at a day, ~50 ps at a day, 34 W, 10 kg
DSAC-1
USO
PIU
+
+
TMT Concept
50 W, 19 kg
34 W, 10 kg
Notes: DSAC-2: SWaP: 50W/19kg -> 34W/10kg
Performance: 1e-13 at 1 s, 2e-13/sqrt(tau), 1e-15 at a day, 10 ps at a day
DSAC-2 selected for VERITAS, but cancelled due to a budgetary shortfall

## Slide 22: Prototype Build
3D Model
Prototype build with electronics

## Slide 23: Prototype Initial Results
Clock transition spectroscopy
Zeeman spectroscopy
Trap lifetime ~ 1 day
Contrast with increasing microwave power
=> excellent coherence
Initial temperature sensitivity: < 2e-14/C
Magnetic sensitivity: in progress
4e-13/sqrt(tau) (maser LO)
Further optimization: SNR supports 3e-13/sqrt(tau)
Preliminary – not optimized

## Slide 24: Prototype Electronics
DSAC synth
Prototype
DSAC trap/EED
Prototype

## Slide 25: Light source aging
Too much Hg
Too little Hg
Mercury Oxide: where does the oxygen come from?
+
+
+
+
+
+
+
+
Hg+/e-
neutral
-
-
-
-
-
-
-
-
E field
SiO2
HgO recombination
“Plasma Sheath”
Separation of charge
Need:
Incorporate sapphire
Methods to calibrate/quantify Hg
High energy Hg+ wall collisions dissociate O2
-
-
-

## Slide 26: Plasma discharge light source development: Incorporating sapphire
See next talk for another approach…
ALD sapphire coated
Bulk sapphire
Sapphire insert

## Slide 27
Optical Hg Drop Imaging
Lamp Content Quantification
Differential Scanning Calorimetry (DSC)
Long-term spectral & life monitoring
Ar and Hg
discharge
Lamp Life and Quantification

## Slide 28: TMT Summary
Performance:
  on track for 2e-13/sqrt(tau), 1e-15 at a day
SWaP:
  on track for 10 L / 34 W
Lifetime:
  have built a sapphire-coated lamp
  Producing best clock signal
  life testing in progress

## Slide 29: Science Applications

## Slide 30: Deep Space Nav using 1-Way Range and a Spacecraft Clock
Deep space navigation: a science enabler
One-way signal with bad clock:
Time delay + Doppler + Fit
Two-way signals:
S/C is transponder, no clock
Distance + Velocity
Real-time nav, but imprecise
Real-time nav and precise
Precise, but not real time
One-way signal with good clock:
Time delay + Doppler + Fit
Notes: Two-way nav: s/c transponder – precise but not real time
One-way with bad clock plus model – in situ real time, but not precise
One-way with good clock plus model – in situ real time AND precise

## Slide 31: Radio Occultation
Credit: JAXA Akatsuki RS Team
RO Experiment Geometry
Must correct for USO drift
Depends heavily on data used
Baseline mean:
- 7mHz (red), -2mHz (blue)
Ingress
absorption
Frequency shift
Data used for drift est.
Uncertainty in RO result due to USO drift estimate
Ingress point
Occultation data
Courtesy Tatiana Bocanegra-Bahamon, JPL
Notes: - RO: atmosphere leads to refraction, absorption, and frequency shift in microwave signal
Measure this => measure atmospheric characteristics
USO used to generate s/c timing
Must characterize USO drift - problematic

## Slide 32: RO Simulation: DSAC vs. USO
Effect of reducing the clock error
Neglecting OD and transmission errors
propagate through the multivariate and multistage occultation measurement model using MC simulations.
1% vs. factor of 2
Part per 1000 vs. 10%
Courtesy Tatiana Bocanegra-Bahamon, JPL
Notes: Tatiana Boanegra-Bahmon (JPL) simulation of RO: compare USO to DSAC
1-2 orders improvement for both refraction and df

## Slide 33: Simulation Results for Venus Cruise LPI Measurement
Parametrization: . LPI test with identical clocks (DSAC in space and on ground)
Multi-arc covariance analysis (using NASA-JPL MONTE software):
formal uncertainty
: accuracy of initial frequency calibration (i.e. expected to be )
: expected 10-day frequency stability
DSAC-1 verified )
(10-day DSAC-2 stability)
Current best
Possible with DSAC-2
LPI/LLI slides courtesy of G. Cascioli, F. de Marchi, and L. Iess
Notes: Perform simulation, including clock, OD, and measurement system noise
Looking just at alpha, may be able to get a 5x improvement over Galileo measurement
Note, Galileo meas. Struggled with temperature and magnetic field sensitivity – not present with DSAC

## Slide 34: Summary
DSAC: first trapped ion atomic clock in space
< 3e-15 at 1 day, 3e-16/day drift, 300 ps at 1 day (meas. system limit)
Lowest environmental sensitivities
A lower SWaP breadboard prototype has been built
Enables: deep space nav, science, and applications requiring autonomy
Science examples: Relativity, Radio Occultation

## Slide 35: Supporting Slides

## Slide 36: DSAC 2019-2021 Operations Overview
DSAC has proven to be very reliable and robust
Mission operations from 8/20/2019 to 9/18/2021 for a duration of 760.2
Operations periods include:
  USO on: 84.6%
  Clock on: 78.2%
  GPS data: 74.9%
10 long runs (67 days longest)
All run terminations due to S/C safe modes
2 known radiation-induced faults that impacted clock control – recovered
No known clock hardware faults
Clock up time

## Slide 37: Long-Term Performance: Drift - 3 long runs, each ~ 2 months
Strength of trapped ion technology: long-term autonomous operations
Confirmed:
Systematic effects “corralled”
Very little linear drift with no thermal regulation – two orders magnitude less than Rb
Time keeping applications
Run #1 (late 2020):
+3.0(0.6)e-16/day
Run #2 (early 2021):
Small ion cloud
correct for ion number
-5.0(0.8)e-16/day
Run #3 (late 2021):
Small ion cloud
-5.8(0.7)e-16/day

## Slide 38: Fundamental Science Enabled by DSAC

## Slide 39: Gateway Provides Precision Time & Frequency for Science: measurements enabled by a precise clock
1/r2 law of gravitation:
  All measurements to date agree that the exponent is precisely “2”,
  some theories of gravity predict a small deviation from this law.
  Precise range of a drag-free spacecraft in deep space using a DSG clock together with laser ranging from the DSG to the moon could yield an improved limit on the correction [1].
Cosmological constant (CC):
  The CC is added to Einstein’s equations to explain an observed accelerating expansion of the universe.
  Theoretical predictions disagree by many orders of magnitude with measured limits.
  Ranging to a drag free spacecraft in deep space together with a DSG clock and laser ranging to the moon could improve limits on the cosmological constant. The measurement could also be enhanced by inferred outer planetary orbits as determined by nearby precise spacecraft ranging [1, 2].
Dark Matter (DM):
  DM in some form constitutes 90% of all matter and is required to explain the structure and dynamics of galaxies. So far unobserved.
  Time variations in spacecraft ranging enabled by a DSG clock relative to a precisely known DSG orbit using laser ranging could be an indication of dark matter fluctuations on the spatial scale of the range [7].
  Similar dark matter searches could also be carried out with a network of clocks, including one on the DSG and others on earth, on the spatial scale of the earth-moon distance.

## Slide 40: Gateway Provides Precision Time & Frequency for Science: measurements enabled by a precise clock
Gravitational Red Shift (GRS):
  This general relativistic effect causes clocks deep in a gravitational potential to tick slower than those higher in the potential. All measurements to date agree with GR, however GR is known to be incomplete and some new theories predict variations in the GRS.
  Due to it’s highly eccentric orbit, the DSG is an ideal platform to measure variations in the GRS by comparing a DSG clock to one on earth [5].
Lorentz Violation (LV):
  Lorentz symmetry states that physics is the same regardless of relative orientation and speed of reference frames.
  Some theories of quantum gravity predict a violation of this symmetry.
  DSG orientation and speed constantly changing relative to Earth => DSG platform is ideal for looking for LV effects by comparing a DSG clock to one on earth.
  Clock comparisons include changes in orientation (Michelson-Morley), velocity (Kennedy-Thorndike), and time dilation (Ives-Stilwell) [5, 6].
Quantum Entanglement (QE):
  correlating quantum states in systems separated by a space-like interval (not causally linked in a classical picture.)
  Entanglement has been demonstrated out to 1200 km [8]. A DSG clock could extend this to earth-moon distance by using precise timing to improve signal to noise in an entangled state measurement apparatus.
Quantum Clock Synchronization (QCS):
  An algorithm to synchronize remote clocks using quantum entanglement.
  Not yet fully demonstrated on the ground

## Slide 41: Gateway Provides Precision Time & Frequency for Science
Test | Observable | Instrument | Current Best | DSAC on DSG
Space VLBI | 230+ GHz microwave signals from SMBH’s | Multiple antennae + clocks + … |  | 
1/r2 cosmological constant DM vs. MOND1[1] | S/C range | DSG clock + drag free S/C + DSG-Moon laser range2 | 10-6 at 100 AU | 5x10-8 at 100 AU2
Cosmological constant [2] | Planetary orbit from S/C range | DSG clock + drag free S/C + DSG-Moon laser range3 |  < 10-36 |  < 10-39
Cosmological constant | Perihelion precession4 | Ranging between DSG and the moon |  < 10-36 | Unlikely to be competitive, but needs more study9
Gravitational Red Shift | Clock compare: DSG to earth | DSG clock | Clock: 2.5x10-5 [3a] AI: 7x10-9 [4] | >3x10-6 [5] ~5e-6 (VERITAS)
Lorentz Violation | Clock compare DSG to earth in different reference frames | DSG clock | 1.1x10-8 [6] | >1x10-7 [5]
Local Position Invariance | Compare two isotopes in- situ in space with elliptical orbit | Dual 199Hg+/201Hg+ clock | TBD | TBD
Dark Matter | Clock compare | network of clocks7 | Energy scale > 105 TEV for a 100 km defect size [7] | >106 TEV [8]

## Slide 42: Gateway Provides Precision Time & Frequency for Science
Test | Observable | Instrument | Current Best | DSAC on DSG
Quantum Entanglement | Entangled photon polarization correlations [8] | DSG clock + high precision clock elsewhere + entangled photon source10 | Demonstrate entanglement over 1200 kilometers | Extend to Earth-moon distance scale
Quantum Clock Synchronization | Time synchronization of two remote clocks | 2 DSG clocks + 1 remote clock + set of entangled qbits + method for transferring qbit phase to clocks11 | Not yet fully demonstrated | 

## Slide 43: Gateway Provides Precision Time & Frequency for Science
Notes:
1DM = Dark Matter, MOND = Modified Newtonian Dynamics
2There is already a proposal to do 1-meter ranging with the DSN on a drag-free S/C at the outer planets (minimal solar gravitational potential). There is also a proposal to do laser ranging from the DSG to the moon at the mm level. If these two were already present, ranging from the gateway might be done at the sub-meter level, thereby improving on the DSN-based measurement.
3This is similar to the previous experiment and uses the same set up, but by virtue of the measured precision drag free satellite range, a planet position is inferred. The planetary orbit would place a limit on the cosmological constant.
4Use the DSG clock to do ranging from the DSG to the moon. The large eccentricity of the DSG orbit is predicted to result in a large perihelion precession as predicted by GR and could be used to place a limit on the cosmological constant. However, laser ranging should be far more precise than that using a clock.
5Might be better than existing clock comparisons, which are led by GPA, but not competitive with atom interferometer (AI) experiments.
6Willams et al., estimate that a Lorentz violation measurement could be improved by 2 orders of magnitude using the DSG highly eccentric orbit and an on board optical clock. A DSAC clock would provide 3 orders of magnitude worse stability, so the effect would be measured with 10x less precision than is already done.
7Cross correlation between two clocks is performed in situ or between remote clocks. For a given time scale, there should be no correlation (random noise).
8This is very speculative. The highly eccentric orbit may increase sensitivity to certain defect size scales not currently available.
9Current best measurements are based on perihelion advances of Earth and Mars. The perihelion advance of the DSG is likely to be less well known, but this needs more study.
10Quantum entanglement experiments send pairs of photons with their polarization states entangled, between two remote sites. Often the signals are greatly attenuated by the transmission distance. Precise timing on both ends enables very weak signal retrieval using narrow band detection.
11Quantum Clock Synchronization protocols synchronize remote clocks using measurements on quantum systems with shared prior entanglement. A quantum measurement on one element of a pair in one location causes a wave function collapse in the other remote partner and subsequent synchronized phase evolution. However virtually all such protocols depend on having a common phase origin in the two separated systems, which is equivalent to already synchronized clocks. More recent protocols attempt to address this loophole, however it is not clear how this could be implemented in an actual DSAC-like clock. The mechanism must be demonstrated in the laboratory using real clocks before attempting a demonstration on the DSG platform.

## Slide 44: Gateway Provides Precision Time & Frequency for Science
References:
[1] B. Buscaino, et al., “Testing long-distance modifications of gravity to 100 astronomical units,” PRD 92, 104048 (2015)
[2] M. Sereno, and P. Jetzer, “Solar and stellar system tests of the cosmological constant,” PRD 73, 063004 (2006).
[3] R. Vessot et al., PRL 45, 2081 (1980)
[3a] S. Herrmann, et al., PRL 121, 231102 (2018)
[4] H. Muller, A. Peters, and S. Chu, “Precision measurement of the gravitational red shift by the interference of matter waves,” Nature 463, 926 (2010).
[5] J.A. Williams and N. Yu, Deep Space Gateway Science workshop (2018)
[6] P. Delva et al., PRL 118, 221102 (2017)
[7] P Wcislo et al., “Experimental constraints on dark matter detection with optical clocks,” Nature Science 1, 0009 (2016).
[8] J. Yin, et al., “Satellite-based entanglement distribution over 1200 kilometers,” Science 356, 1140 (2017)

## Slide 45: Measuring the GR gravitational red shift using DSAC on VERITAS
Measuring the GRS: Engineering vs. science
GRS currently known at 1e-5 level
Use depends on which is more precise: knowledge of GRS OR knowledge of OD, potential, and clock
For instance:
  DSAC orbit varied by 2e4 m altitude for each 100 minute orbit => GRS = 2e-12
  In DSAC orbit, GRS could only be measured at 1e-3 level – engineering
BUT, in deep space, potential change can be much larger
Eg., in cruise phase to Venus, sample *solar* potential over 4e10 meters!
Earth-Venus GRS ~ 1e-9 (solar effect at Venus is ~1e-19/m)
  Venus cruise GRS could be measured at 1e-6 level - science
LPI/LLI slides courtesy of G. Cascioli, F. de Marchi, and L. Iess
F. De Marchi, et al., “Testing the Gravitational Redshift with an Inner Solar System Probe: The VERITAS Case.” Physical Review D 107, no. 6 (March 13, 2023): 064032. https://doi.org/10.1103/PhysRevD.107.064032.
Notes: NOTE: this is a direct measurement of the GRS – there are so-called “null red shift” measurements that place stricter limits but are technically not the same (eg., Ashby et al. PRL, 2007)
NOTE 2: Do not need an accurate clock, but do need to have full understanding of systematic effects – must show that systematics don’t vary on the same time scale as a possible LPI violation

## Slide 46: LLI – LPI on VERITAS: Iess et al. analysis
This parametrization can be simplified under certain assumptions:
If the two clocks are identical or based on the same kind of atomic transition:
LPI violation only (e.g. Delva et al. 2018):
Local Position Invariance (LPI)
Local Lorenz Invariance (LLI)
Space craft GRS
Station GRS
Station 2nd order Doppler
Space craft 2nd order Doppler
LPI/LLI slides courtesy of G. Cascioli, F. de Marchi, and L. Iess, University of Rome, VERITAS science team
Notes: Lucian Iess and team, on VERITAS science time, use standard LPI/LLI violation parameterization

## Slide 47: RO Timing Methods
USO:
  Traditional approach
  High TRL
  Large drift
2-way Ka-band microwave link to ground maser
  Emerging approach
  Maser-like stability
  May not always be possible
  Resynch overhead (“lock up time” = 2x light travel time to Earth)
DSAC
  Near-maser stability in-situ
  Vs. USO and other clocks: Lowest environmental sensitivity
  Vs. USO: No drift
  Vs. 2way: 2x lower resynch time (1-way vs. 2-way)
  Vs. 2way: Better localization (compared to 2-way)
  Vs. 2way: Reduction of solar plasma effects by 2x
Notes: Alternatives to USO
Two-way: lock up time: 2x light travel time to earth: lost data
DSAC: no drift, no re-lock, better localization, less solar plasma