## Slide 1: UTC(k) steered by intermittent operation of an optical clock
Tetsuya Ido
H. Hachisu, N. Nemitz, H. Ito, N. Ohtsubo, Y. Miyauchi, M. Morikawa, M. Tonnes, T. Gotoh, Y. Hanado, K. Matsubara
Space-Time Standards Laboratory,
National Institute of Information and Communications Technology (NICT)

## Slide 2: Outline
UTC(k) generation using an optical clock
UTC(k) : real signal generated in laboratory k
Optically steered timescale (1): Source oscillator + Frequency reference
Optically steered timescale (2): Source oscillator + MW clock ensemble + Frequency reference
  Optically steered UTC(NICT) for 2 years
Recent improvements
Option 2 in the redefinition of the SI second
Graphical picture of the option 2a and 2b for the redefinition of the SI second

## Slide 3: UTC(k) generation using an optical clock

## Slide 4: How is UTC maintained?
institute k
NICT
BIPM
UT1
leap second
UTC(NICT)
Cs (or Rb) fountains
UTC(k)
adjustment
Japan Standard Time (JST)
>400 clocks
H-Maser
Local Ensemble
18 clocks
+ 9h
PMS
adjustment
adjustment
UTC
TAI
International Atomic Time
EAL (ensemble)
Generating a real signal in institute k
Generating a paper clock “UTC”

## Slide 5: JST generation
Atomic clocks
18 Cs
3 HM
1PPS generation
JST
Time, frequency adjustment
Ｈメーザ
Cs clock
Source oscillator
Ｈメーザ
HM
Phase micro stepper
Synthesis
Paper clocks as ensemble of 18 clocks
3 to 1 multiplexer
+
-
HM frequency was steered to ensemble of 18 Cs clock

## Slide 6: DMTD system performance
Frequency difference
since 2012, over 8 HMs:
We expect persistent frequency errors < 3
Flywheel HM – UTC(NICT)
HM rf signal transfer
comb up-conversion
We similarly examine the performance of the Japan Standard Time system dual-mixer time-difference system.
Agreement of TIC and DMTD
Instability over averaging time τ (s)
Constant errors are bounded by comparing DMTD measurement against a time interval counter (TIC).
HM instability

## Slide 7: Reliability (Power & Signals)
Atomic
Clocks
AOG frequency steering
Opt one of signals
Phase difference measurement
between clocks
RTA
calculation
Data storage
& steering algorithm
5MHz
1PPS
Control
LAN
Status monitor
(Cs, HM, AOG, multiplexer, room temp. humidity)
Real-time monitoring
Measurement, steering, and choosing signals
Multi
-plexer
Multi
-plexer
Manual change of signal selection
3 signals for redundancy
UPS & Two backup power generators

## Slide 8: Hydrogen masers at NICT
NICT operates multiple HMs that has Hadamard deviations at the 10-16 level.
Mid 10-16 level of Hadamard deviation extends to 10 days or more.
Mote than half of HMs reach less than 2×10-16 at the bottom.

## Slide 9
10-12
10-13
10-14
10-15
10-16
100 101 102 103 104 105 106 107 108
Timescale generation
Japan Standard Time before 2020
One hydrogen maser is source oscillator for UTC(NICT)
Two more masers serve as backup.
An ensemble of 18 cesium 5071A clocks steers the timescale.
Behavior of UTC(NICT)
Source oscillator HM4
The deviation from UTC is typically less than 20 ns.
Instability is 2×10-15 over 10 to 30 days.
UTC(NICT) = HM4+AOG
Allan deviation
Averaging time [s]
20 ns
(3-corner hat)
UTC(NICT)-HM3
NICT Cs ensemble
timescale
HM3-HM4
JST system
backup generator
UTC – UTC(NICT) [ns]
Date
20 ns
5 years

## Slide 10: Optically steered timescale in 2016: Source oscillator + Frequency reference
6 months
1×10-13
104 s measurement
Source osc.
(HM)
Phase micro stepper (PMS)
f(HM)
based on Sr
TA(Sr)
TA(Sr): Frequency and drift rate of HM signal is adjusted with reference to Sr
Adjustment of PMS offset frequency every 4 hours
No servo to reduce the time offset UTC ─ TA(Sr)

## Slide 11: Steering
Interval for the drift estimation: T=25days
Number of Optical frequency standard (OFS) operation in T : N+1 >4
(once per week or more frequently)
One HM free evolution time: DT = T/N

## Slide 12
Comparison against UTC & TT(BIPM16)
the frequency offset of UTC is clearly identified.
Phase difference against TT(BIPM16) < 1ns after 5months
Stability 4e-16 @ 20 days
TA(Sr)-UTC
TA(Sr)-TT(BIPM16)
TT(BIPM16)-UTC
H. Hachisu et al., Sci. Rep. 8, 4243 (2018)
H. Hachisu

## Slide 13: Upgrade of UTC(NICT) in 2021
2006
JST system built in 2006 with rich redundancies for reliability.
  Three independent HMs signals are steered to ensemble clocks made of 18 Cs
  3 DMTD & 1 TIC for reliable measurement of clock phase differences
2016
Test operation of a Sr-steered timescale (1HM + Sr) showed great potential in accuracy.
2020
Reliability of Opt.-> MW conversion improved. Monitor, control, and evaluation have become systematic. Redundancy in various parts established.
→ Reliable operation of a Sr lattice clock once per week now feasible.
2021 ─ UTC(NICT) is finally steered by NICT-Sr1
Y. Hanado
N. Nemitz
H. Hachisu
K. Imamura
F. Nakagawa

## Slide 14: Strategy
To keep the benefit of reliability in legacy system…
  Keep the steering of 3 HM signals to Cs ensemble
  →Occasional adjustment of the parameter of frequency offset with respect to Cs ensemble timescale
  To keep the intermittent operation of the strontium lattice clock (once per week)
    →Leaves four (or more) working days for any necessary. Reduced operation time to extend the durability
Operations once per week or more has continued since September 2021, and is still ongoing.
Measurement Interval
mostly ≤ 1 week for 2 years
120 operations

## Slide 15: Recipe of steering
Frequency adjustment of UTC(NICT) twice per week to remove y(UTC(NICT)-TA(Sr)) in the one past day
The instability of Cs ensemble often requires >3e-15 adjustment
Time difference DT=UTC(NICT)-TA(Sr) gradually removed by additional fractional frequency adjustment of
DT/(60 days)

## Slide 16: Steering using optical clocks: before & after
2 years
Standard deviation: 2.3ns
Standard deviation: 8.5ns
Incorporation of Sr reduced the fluctuation to 1/4.
The magnitude of the reduced fluctuation reflects the stability of mean timescale composed of commercial Cs clocks
(Poster No. 110)
Measurement Interval
mostly ≤ 1 week for 2 years

## Slide 17: Incorporation of HMs to TA
HMs have been used only as source oscillators. But predictions of the frequency drift in HMs allow us to incorporate HMs to the ensemble clocks.
BIPM did similar improvement in early 2010s*. EAL incorporate the 2nd-order term (frequency drift) of atomic clocks
  → Cs clocks have lost the weight, but HM gained the weight
  Similar strategy possible by adding HMs to ensemble clocks as well as changing the dedrifting process in mean free time scale.
H. Ito
Note that NICT prefers to depend on our own clocks as much as possible for resilience against the GNSS non-availability
*G. Panfilo, et al., Metrologia 49, 49 (2012)
G. Panfilo et al., Metrologia 51, 285 (2014)

## Slide 18: Actual calculation (an example)
Suited method and value for the purpose are important.
y (rate) and w (weight) define ’s character
Clock rate
:Measured :Predicted :Output
:Calculated from xi
…(6)
…(8)
…(9)
From Y. Hanado, IFCS-EFTF 2023 tutorial (2023)
Detrend term
Linear prediction
based on the past duration T
with respect to TA itself
Could be self divergent

## Slide 19: Actual calculation (an example)
Variation of de-trending
The optimal way to detrend depends on the clock’s behavior.
Cs clocks
H-maser
(A)
(B)
(C)
…(8)
detrend term
No trend
Linear fit
Quadratic fit
From Y. Hanado, IFCS-EFTF 2023 tutorial (2023)

## Slide 20: Consideration for independent timescale UTC(k)
Reference for detrending
  BIPM uses not EAL itself, but TT(BIPM) as the prediction reference
  Local source is ideal for UTC(k), considering the resiliency
  Independent signal, TA(Sr)? Or the ensemble timescale itself?
  Weighting
    Allan deviation → Hadamard deviation
 | BIPM (-2011) | BIPM (2011-) | UTC(NICT) (-2022) | UTC(NICT) (2023-)
Detrend (in phase) | linear | quadratic | linear | quadratic
Reference for prediction | EAL | TT | RTA | RTA@NICT
weight | Allan variance | predictability | Allan deviation | Hadamard deviation

## Slide 21
Similar level of stability with TA(Sr) up to one month
→ Steering using Circular T becomes an option
We need to provide (or adjust) three parameters.
Phase offset, frequency offset, frequency drift offset
Algorithm to incorporate Sr data will be explored.

## Slide 22: Improvement after quadratic detrending
4 or 5 stable HM has weight in similar level.
(with Sr steering)

## Slide 23: Next step & outlook
Optimization of TA(HM)
  Reference for dedrifting
  weight
Adjusting three parameters (phase, frequency, frequency drift) based on intermittent operation of the optical clock and Circular T for phase alignment
  Kalman filter which can manage both frequency and phase data?
  Second or possibly commercial lattice clocks
  Redundancy for public service could be achieved by commercial system (For instance, Poster No. 82 by Shimadzu)
(Poster No. 112)

## Slide 24: Graphical picture of the option 1, 2a and 2b in the redefinition of the SI second

## Slide 25
Choosing multiple transitions as primary transitions?
fCs /Hz
nSr /Hz
Constraint by LSA
(least square analysis)
9 192 631 770
Constraint
by definition ( = definition line)
,
uSr
Current situation
All measurements and least square analysis tell us that point falls on the red line with uncertainty.
“Definition” determines where the point is on the red line.
Currently, is fixed to be 9192631770 Hz.
We find a intersection with the redline and vertical blue dot line.
The definition is an ARTIFICIAL constraint to determine values.
This constraint does not need to be vertical line, as long as it provides intersection with the red (LSA) line.

## Slide 26
Option 1:
Definition LINE : n1=50
→Uncertainty only occurs in n2, n3, …
(uncertainty of the ratio 0.1)
Another possibility: using arithmetic mean of two transitions
Option 1 (n1=50)
=0.9(1)
（note that →Definition1))
（1）
If n1 and n2 has same uncertainty, it is fair to set the definition line as
Weight & nominal frequencies are mixed, and we cannot separate them…
Definition LINE:
Further simplified picture

## Slide 27
Indeed, curves that “Geometric mean = constant” is curve
Revision process
New experiments shift black line and confidence band.
Option 2.a: the definition curve is not renewed.
(point stays on the original curve)
Option 2.b: the definition curve is renewed after (possibly choosing transitions), fix the intersection point, and finally setting weights
This formula holds weights and nominal frequencies separately !
Definition line can also be a curve.
If weight is inversely proportional to the square of uncertainty, the projection of the confidence band to each axis determines the uncertainty of each transition frequency.
k=0.5
k=0.6

## Slide 28: Summary
Intermittent operation of an optical clock (once in a week) realized an optically steered timescale for two years.
Ensemble of HMs with quadratic de-drifting has good stability. We need a recipe to determine the offset parematers in phase, frequency, and frequency drift.
Option 2 in the redefinition of the SI second does not bring you to a totally different regime. Option 1 could have been just a special case of option 2.

## Slide 29: Teams
Members of the space-time standards laboratory, including those for telecom-related T&F research, public service, and technical & administrative support

## Slide 30
N. Nemitz
H. Hachisu
H. Ito
M. Tonnes
K. Matsubara
N. Ohtsubo
Y. Miyauchi
M. Morikawa
Lattice clock
Generation of Japan Standard Time
In the audience
T&F comparison
T. Gotoh
Y. Kozuki
R. Ichikawa