## Slide 1
Optical and microwave metrology with frequency combs
Tara Fortier
NIST Time and Frequency Division
Oct 10, 2023

## Slide 2
Introduction to atomic clocks.
What is a frequency comb and how does it work?
Combs for optical and microwave metrology.
Time & frequency applications with combs.
Outline

## Slide 3
2. The second is then defined as a fixed number of cycles of radiation of the atomic transition.
The second is currently defined as 9,192,631,770 oscillations of an electronic hyperfine transition in 133Cs
To calibrate the passage of time accurately you need to accumulate cycles of a phase continuous signal since t0 = 0s.…. Due to their complexity most atomic clocks are only run intermittently.
Atomic clocks are frequency references, not time references!
1. Cycles of microwave EM radiation are stabilized to a hyperfine transition in 133Cs
SI second
How the atomic clock works
How atomic clocks define time

## Slide 4
133Cs
Voltage
Time
1 s
Unknown frequency in
Frequency is determined by counting # zero crossings in 1s
2. The second is then defined as a fixed number of cycles of radiation of the atomic transition.
The second is currently defined as 9,192,631,770 oscillations of an electronic hyperfine transition in 133Cs
1. Cycles of microwave EM radiation are stabilized to a hyperfine transition in 133Cs
SI second
How the atomic clock works
How atomic clocks define time

## Slide 5
The microwave timescale
The steered u-wave ensemble signal is used by the NMI to provide stable local signals for synchronization of oscillators in mobile devices, GPS, etc.
133Cs corrects the frequency drift of a weighted ensemble of microwave oscillators (usually H-masers and commercial atomic clocks) that act as a continuously running flywheel for local time.
The microwave flywheel creates continuous timing “ticks”
133 Cs
Frequency steering
The local Cs clock, however, can only provide frequency steering of the maser ensemble…. 133Cs cannot tell if the ensemble is “on time”
How atomic clocks define time

## Slide 6
PSFS
PSFS
400 atomic clocks
50 NMIs
How atomic clocks define time
NIST’s microwave timescale information is broadcast via satellite to the BIPM to compare against the international ensemble of primary and secondary standards (PSFS)
The BIPM sends a document to NIST, once a month, to tell it how many nanoseconds off UTC(NIST) is with respect to PSFS.
NIST uses the information from BIPM to steer UTC(NIST) on time with PSFS.
Fun fact: the NIST time server receive 10 time more combined hits per day than the Google Search engine!

## Slide 7
Fractional Accuracy
Year
Al+
Yb
Sr
Al+
Sr
Historical performance of atomic clocks
Microwave clocks
Optical clocks
N. Poli et al, Rivista del Nuovo Cimento, 36 (2013)
Ludlow Rev. Mod. Phys. 87 (2015)
Clocks accuracy Dn/n
… operating at a higher transition frequency naturally increases the precision and accuracy.
1 part/million control of u-wave a cycle!
Accuracy Limit 10-16
1 part/1000 control of an optical cycle
Accuracy limit
10-18
Notes: Steeper slope for optical clocks
Reason: microwave clocks require much better control over their systematics.
Accuracy and potential of optical clocks…. Redefinition of second.
If we look at the performance of clocks over time we see a steeper slope associated with improvements in the accuracy of optical as opposed to microwave clocks. There is a reason for this. As I mentioned before, the higher the frequency of your clock the better the accuracy and stability. Said differently, the lower your clock frequency, the higher the constraint on the your systematics.
To achieve parts in 10-18 requires knowledge of an optical transition to 1 mHz. On a microwave transition 10-18 requires that you know your transition frequency to a fraction of a micro-Hz. Because the difference in constraints, microwave clocks are not likely to make significant gains in accuracy over their current performance, but optical clocks which have been demonstrated to be 100 more accurate are still improving.
The question that must be on everyone’s minds is…. Given the higher accuracy and potential for optical clocks…. Is it the time right for optical redefinition of the SI second?

## Slide 8
Up to 10 GHz
3 GHz processor
200 Tera-oscillations/s
~10,000 to 100,000 x faster
than a state-of the-art
counter!
The big challenge of optical clocks
The big obstacle to optical clock development: light waves are too fast to count!

## Slide 9
V
t
Photodector signal
Wave 1
(nopt1)
PD will produce a voltage proportional to the light’s intensity
The carrier frequency of a light wave is too fast for a photodiode to resolve….
Difference frequency measurement

## Slide 10
+
Wave 2
(nopt2)
t
Photodector signal
V
Beat frequency
Df =(n2 - n1)
=
Interference of two optical waves close in frequency results in an amplitude modulated optical carrier
Wave 1
(nopt1)
+
T. Udem et al., "Phase-coherent measurement of the hydrogen 1S-2S frequency with an optical frequency interval divider chain," in IEEE Transactions on Instrumentation and Measurement, 46 (1997)
Difference frequency measurement

## Slide 11
A brief history of optical frequency measurement
Atomic transition
~ 300 THz
5 orders of magnitude
different in frequency!
Problem!
Counter
Conversely, to define the second to Optical atomic time we need to count the cycles of the clock transition.
Cs clock
9.192 GHz
To measure the frequency of the optical transition we need to measure it against the definition of the Hz.

## Slide 12
Atomic transition
~ 300 THz
Counter
Cs clock
9.192 GHz
Frequency chain enables multiplication of the 133Cs microwave reference to the optical domain via a series of 20 intermediate and 50 phase locked oscillators.
Schnatz, et al., Phys. Rev. Lett. 76, 19 (1996)
Schnatz, et al., Phys. Rev. Lett. 76, 19 (1996)
Schnatz, et al., Phys. Rev. Lett. 76, 19 (1996)
Schnatz, et al., Phys. Rev. Lett. 76, 19 (1996)
Frequency multiplication chains: prior to 2000
A brief history of optical frequency measurement

## Slide 13
Atomic transition
~ 300 THz
Counter
Cs clock
9.192 GHz
In 2000, the first optical frequency comb was demonstrated:
direct optical-to-microwave conversion with a single oscillator!
OFC
2005 Nobel Prize: Jan Hall and Ted Haensch
A brief history of optical frequency measurement

## Slide 14
Optical IN
250-500 THz
Optical OUT
150 – 280 THz
Photonic Synthesizer
Reference
DT < 1fs
Microwave OUT
DC – 100 GHz
optical
m-wave
The frequency comb: a photonic pulsed synthesizer
Ultrabroad frequency synthesis:
Wideband optical & microwave synthesis
Direct optical-to-microwave conversion
Near perfect frequency multipliers/dividers
  High precision frequency counter
Ultra-short optical pulse synthesis
fs-level pulse-to-pulse timing jitter
Control/shaping of pulse electric field
Arbitrary waveform synthesis

## Slide 15
What is a frequency comb and how does it work?

## Slide 16
L – cavity length
2 fr
fr
3 fr
4 fr
Spectrum of lasing light
n
Artist rendition -oscillations should be 105 times faster
Modelocked lasers: Time/frequency representations
resonant modes are equally spaced by the laser repetition rate:
nn = n (c/ 2L)= n fr

## Slide 17
Spectrum of lasing light
n
Coherent addition of longitudinal cavity modes
Pulse formation
Modelocked lasers: Time/frequency representations
resonant modes are equally spaced by the laser repetition rate:
nn = n (c/ 2L)= n fr

## Slide 18
Spectrum of lasing light
n
Modelocked lasers: Time/frequency representations
resonant modes are equally spaced by the laser repetition rate:
nn = n (c/ 2L)= n fr

## Slide 19
Spectrum of lasing light
n
Modelocked lasers: Time/frequency representations
resonant modes are equally spaced by the laser repetition rate:
nn = n (c/ 2L)= n frep

## Slide 20
Spectrum of lasing light
n
Modelocked lasers: Time/frequency representations
resonant modes are equally spaced by the laser repetition rate:
nn = n (c/ 2L)= n fr

## Slide 21
Modelocked laser-based frequency comb
Optical carrier
Pulse envelope
Tr

## Slide 22
Modelocked laser-based frequency comb
Optical carrier
Pulse envelope
Tr

## Slide 23
Modelocked laser-based frequency comb
Optical carrier
Pulse envelope
Tr

## Slide 24
Modelocked laser-based frequency comb
Tr
2 DfCE
DfCE
0
3 DfCE
4 DfCE
The CE phase slip results from a difference in dispersion: the group and phase velocities are different as the pulse travels through a dispersive medium, such as the laser crystal or optical fiber.
Optical carrier
Pulse envelope
2023 Nobel Prize: Ultrafast studies of electron dynamics in atomic and molecular systems.

## Slide 25
Modelocked laser-based frequency comb
frequency
fr
Microwave
Optical
fCE < fr
2 DfCE
DfCE
0
3 DfCE
4 DfCE
Tr
L
DfCE = 2p (fCE / fr )
Tr = 1 / fr = 2L/c

## Slide 26
Modelocked laser-based frequency comb
frequency
fr
Microwave
Optical
fCE < fr
2 DfCE
DfCE
0
3 DfCE
4 DfCE
Tr
L
nN= N x fr+ fCE
The comb equation
DfCE = 2p (fCE / fr )
Tr = 1 / fr = 2L/c

## Slide 27
The comb equation:
Directly connects optical and microwave frequencies.
Optical modes are related in a deterministic fashion (coherently related spectral components, or every optical mode shares the same phase)
fr is inversely to the pulse round trip time:
cavity length (temp changes)
Control: PZT actuated mirror
fCE connects the optical spectrum to DC (0 Hz).
It is dependent on dispersion in the laser cavity
air currents
material dispersion
changes in gain
Control: modulation in pump power (gain modulation)
nN= N x fr + fCE
Optical frequency (500 THz)
The comb equation

## Slide 28
nN= N x fr + fCE
Optical frequency (500 THz)
N takes fr to the optical domain (can be determined using a wavemeter Dn/n ~ 10-7)
Tr =1/fr= 2L/c
Tr =1/fr
The comb equation: fr measurement

## Slide 29
nN= N x fr + fCE
Optical frequency (500 THz)
Tr =1/fr= 2L/c
fr
Frequency counter
The comb equation: fr measurement
N takes fr to the optical domain (can be determined using a wavemeter Dn/n ~ 10-7)

## Slide 30
Hard to measure!
Is related to the phase of the optical carrier)
If you can’t measure the optical carrier, so how do you measure its phase?
Without access to fCE……. Dnatom/natom = 1 part per million = 109/1015 !
2 DfCE
DfCE
0
3 DfCE
4 DfCE
Tr
DfCE /(2p Tr) = fCE
nN= N x fr + fCE
Optical frequency (500 THz)
The comb equation: fCE measurement

## Slide 31
Optical frequency combs
n2n = 2n frep + fCE
nn = n frep + fCE
x 2
fCE
_
n2n = 2n frep + fCE
2nn = 2n frep + 2fCE
fCE
Measuring f0 requires an optical octave of bandwidth!... Poses significant challenges to systems with low pulse energies (high fr systems).
J.Bethe (2009)
Continuum generation
High confinement
Small mode volume = high intensities
Long interaction lengths
dispersion is managed by balancing material and geometrical dispersions.
Carlson et al. Optics Letters Vol. 42, 2314 (2017)
Cores =2 - 5 mm
The comb equation: fCE measurement

## Slide 32
fCE
When detecting f0, the photodetector detects a pulse train with slow additional amplitude modulation, with modulation frequency, f0
fr
2 fr
3 fr
4 fr
f
RF spectrum
fCE conjugate beats
Power (dBm)
MLL
+
Nonlinear fiber
x2
n2n, 2nn
bandpass
Optical amp.
nn
n2n
Optical setup: fCE measurement

## Slide 33
nN= N x fr + fCE
The comb equation
Important note: The comb equation holds for any coherently modelocked laser spectrum!
Question: What is a frequency comb?
Answer: A modelocked optical spectrum that is
Deterministic (coherent) and
Accessible or controllable in frequency and phase.

## Slide 34
Modelocked sources and clocks vs wavelength
ErYb:glass
Er:fiber
Cr:LiSaF
Yb:KYW
Yb:fiber
Tu:fiber
Cr+:ZnSe
Cr+:ZnS
Ti:sapphire
Wavelength
1550 nm
1000 nm
2 mm
2-3 mm
200 nm
100 nm
700 - 400 nm
VISIBLE
Any modelocked laser can be turned into a frequency comb as long as you can measure fCE:
usually requires optical amplification + non-linear broadening
c3
Gain
Pump
High-reflector
Partial reflector
SA
n(I) = n0+ I n2
Preferential gain for pulsed operation

## Slide 35
Modelocked sources and clocks vs wavelength
https://www.bipm.org/en/publications/mises-en-pratique/standard-frequencies
171Yb
(578 nm)
27Al+ (267 nm)
115In+ (237 nm)
199Hg+ (282 nm)
87Sr (698 nm)
88Sr+ (674 nm)
40Ca (657 nm)
171Yb+ (436nm/ 467 nm)
288Th nuclear
(~149 nm)
ErYb:glass
Er:fiber
Cr:LiSaF
Yb:KYW
Yb:fiber
Tu:fiber
Cr+:ZnSe
Cr+:ZnS
Ti:sapphire
Wavelength
1550 nm
1000 nm
2 mm
2-3 mm
200 nm
100 nm
700 - 400 nm
VISIBLE

## Slide 36
Micro-resonators:
Compact nonlinear optical filters/cavities
CMOS compatible
Chip-scale
  Inexpensive to manufacture
Credit: Image by Jiangang Zhu and Jingyang Gan/WUSTL
200 um
Silica
CaF2
SiN
Vahalla et al
Gaeta et al
Maleki et al
Smaller: Optical frequency combs
Chip-based combs: Kerr-resonators
Pump
c3
Pump
High-Q resonator
SA
n(I) = n0+ I n2
Comb modes formed via degenerate and non-degenerate 4 wave mixing

## Slide 37
Micro-resonators:
Compact nonlinear optical filters/cavities
CMOS compatible
Chip-scale
  Inexpensive to manufacture
Credit: Image by Jiangang Zhu and Jingyang Gan/WUSTL
200 um
Silica
CaF2
SiN
Vahalla et al
Gaeta et al
Maleki et al
Smaller: Optical frequency combs
Current limitations:
Low optical coupling < 5%
Narrow optical bandwidth ~ 10 nm
Optical loss due to Brownian noise limits stability
Difficult to broaden-> low energy/pulse
Chip-based combs: Kerr-resonators
Pump
c3
Pump
High-Q resonator
SA
n(I) = n0+ I n2
Comb modes formed via degenerate and non-degenerate 4 wave mixing

## Slide 38
Electro-optic comb generators
EDFA
Tunable
CW laser (fc)
fc
frequency
Optical spectrum IN
Optical spectrum OUT
fc
fsynth
frequency
A. J. Metcalf, V. T-. Campany, D. E. Leaird, and A. M. Weiner, JSTQE 19(6) (2013).
PM
+
IM
Comb output
PM
PM
+
+
+
fsynth (< 40 GHz)

## Slide 39
3. Combs for optical and microwave metrology

## Slide 40
Combs for measuring optical frequencies
< 1 GHz
nN= N x fr+ fCE
Df
frequency
fr
Microwave
Optical
fCE
natom (Hz)= N x fr +fCE+ Df
Wavemeter
~ 500,000

## Slide 41
natom (Hz)= N x fr +fCE+ Df
133Cs
133Cs
Frequency counter
Combs for measuring optical frequencies
133Cs
Df
frequency
fr
Microwave
Optical
fCE
nN= N x fr+ fCE
Wavemeter
~ 500,000

## Slide 42
1/Df
Accuracy in microwave frequency counting
Df * dnCs /nCs
Df
=
1 x 10-16
Counter accuracy is only as good as its reference (”Hz”)
dnCs/nCs ~ 10-16
Normalization frequency
Counted frequency

## Slide 43
1/Df
Counter accuracy is only as good as its reference (”Hz”)
dnCs/nCs ~ 10-16
1/Df
Df * dnCs/nCs
=
Normalization frequency
100 MHz * 10-16
=
10-7 x 10-16
1 x 10-23 !
=
nopt
1 PHz
Because of the large normalization frequency, optical difference frequencies can be counted with extremely high resolution (10-19 t-1) and accuracy using a microwave reference and standard frequency counter.
Accuracy in optical difference frequency counting

## Slide 44
+
Nonlinear fiber
x2
n2n, 2nn
bandpass
Optical amp.
nn
n2n
Mode-locked laser
Dichroic beamsplitter
fCE
fr
Optical beat signal measurement requires the following:
Spectral overlap
Spatial overlap
Polarization overlap
With 1 mW of CW light + 100 nW in a single comb mode Df ~ 40 dB SNR
Combs for measuring optical frequencies
Optical band pass filter:
lopt ± 5 nm ~ 800 GHz @ 1550 nm
l/2
Pol.
Df
> 1 mW
Unknown optical frequency, lopt

## Slide 45
frequency
Stabilizing combs with optical clocks
Df
L
fCE
By stabilizing the comb at frequencies nN and at fCE, you are pinning the spectrum at two extremes: 1) in the optical domain and 2) near 0 Hz. This rigidly stabilizes the entire optical spectrum, transferring the fractional clock stability to all modes.
dTr/Tr < 10-15
fr ∝ nopt/N
dfr/fr < 10-15
Fortier et al. “Generation of ultrastable microwaves via optical frequency division,” Nat. Photon. 5 (2011)
nN= N x fr+ fCE
Pump
Pump current

## Slide 46
Summary: Frequency combs
Optical frequency combs (OFCs)
Ultra-broad spectral coverage (optical octave)
High-precision optical frequency counters (perfect frequency multipliers/dividers)
Provide fs-level pulse-to-pulse timing
Sub-radian phase control of the carrier-envelope phase for deterministic evolution/control of field sensitive experiments
MID IR/ Near IR/optical
Optical reference
UV/ XUV
Harmonic generation
Transfer phase and frequency information across the optical domain and down to m-wave frequencies.
m-wave
Photomixing
DFG
THz /Mid IR
Review: T.M. Fortier & E. Baumann, “20 years of developments in optical frequency comb technology and applications,” Communications Physics 2 (2019)

## Slide 47
4. Time & frequency applications with combs

## Slide 48
Fortier et al Nat. Photonics (2011)
Fortier et al Laser and Photonic Reviews (2016)
Low noise microwaves via optical division
fr
CW laser
Optical cavity
Length reference
DL/L = 10-15
Frequency reference
Df/f = 10-15
Timing stability
Dt/t = 10-15
fr ~
Sm-wave(f) = Sf_opt(f)/N2
1000 x narrower linewidth than a H-maser

## Slide 49
Df
nYb (local time)= N x fr +Df
Microwave timescale calibration
nYb(SI) = 518,295,836,590,863.6 Hz
A constant of nature
fr
H-maser (local time), tells the counter what it thinks 1-s is for gating the counters.
From the counters, we calculate:
“Local time” is calibrated to Yb atomic time by assessing the difference:
nYb (local time) - nYb(SI)
Time base
Dfr/fr < 10-15

## Slide 50
2-way time transfer: Clock synchronization/comparison
remote clock synchronization
relative measurement of remote clocks
Turbulence is common to both pulse trains.
Clock timing difference = interference between the red and blue pulse trains.
Enables free-space comparison of clocks in the absence of fiber networks.
Requires line of sight and good weather
F. Giorgetta Nature Photonics (2013), I. Khadar Optica (2018)

## Slide 51
87Sr
359 THz
2 x10-18
171Yb
518 THz
2 x10-18
27Al+
1.1 PHz
9 x10-19
Comparing atomic clocks
Modes act like tick marks to measure frequency differences.
Comparing clocks:
To perform measurements beyond the accuracy defined by the Hz (10-16)
Regular clock comparisons, local and international, are necessary for redefinition of the SI second.
Tests of fundamental physics.
Optical mode spacing can be split by parts in 1014!
resolution, s(t) ~ 10-18/t-1/2
100’s THz
nYb
nAl
Notes: And so, with our photonic synthesizer, we can span all of our favorite clock transition frequencies. And, we all say optical frequency comb but a better name might be optical frequency ruler because, much the same as a ruler has ticks to mark distance, the comb has lines to mark frequency, and these ticks span the entire spectrum and allow us to measure the difference in frequency between two or more optical clocks but with readily available microwave electronics. And much the same as the tick marks of a ruler can be split to measure millimeters, the spacing between modes in a frequency comb can be split by 10^14 to measure as small as micro-Hz!

## Slide 52
BRAN fiber link
Boulder Atomic Clock Optical Network
Yb lab
laser
Al+ lab
laser
Er comb
TiS comb
JILA
Sr lab
laser
Er comb
>5 frequency combs
1) Optical frequency measurement/comparisons
2) Frequency conversion
  Optical synthesis
  Optical-to-microwave conversion
3) Time/frequency distribution over free-space
NIST
Free-space link
BACON “Frequency ratio measurements at 18-digit accuracy using an optical clock network,” Nature 591 (2021).

## Slide 53
Comb applications across the E-M spectrum
Optical & nuclear
atomic clocks & spectroscopy
Optical free-space communications
Coherent Lidar
AOWFG
Ranging
Astronomical
spectrographs
Molecular spectroscopy
Microwave
Atomic Clocks
Low-noise
Microwave generation
Frequency
Attosecond Science
Photo-electronic spec.
Review: T.M. Fortier & E. Baumann, “20 years of developments in optical frequency comb technology and applications,” Communications Physics 2 (2019)

## Slide 54
Summary
1) OFCS enable down-conversion of optical frequencies to the electronic domain where signals can be counted, conditioned and synthesized for definition of the SI second, test & measurement, calibration services and GPS.
Microwave signals derived from optical reference are ultra-low noise and carry the stability and accuracy of the reference.
2) OFCs enable synthesis of harmonically related optical modes/optical pulses that carry the stability and accuracy of the reference.
The optical spectrum can be used as a high-precision frequency ruler for optical measurement of unknown signals and for comparison of optical clocks.
The stable optical pulse train can be used for free-space time transfer.
Optical IN
250-500 THz
Optical OUT
150 – 280 THz
Photonic Synthesizer
Reference
DT < 1fs
Microwave OUT
DC – 10 GHz
optical
m-wave
OFCs are mode-locked optical systems with a deterministic (coherent) relationship between optical and microwave modes.
Optical harmonics: nN = N fr+ fCE
Microwave harmonics: fN = N fr

## Slide 55
Thank you for your attention!
Have additional questions, please feel free to email me:
Tara.Fortier@nist.gov
T.M. Fortier & E. Baumann, “20 years of developments in optical frequency comb technology and applications,” Communications Physics 2 (2019)
N. Newbury, “Searching for applications with a fine-tooth comb,” Science 5 (2011)
S.A Diddams, K. Vahal & T. Udem “Optical frequency combs: Coherently uniting the electromagnetic spectrum,” Science (2020)