## Slide 1: Optical clocks for international timekeeping
Helen Margolis
Summer School on Frequency Standards, Precision and Quantum Measurement
Gold Coast, Queensland, Australia, 11th October 2023

## Slide 2: Timekeeping based on Cs atoms
The second, symbol s, is the SI unit of time. It is defined by taking the fixed numerical value of the caesium frequency Cs, the unperturbed ground-state hyperfine transition frequency of the caesium 133 atom, to be 9 192 631 770 when expressed in the unit Hz, which is equal to s1.
NPL-CsF2 , accuracy 2.0 1016
Definition is realised using a Cs fountain primary frequency standard
Used to contribute to the global time scales TAI and UTC
Used to steer the physical realisations of UTC maintained by national timing laboratories, e.g. UTC(NPL)

## Slide 3: Coordinated Universal Time (UTC)
BIPM
 80 National
Timing Institutes
Local UTC(k)
time scales
~ 500 commercial
atomic clocks
Free Atomic Time
(EAL)
leap seconds
Earth rotation
measurements
Universal Time
(UT1)
IERS
Coordinated Universal Time
(UTC)
International Atomic Time (TAI)
~ 12 Cs primary
standards
BIPM Circular T
UTC is a paper time scale, computed monthly

## Slide 4: UTC(k) time scales
Satellite links to other UTC(k) laboratories
Cs beam clocks
Hydrogen masers
UTC(NPL)
Time and frequency dissemination services
BIPM Circular T
Cs fountain primary frequency standard, NPL-CsF2
Accuracy 2.0 1016

## Slide 5: Optical atomic clocks
Theoretically achievable fractional frequency instability  Q 1
Q =
f0
f
= line quality factor
 | Microwave | Optical
f0 | ~ 1010 Hz | ~ 1015 Hz
f | ~ 1 Hz | ~ 1 Hz
5 orders of magnitude improvement in stability (in principle)
Ultrastable laser
Atomic reference (narrow optical transition in an ion or atom)
Femtosecond optical frequency comb
Frequency control

## Slide 6: Optical atomic clock technology
l
1D lattice, site spacing /2
Neutral atom optical lattice clocks
1S0 – 3P0 clock transitions (mHz natural linewidth)
Atoms confined in an optical lattice
N atoms, stability  N1/2
AC Stark shift eliminated by operating at “magic” wavelength
1P1
1S0
3P0
3P1
3P2
1st-stage cooling
2nd-stage cooling
clock transition
Trapped ion optical clocks
Based on “forbidden” (e.g. E2) transitions in single trapped ions (Q ~ 1015)
Long interrogation times possible
Rf ion trap provides a low perturbation environment
cooling transition
reference (“clock”) transition
 10 ns
 1 s
ground state

## Slide 7: Candidates for the atomic reference
H
Li
Na
K
Rb
Cs
Fr
Be
Mg
Ca
Sr
Ba
Ra
Sc
Y
La
Ac
Ti
Zr
Hf
Rf
V
Nb
Ta
Db
Cr
Mo
W
Sg
Mn
Tc
Re
Bh
Fe
Ru
Os
Hs
Co
Rh
Ir
Mt
Ni
Pd
Pt
Ds
Cu
Ag
Au
Zn
Cd
Hg
B
Al
Ga
In
Tl
C
Si
Ge
Sn
Pb
N
P
As
Sb
Bi
O
S
Se
Te
Po
F
Cl
Br
I
At
He
Ne
Ar
Kr
Xe
Rn
Ce
Th
Pr
Pa
Nd
U
Pm
Np
Sm
Pu
Eu
Am
Gd
Cm
Tb
Bk
Dy
Cf
Ho
Es
Er
Fm
Tm
Md
Yb
No
Lu
Lr
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
72
73
74
75
76
77
78
79
80
81
82
83
84
85
86
87
88
89
104
105
106
107
108
109
110
58
59
60
61
62
63
64
65
66
67
68
69
70
71
90
91
92
93
94
95
96
97
98
99
100
101
102
103
Neutral atom and single ion
Neutral atom
Single ion
Desirable features:
Narrow optical transitions (for low instability)
Insensitive to external perturbations (for high accuracy)
Accessible clock transition wavelengths (practical to build)

## Slide 8: Improvements in clock performance
Optical atomic clocks: absolute frequency measurements
Optical atomic clocks: estimated systematic uncertainties
Cs microwave atomic clocks
Optical redefinition of the SI second is anticipated

## Slide 9: Secondary representations of the second
Frequency standards that can be used to realise the SI second, and to contribute to TAI
Uncertainty cannot be better than Cs primary standards
List contains 1 microwave transition (Rb) and 10 optical transitions
Atom or ion | Transition | Wavelength | Recommended fractional uncertainty (2017)
199Hg | 1S0 – 3P0 | 266 nm | 2.4 x 10–16
27Al+ | 1S0 – 3P0 | 267 nm | 1.9 x 10–16
199Hg+ | 2S1/2 – 2D5/2 | 282 nm | 2.2 x 10–16
40Ca+ | 2S1/2 – 2D5/2 | 411 nm | 1.8 x 10–15
171Yb+ | 2S1/2 – 2D3/2 | 436 nm | 2.0 x 10–16
171Yb+ | 2S1/2 – 2F7/2 | 467 nm | 1.9 x 10–16
171Yb | 1S0 – 3P0 | 578 nm | 1.9 x 10–16
88Sr+ | 2S1/2 – 2D5/2 | 674 nm | 1.3 x 10–15
87Sr | 1S0 – 3P0 | 698 nm | 1.9 x 10–16
88Sr | 1S0 – 3P0 | 698 nm | 2.0 x 10–16
https://www.bipm.org/en/publications/mises-en-pratique/standard-frequencies

## Slide 10: Integration of optical clocks into UTC and UTC(k)
Secondary representations of the second can be used for:
Steering International Atomic Time (TAI)
Steering hydrogen masers in UTC(k) time scales

## Slide 11: Roadmap towards the redefinition
2020:
CCTF task group formed
CGPM 2022:
Validated roadmap
Different options for redefinition
CGPM 2026:
Redefinition option proposed
Mandatory criteria achievable by 2029
CGPM 2030:
Redefinition
CGPM 2034: (backup scenario)
Dedicated CCTF task force established June 2020, chaired by Noel Dimarcq and Patrizia Tavella
3 sub-groups bringing together 40 worldwide contributors to prepare a detailed roadmap towards the redefinition of the second
Paper submitted to Metrologia, available at https://arxiv.org/abs/2307.14141

## Slide 12: Mandatory criteria and ancillary conditions
Status should be advanced, even if not completely achieved at the time of redefinition
Ancillary conditions
Must be achieved before changing the definition
Mandatory criteria
High reliability of optical frequency standards
High reliability of ultra high stability T/F links
Continuous improvement of the realization and time scales after redefinition
Regular contributions of optical clocks to UTC(k)
Availability of commercial optical clocks
Improved quality of the dissemination towards users
Validation that optical frequency standards (OFS) are at a level 100 times better than Cs
Continuity with the definition based on Cs
Regular contributions of OFS to TAI as secondary representations of the second
Availability of sustainable techniques for OFS comparisons
Knowledge of the local geopotential with a sufficient uncertainty level
Definition allowing future more accurate realisations
Access to the realisation of the new definition
Achieved
In progress
https://arxiv.org/abs/2307.14141

## Slide 13: Mandatory criteria and ancillary conditions
Status should be advanced, even if not completely achieved at the time of redefinition
Ancillary conditions
Must be achieved before changing the definition
Mandatory criteria
High reliability of optical frequency standards
High reliability of ultra high stability T/F links
Continuous improvement of the realization and time scales after redefinition
Regular contributions of optical clocks to UTC(k)
Availability of commercial optical clocks
Improved quality of the dissemination towards users
Validation that optical frequency standards (OFS) are at a level 100 times better than Cs
Continuity with the definition based on Cs
Regular contributions of OFS to TAI as secondary representations of the second
Availability of sustainable techniques for OFS comparisons
Knowledge of the local geopotential with a sufficient uncertainty level
Definition allowing future more accurate realisations
Access to the realisation of the new definition
Achieved
In progress
https://arxiv.org/abs/2307.14141

## Slide 14: Robust optical clocks
Monitoring and automatic control systems
Robust & autonomous frequency stabilisation of laser systems
Automated optical alignment
Remote monitoring and control
Software to schedule and perform operational checks, e.g. minimisation of micromotion in trapped ion clocks, magic wavelength calibration in lattice clocks
Tools for on-the-fly correction of systematic frequency shifts
With sufficiently low latency for UTC(k) steering
Also useful for timely contributions to TAI and faster data analysis in international comparisons
e.g. uptimes exceeding 80 % over several weeks demonstrated for NPL-Sr1

## Slide 15: Uptime of the NMIJ optical lattice clock over 6 months
T. Kobayashi et al., Metrologia 57, 065021 (2020)

## Slide 16: Mandatory criteria and ancillary conditions
Status should be advanced, even if not completely achieved at the time of redefinition
Ancillary conditions
Must be achieved before changing the definition
Mandatory criteria
High reliability of optical frequency standards
High reliability of ultra high stability T/F links
Continuous improvement of the realization and time scales after redefinition
Regular contributions of optical clocks to UTC(k)
Availability of commercial optical clocks
Improved quality of the dissemination towards users
Validation that optical frequency standards (OFS) are at a level 100 times better than Cs
Continuity with the definition based on Cs
Regular contributions of OFS to TAI as secondary representations of the second
Availability of sustainable techniques for OFS comparisons
Knowledge of the local geopotential with a sufficient uncertainty level
Definition allowing future more accurate realisations
Access to the realisation of the new definition
Achieved
In progress
https://arxiv.org/abs/2307.14141

## Slide 17: Frequency combs for local optical clock comparisons
Sr lattice clock
Ultrastable lasers
Cs fountains
Hydrogen masers
Yb+ optical clock
Femtosecond optical frequency combs can be used for
Absolute frequency measurements (referenced to the SI unit of frequency)
Optical frequency ratio measurements
Frequency combs themselves contribute negligible uncertainty to either type of measurement

## Slide 18: Absolute frequency measurements
Optical clock
frep
f0
Femtosecond comb
H-maser
Caesium fountain
Independent measurements made relative to local Cs fountain primary frequency standards or TAI
Limited by uncertainty of the best Cs primary standards (low parts in 1016 after 20 – 30 days)
PTB 2012
NPL 2012
NPL 2014
PTB 2014
NPL 2018
PTB 2021
171Yb+ E3
87Sr
JILA 2007
JILA 2008
JILA 2020
SYRTE 2008
SYRTE 2013
SYRTE 2016
NMIJ 2009
NMIJ 2015
NMIJ 2014
PTB 2011
PTB 2014
PTB 2014
PTB 2015
NICT 2012
NICT 2017
NICT 2020
NPL 2020
PTB 2020
NIM 2015

## Slide 19: Optical frequency ratio measurements
frep
f0
Femtosecond comb
Optical clock 2
Optical clock 1
Optical frequency ratio measurements can have better stability and accuracy than absolute frequency measurements
where i
=
f0 + fbi
mi frep
~
106 – 107
f2
f1
m1 frep + f0 + fb1
m2 frep + f0 + fb2
=
=
1 + 2
1 + 1
m2
m1
i.e. frep cancels to first order

## Slide 20: Local optical frequency ratio measurements
LNE-SYRTE
RIKEN
Yb/ Sr
Yb/ Sr
NIST/JILA
INRIM/PTB
RIKEN
PTB
NPL
Yb+/ Sr
Limited only by the stability and accuracy of the optical clocks themselves
  Low parts in 1018 achievable in a few hours
  Can demonstrate reproducibility to an accuracy beyond that of the realisation of the SI second
Clock comparisons without a direct link!

## Slide 21: Remote clock comparisons
Clock A
Clock B
Transportable optical clocks

## Slide 22: Transportable optical clocks
M. Takamoto et al., Nature Photonics 14, 411 (2020)
RIKEN transportable Sr optical lattice clock
J. Stuhler et al., Measurement: Sensors 18, 100264 (2021)
Opticlock: transportable Yb+ (E2) optical clock
J. Grotti et al., Nature Physics 14, 437 (2018)
PTB transportable Sr optical lattice clock

## Slide 23: Remote clock comparisons
Clock A
Clock B
Transportable optical clocks
GEO
Two-way satellite time and frequency transfer (TWSTFT)
MEO
GNSS links

## Slide 24: Optical clock comparisons via satellite
Yb+
Yb+
PTB – NPL
GPS Precise Point Positioning (PPP)
Remote comparison of 171Yb+ (E2) optical clocks
Fractional difference (1.3  1.2)1015
J. Leute et al., IEEE Trans. UFFC 63, 891 (2016)
NPL – SYRTE – PTB
Broadband two-way satellite time and frequency transfer (TWSTFT)
Dedicated transponder (expensive!)
Remote comparison of 5 optical clocks in 3 European laboratories
Comparison uncertainties at the low parts in 1016 level
F. Riedel et al., Metrologia 57, 045005 (2020)
Long averaging times required
Uncertainty achieved does not match clock performance
Sr
Sr
NICT – PTB
Carrier-phase-based two-way satellite time and frequency transfer (TWCP)
Remote comparison of 87Sr lattice clocks separated by a baseline of 9000 km
Fractional difference (1.1±1.6)×10−15
Hachisu et al, Opt. Lett. 39, 4072 (2014)
Sr
Sr

## Slide 25: Remote clock comparisons
Clock A
Clock B
Transportable optical clocks
GEO
Two-way satellite time and frequency transfer (TWSTFT)
MEO
GNSS links
LEO
Optical or microwave links to satellites in low Earth orbit

## Slide 26: ACES microwave link
Part of the Atomic Clock Ensemble in Space (ACES) payload to be installed on the International Space Station
Bidirectional Ku-band link with high modulation rate of 100 Mchips/s
Additional S-band downlink allows ionospheric delay to be determined and cancelled
Specified performance ADEV  1.61016 @ 1 day, but intercontinental frequency comparison accuracies at the 1017 level achievable with one week of averaging if no cycle slips
Free-space optical links also being investigated, e.g. at NIST
Image: W. Schäfer, T. Feldmann, J. Phys.: Conf. Ser., 723 012038, 2016, CC BY 3.0.
Image: ESA–D. Ducros

## Slide 27: Remote clock comparisons
LEO
Optical or microwave links to satellites in low Earth orbit
Clock A
Clock B
Transportable optical clocks
GEO
Two-way satellite time and frequency transfer (TWSTFT)
MEO
GNSS links
Optical fibre links

## Slide 28: European optical fibre network for clock comparisons
NPL
PTB
LNE-SYRTE
INRIM
LPL
Sr
Sr
Sr
Yb
Hg
Yb+
Yb+
Uses the 1.5 µm transmission band of standard telecommunications fibre
Environmentally induced phase fluctuations must be measured and corrected
Sources of noise:
Vibrations
Temperature changes
Frequency comparison uncertainties currently reaching the low part in 1017 level
< 1 part in 1018 possible

## Slide 29: Tests of fundamental physics using the network
NPL
PTB
LNE-SYRTE
INRIM
LPL
Sr
Sr
Sr
Hg
Yb+
Yb+
Test of Lorentz invariance
Comparison of Sr optical lattice clocks in 3 laboratories
Search for daily variations in clock frequency differences tests the assumption that there is no preferred frame of reference in the universe
Factor of 2 improvement on previous tests
P. Delva et al., Phys. Rev. Lett. 118, 221102 (2017)
Search for transient variations of the fine structure constant and dark matter
Comparison of Sr, Hg and Yb+ optical clocks in 3 laboratories
Search for variations in optical frequency ratios that are coherent across the entire network
Improved constraints on transient variations of , especially for timescales > 100 s
B. M. Roberts et al., New J. Phys. 22, 093010 (2020)

## Slide 30: March 2022 comparison campaign
11 optical clocks in 7 different countries
Used satellite links as well as fibre links, so more optical clocks could participate
Optical fibre links between INRIM, LNE-SYRTE & PTB
Local links with ultrastable lasers
Optical clocks

## Slide 31: Towards intercontinental clock comparisons via fibre?
Transatlantic fibre linking Halifax and Southport (5860 km long)
So far mainly used for environmental sensing
  Initial demonstration: G. Marra et al., Science 361, 486 (2018)
  Isolation of signals from individual spans between optical repeaters: G. Marra et al., Science 376, 874 (2022)
Also offers prospects for intercontinental clock comparisons
  Environmentally induced noise lower than for terrestrial fibres
  Encouraging predictions of stability achievable using two adjacent fibres for noise cancellation
Actual number of repeaters (128) reduced for illustrative purposes

## Slide 32: Mandatory criteria and ancillary conditions
Status should be advanced, even if not completely achieved at the time of redefinition
Ancillary conditions
Must be achieved before changing the definition
Mandatory criteria
High reliability of optical frequency standards
High reliability of ultra high stability T/F links
Continuous improvement of the realization and time scales after redefinition
Regular contributions of optical clocks to UTC(k)
Availability of commercial optical clocks
Improved quality of the dissemination towards users
Validation that optical frequency standards (OFS) are at a level 100 times better than Cs
Continuity with the definition based on Cs
Regular contributions of OFS to TAI as secondary representations of the second
Availability of sustainable techniques for OFS comparisons
Knowledge of the local geopotential with a sufficient uncertainty level
Definition allowing future more accurate realisations
Access to the realisation of the new definition
Achieved
In progress
https://arxiv.org/abs/2307.14141

## Slide 33: Clocks and gravity
f = f W / c2
W = gravity potential difference (gravitational + centrifugal components)
f = clock frequency
c = speed of light
Relativistic redshift between two Earth-bound clocks at rest is
Correction terms < 1018 as long as time-variable effects are considered
Frequency shift ~ 1 part in 1016 per metre difference in height at the Earth’s surface
Absolute potentials needed for contributions to international time scales (relative to conventional W0)
Potential differences W sufficient for clock comparisons

## Slide 34: Definition of International Atomic Time (TAI)
Proper quantities are the direct results of observations, e.g. measurement of local time  by an ideal clock
Coordinate quantities are dependent on conventional choices, e.g. International Atomic Time (TAI)
Recommendation CCTF 3 (2017)
TAI is a continuous time scale produced by the BIPM based on the best realizations of the SI second …
In the transformation from the proper time of a clock to TAI, the relativistic rate shift is computed with respect to the conventionally adopted equipotential W0 = 62 636 856.0 m2s−2 of the Earth’s gravity potential …
UTC … is the only recommended time scale for international reference and the basis of civil time in most countries
UTC differs from TAI only by an integral number of seconds …

## Slide 35: Physical geodesy
Determination of W(r) is one of the primary goals of geodesy
  Cannot be computed directly
  Has to be determined indirectly through measurements taken at or near the Earth’s surface
A complicated problem, closely related to the definition of heights, equipotential surfaces, and the geoid
Required uncertainty for optical clock comparisons ~ 0.1 m2s−2
H. Denker et al., Journal of Geodesy 92, 487 – 516 (2018)
h = ellipsoidal height (purely geometric) e.g. from GNSS
H = orthometric height (physical height), measured along the curved plumb line
N = geoid height or geoid undulation

## Slide 36: Gravity potential for optical clock comparisons
ITOC project
Geometric levelling approach
Based on geometrical levelling and gravity observations
A differential technique – gives only potential differences
Used for local clock comparisons at INRIM, LNE-SYRTE, NPL and PTB
Sub-mm accuracy over short distances, but systematic errors accumulate over larger distances
Usually connected to a national tide gauge, leading to inconsistencies of > 0.5 m across Europe
GNSS/geoid approach
Combines results from gravity field modelling with GNSS position measurements
Can deliver absolute potential values

## Slide 37: Gravity measurement campaigns
1 absolute gravity measurement
35 relative gravity measurements
INRIM (Sept 2013)
1 absolute gravity measurement
83 relative gravity measurements
PTB (Mar and Oct 2014)
2 absolute gravity measurements
64 relative gravity measurements
NPL (Mar 2014)
LNE-SYRTE (Oct 2014)
3 absolute gravity measurements
99 relative gravity measurements
Map ©iStock.com/scibak
Consistency checks of historical gravity data
Filled gaps in coverage

## Slide 38: New European Gravimetric (Quasi)Geoid Model (EGG2015)
Accuracy ~ 0.2 m2s−2
Long wavelength components computed from a global geopotential model (GOCO05S)
Short wavelength components computed from high resolution digital elevation models (topographic data)
Medium wavelength structures recovered from terrestrial gravity data
H. Denker et al., Journal of Geodesy 92, 487 – 516 (2018)
Equivalent to 2 cm or 2×10−18 in fractional accuracy

## Slide 39: Time-variable components of the gravity potential
Largest effect is Solid Earth tides:
Gravity potential differences between NPL and PTB
Peak-to-peak range ~ 5 m2/s2
Differences increase as East-West separation increases
C. Voigt et al., Metrologia 53, 1365 – 1383 (2016)
1.3 m2/s2
(1.31017)
7 hours

## Slide 40: Mandatory criteria and ancillary conditions
Status should be advanced, even if not completely achieved at the time of redefinition
Ancillary conditions
Must be achieved before changing the definition
Mandatory criteria
High reliability of optical frequency standards
High reliability of ultra high stability T/F links
Continuous improvement of the realization and time scales after redefinition
Regular contributions of optical clocks to UTC(k)
Availability of commercial optical clocks
Improved quality of the dissemination towards users
Validation that optical frequency standards (OFS) are at a level 100 times better than Cs
Continuity with the definition based on Cs
Regular contributions of OFS to TAI as secondary representations of the second
Availability of sustainable techniques for OFS comparisons
Knowledge of the local geopotential with a sufficient uncertainty level
Definition allowing future more accurate realisations
Access to the realisation of the new definition
Achieved
In progress
https://arxiv.org/abs/2307.14141

## Slide 41: Optical clock contributions to TAI
Sr lattice clock
frep
f0
Femtosecond comb
H-maser
Calibrate a commercial clock for which phase comparison data is submitted on a continuous basis to the BIPM
Use the frequency of the optical clock as listed in the “Recommended values of standard frequencies” http://www.bipm.org/en/publications/mises-en-pratique/standard-frequencies-second
For first submission of a new standard
  Several calibration values must be provided (do not need to be contiguous)
  Operation of the frequency standard and its uncertainty evaluation must be described in a peer-reviewed publication
Material must be reviewed and approved by the CCTF WGPSFS before the clock can be used for steering of TAI

## Slide 42: List of recommended frequency values
Recommended frequencies and uncertainties are assigned by the CCL-CCTF Frequency Standards Working Group (WGFS)
Values are periodically updated and published at https://www.bipm.org/en/publications/mises-en-pratique/standard-frequencies.html
Straightforward with only absolute frequency measurements
2017
2015
2021
Direct optical frequency ratio measurements make the situation more complicated
No longer possible to treat each optical clock in isolation when considering the available data

## Slide 43: Handling over-determined sets of clock comparison data
H. S. Margolis and P. Gill, Metrologia 52, 628 (2015)
Methods have been developed for analysing over-determined sets of clock comparison data
  To check the level of internal self-consistency
  To derive a self-consistent set of optimal values for the ratios between their operating frequencies
All data stored as frequency ratios (optical frequency ratios, microwave frequency ratios or optical-microwave frequency ratios)
Correlations between measured quantities can be included (where known)
Use a least-squares adjustment procedure, based on the approach used by CODATA to provide a self-consistent set of recommended values of the fundamental physical constants
First used to update the list of recommended frequency values in September 2015
L. Robertsson, Metrologia 53, 1272 (2016)
  By 2017 a second independent technique was also available

## Slide 44: Importance of correlations
Neglecting correlations can lead to biased frequency values and underestimated uncertainties
Correlations can arise from both statistical and systematic uncertainties
Consider
  Correlations between clocks based on the same atomic species
    E.g. due to common theoretical or experimental values of atomic coefficients
  Correlations between different clocks in the same institution
    E.g. due to common gravitational redshift correction
  Correlations arising from common data
    Several ratio measurements involving the same clock, performed at the same time
    Absolute frequency measurements performed using TAI as a reference, even if several months apart
H. S. Margolis and M. Pizzocaro, Guidelines on the evaluation and reporting of correlation coefficients between frequency ratio measurements (2020)
http://empir.npl.co.uk/rocit/project-outputs/

## Slide 45: 2021 update to the list of recommended frequency values
105 input data points (72 absolute frequencies, 33 ratios)
Correlations included for the first time (483 correlation coefficients)
6 secondary representations of the second now have uncertainties  2  1016

## Slide 46: Optical clock contributions to TAI
https://webtai.bipm.org/database/show_psfs.html

## Slide 47: Mandatory criteria and ancillary conditions
Status should be advanced, even if not completely achieved at the time of redefinition
Ancillary conditions
Must be achieved before changing the definition
Mandatory criteria
High reliability of optical frequency standards
High reliability of ultra high stability T/F links
Continuous improvement of the realization and time scales after redefinition
Regular contributions of optical clocks to UTC(k)
Availability of commercial optical clocks
Improved quality of the dissemination towards users
Validation that optical frequency standards (OFS) are at a level 100 times better than Cs
Continuity with the definition based on Cs
Regular contributions of OFS to TAI as secondary representations of the second
Availability of sustainable techniques for OFS comparisons
Knowledge of the local geopotential with a sufficient uncertainty level
Definition allowing future more accurate realisations
Access to the realisation of the new definition
Achieved
In progress
https://arxiv.org/abs/2307.14141

## Slide 48: Optical steering of UTC(k) time scale
Phase comparator
Frequency offset generator
UTC(NPL)
HM4
Frequency steers applied 3 times a week, based on data from NPL-CsF2 and Circular T
UK time scale
Frequency steers applied hourly, based on data from Sr and Yb+ optical clocks
Frequency offset generator
UTCx(NPL)
HM6
Experimental prototype

## Slide 49: Automated data validation and steering
Automated validation of hour-long chunks of data
Maser frequency deviation from 10 MHz computed using the recommended frequency values for each optical clock, and corrected using the external frequency offset generator
Optical clock deadtimes bridged with predicted values based on previous maser behaviour.

## Slide 50: Stability of the optically steered time scale
Phase step in output of frequency offset generator
Incorrect steers applied due to improperly flagged optical clock uptimes
Manual correction
Maximum 0.64 ns deviation from UTC, despite gaps in optical clock data (32.3%)
Similar tests performed at LNE-SYRTE
Direct comparison between the two prototype time scales
Agreement better than between the corresponding operational time scales

## Slide 51: Mandatory criteria and ancillary conditions
Status should be advanced, even if not completely achieved at the time of redefinition
Ancillary conditions
Must be achieved before changing the definition
Mandatory criteria
High reliability of optical frequency standards
High reliability of ultra high stability T/F links
Continuous improvement of the realization and time scales after redefinition
Regular contributions of optical clocks to UTC(k)
Availability of commercial optical clocks
Improved quality of the dissemination towards users
Validation that optical frequency standards (OFS) are at a level 100 times better than Cs
Continuity with the definition based on Cs
Regular contributions of OFS to TAI as secondary representations of the second
Availability of sustainable techniques for OFS comparisons
Knowledge of the local geopotential with a sufficient uncertainty level
Definition allowing future more accurate realisations
Access to the realisation of the new definition
Achieved
In progress
https://arxiv.org/abs/2307.14141
Significant progress is being made towards fulfilment of the mandatory criteria…
… but need to reach consensus on best option for the redefinition

## Slide 52: NPL e-learning (free)
https://elearning.npl.co.uk
~ 1 day
Provides an introduction to measurement and data analysis methods used to quantify clock performance.
~ 1 day
Provides an introduction to the measurement of time and frequency, time dissemination methods and time scales.

## Slide 53: Time and Frequency department at NPL
 70 staff across 3 groups:
National time scale
Optical frequency metrology
Atomic clocks and sensors
https://jobs.npl.co.uk/vacancies/