## Slide 1: Electro-Optic Sampling Beam Positioning Monitor for Relativistic Electron Beams
APS Global Summit
March 17, 2026
Elena L. Ros1, Michael L. Litos1, Junzhi Wang2 , Alexander Knetsch2, Brendan O’ Shea2, Claire Hansel1, Valentina Lee1, Daniel Matteo3, Gerard Andonian3,4, Mark Hogan2, Robert Ariniello2, Shutang Meng1, Tara Hodgetts3,
[1] CU Boulder, Physics, [2] SLAC National Accelerator Laboratory, [3] RadiaBeam Technologies, [4] ULCA, Physics
Notes: Done

## Slide 2
WARG Group:
Mike Litos
Claire Hansel
Valentina Lee
Shutang Meng
Benjamin Braun
Radiabeam:
Daniel Matteo
Tara Hodgetts
Gerard Andonian*
FACET-II:
Alexander Knestch
Junzhi Wang
Yiheng Ye
& Radiation Generation
Wakefield Acceleration
WARG Group
Junzhi Wang
*also at UCLA
Acknowledgements

## Slide 3
U.S. Department of Energy, Office of Science, Office of High Energy Physics, under Award Number DE-SC001796 and award number DE-SC0023977.
National Science Foundation under Award Number PHY-1806053 and Faculty Early Career Development Program Award Number PHY-2047083.
Our Research is Funded By…

## Slide 4
Motivation
Independent transverse and longitudinal measurement of drive and witness beam for Plasma Wakefield Acceleration (PWFA).
Requires precise transverse alignment for emittance preservation.
Requires precise longitudinal separation to minimize energy spread and maximize efficiency.
Non-destructive shot by shot measurement
PWFA Simulation
Witness Beam
Drive Beam
z
~100μm
~10μm

## Slide 5
EO Crystal
γ
e- beam
Laser
k
EO Spatial Encoding Scheme1
Laser
k
EO Crystal
γ
e- beam
EO Crystal
γ
e- beam
Laser
k
EO Crystal
γ
e- beam
Laser
k
EO Crystal
γ
e- beam
Laser
k
Rotated Polarization
Temporal information mapped to transverse profile of laser.
Background: EO Sampling

## Slide 6
k
Laser
Polarizer
Horizontally polarized
γ
e- beam
Polarizer
Vertically Polarized
Image on camera
Witness
Drive
CMOS Camera
EO Cross Polarization Scheme

## Slide 7
EOS BPM Mk.1
EO Crystal
Laser
EO Crystal
e- beam
Leverages correlation between strength of phase retardation and location of e- beam with respect to crystal2.
Electro Optic Sampling Beam Position Monitor

## Slide 8
Crystal 1 Drive Bunch Signal [arb]
Crystal 2 Drive Bunch Signal [arb]
Drive Bunch Δx at EOS-BPM inferred from DTOTR (μm)
Δx from DTOTR2 (μm)
Δx from EOS-BPM (μm)
Publication in progress
Claire Hansel
65 μm Resolution!
Crystal 1
Crystal 2
EOS BPM Mk1. / Experimental Data

## Slide 9
4 Crystal Design
Axicon
e- beam
θ
EO Crystal
k
k
Tilted Pulse Front
Axicon creates circular donut
4 crystals around e- beam axis
Easier alignment more uniform laser intensity
More complete azimuthal coverage allows for position measurement in x and y.
Enhanced signal to noise leads to better temporal and transverse resolution.
5o angle of incidence (θ)
3 ps temporal window
4 10mm x 10mm 100μm thick GaP wafers
EOS BPM Mk.2 Design

## Slide 10
CAD Model/Pictures
EOS BPM CAD
EOS BPM Installed at FACET-II
ebeam
ebeam
laser
laser

## Slide 11
Parameter | Value | Unit
Laser Wavelength | 800 | nm
Laser Angle | 5 | degree
GaP Wafer Thickness | 100 | μm
Drive Beam Charge | 800 | pC
Witness Beam Charge | 800 | pC
Time Delay btw. Beams | 450/135 | fs/μm
Transverse Beam Offset Drive | ~500 | μm
Transverse Beam Offset Witness | ~500 | μm
Bunch Length (Both) | 100/30 | fs/μm
Drive
Witness
EOS BPM Mk.2 Simulation

## Slide 12
EOS BPM Signal
EOS Signal
EOS Signal with Background Subtract

## Slide 13
EOS BPM Mk.2 Time of Arrival Analysis

## Slide 14
OTR Screen
EOS BPM
XTCAV
X Corrector
Dipole
ebeam
OTR XTCAV Off
Drive
Witness
EOS and XTCAV at FACET-II
Time
Energy
OTR XTCAV On
Time
Energy

## Slide 15
EOS and TCAV comparison
Radially Binned Signal
Radius [px]
Radially Binned Signal
Bunch Separation [μm]
Shot Number
Bunch Separation Results
2 μm mean difference between EOS and TCAV

## Slide 16
Signal Intensity on Left Crystal
Drive Intensity
Radius [px]
Corrector Dipole Strength
Binned Intensity [abu]
Preliminary BPM Results

## Slide 17
Expected Signal
α1 = α +π/2
Mask to get 1 ring
Piecewise fit to model
Extract r and α to find bunch position
Repeat for 2nd ring
α1
e-beam
α
Crystal rotated 90o
b
a
a
b
α
α1
α
b
a
a
b
EOS BPM Analysis with Rotated Crystals

## Slide 18
Summary
Novel ultra-fast optical device for measuring beam centroid location with fs/μm level precision.
  Required for emittance preservation of PWFA.
  Required to minimize energy spread and maximize efficiency.
Improvements:
  Flatten wavefront by installing a spatial light modulator.
  Improve analysis algorithms and background subtraction to account for laser fluctuations.
Future Work:
  BPM and TOA measurements with improved wavefront.
  Live analysis algorithms for BPM in addition to TOA.
  Explore the utility of EOS BPM to diagnose tilted beams.
See posters T14 and T45 for more information!

## Slide 19
Questions

## Slide 20
References
[1] A. L. Cavalieri et al., “Clocking Femtosecond X Rays,” Phys. Rev. Lett., vol. 94, no. 11, Mar. 2005, doi: 10.1103/physrevlett.94.114801.
[2] K. Hunt-Stone, R. Ariniello, C. Doss, V. Lee, and M. Litos, “Electro-optic sampling beam position monitor for relativistic electron beams,” Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, vol. 999, p. 165210, May 2021, doi: 10.1016/j.nima.2021.165210.
[3] S. Casalbuoni, H. Schlarb, B. Schmidt, P. Schmüser, B. Steffen, and A. Winter, “Numerical studies on the electro-optic detection of femtosecond electron bunches,” Phys. Rev. ST Accel. Beams, vol. 11, no. 7, Jul. 2008, doi: 10.1103/physrevstab.11.072802.
[4] B. Steffen et al., “Electro-optic time profile monitors for femtosecond electron bunches at the soft x-ray free-electron laser FLASH,” Phys. Rev. ST Accel. Beams, vol. 12, no. 3, Mar. 2009, doi: 10.1103/physrevstab.12.032802.
[5] R. Ariniello, C. E. Doss, V. Lee, C. Hansel, J. R. Cary, and M. D. Litos, “Chromatic transverse dynamics in a nonlinear plasma accelerator,” Physical Review Research, vol. 4, no. 4, p. 043 120, 2022, Publisher: American Physical Society.
doi:10.1103/PhysRevResearch.4.043120
[6] V. Yakimenko et al., “FACET-II facility for advanced accelerator experimental tests,” en, Physical Review Accelerators and Beams, vol. 22, no. 10, p. 101 301, 2019.doi:10.1103/PhysRevAccelBeams.22.101301

## Slide 21
Timing Calibration E301 shift 05/26
Total Time Window: 2 ps
Slope: ~0.51 ps/mm
Resolution: 5.85 fs/pixel
Calculated Slope: 0.27 ps/mm

## Slide 22
Center Finding

## Slide 23
EOS BPM Measurement at FACET-II
OTR Screen
EOS BPM
XTCAV
X Corrector
Dipole
ebeam

## Slide 24
TCAV E301 shift 05/26

## Slide 25
Low witness charge E301 shift 05/26

## Slide 26
EOS BPM Mk.2 Time of Arrival Analysis

## Slide 27
Old EOS Signal

## Slide 28
EOS BPM Mk.2 Time of Arrival
SNR Drive Beam: 2.20
SNR Witness Beam: 0.92
SNR Drive Beam: 73.1
SNR Witness Beam: 50.4

## Slide 29
Time of Arrival with 1 crystal
SNR Drive Beam: 1.58
SNR Witness Beam: 0.74
SNR Drive Beam: 11.8
SNR Witness Beam: 4.9

## Slide 30
Study of Longitudinal Separation Resolution
Actual Bunch Separation | Fit
400 fs | 397 fs
405 fs | 402
410 fs | 407
415 fs | 412
420 fs | 418
425 fs | 422
430 fs | 428
~ 3 fs resolution with 15 % noise

## Slide 31
EOS BPM Mk.2 BPM and TOA Results
3 fs relative TOA resolution
<5 μm transverse resolution

## Slide 32
Example of Ringing

## Slide 33
EOS BPM Analysis
α
e-beam
Expected Signal
Temporal window
θ is laser angle of incidence
r

## Slide 34
EOS BPM Analysis with Rotated Crystals
α1
e-beam
Expected Signal
Temporal window
θ is laser angle of incidence
α
α1 = α +π/2
Crystal rotated 90o
Crystal rotated 90o
a
a
b
b

## Slide 35
EOS BPM Analysis with Rotated Crystals
Expected Signal
α1 = α +π/2
Mask to get 1 ring
Piecewise fit to model
Extract r and α to find bunch position
Repeat for 2nd ring
α1
e-beam
α
Crystal rotated 90o
b
a
a
b
α
α1
α
b
a
a
b

## Slide 36
BPM Analysis
Region I
Region II
Region I
α2
α1
r2
r1

## Slide 37
Signal Ringing for Short Bunches
Shorter electron bunches correspond to higher THz frequencies. For GaP the cutoff is ~ 10 THz due to the phase velocity mismatch.
This is better than ZnTe, but there is also the trade off, of lower signal.
The phase mismatch can cause ringing in the signal -> worse TOA resolution.
For FACET-II 100 μm GaP is optimal.

## Slide 38
Study of Longitudinal Separation Resolution
Actual Bunch Separation | Fit
400 fs | 397 fs
405 fs | 402
410 fs | 407
415 fs | 412
420 fs | 418
425 fs | 422
430 fs | 428
~ 3 fs resolution with 15 % noise

## Slide 39
Study of Low Charge Bunches
Bunch Charge (pC) | Fit (fs) | Drive | Witness | Drive | Witness
800 | 447 | 19 | 12 | 450 | 350
700 | 447 | 14 | 9 | 300 | 230
600 | 447 | 11 | 7 | 260 | 200
500 | 447 | 7 | 4 | 190 | 150
400 | 448 | 5 | 3 | 120 | 100
300 | 448 | 3 | 2 | 80 | 70
200 | 450 | 1 | 1 | 50 | 50
100 | 440 | 0.5 | 0.1 | 35 | 35
Lineout SNR
Azimuthal Summing SNR
Actual: 450 fs
5% noise at 800pC

## Slide 40
Phase Retardation of Crystal
EO Crystal
e-beam
α
Signal from Phase Retardation

## Slide 41
Phase Retardation of Rotated Crystal Orientation
EO Crystal
e-beam
α
Signal from Phase Retardation
EO Crystal rotated 90o
αr
Ir
Ir
I
I

## Slide 42
Chromatic Phase Spreading (Dephasing)
When witness beam is offset from drive beam: different energy slices of witness beam will rotate at different frequency in transverse phase space -> emittance growth (size of ellipse => emittance5.

## Slide 43
Limiting Aperture
- Crystal minimum aperture is 1 cm
Mirrors have 4 mm hole diameter and are 6.7 mm thick.

## Slide 44
EOS BPM Preliminary Experimental Results
Single 60 μmx 60 μm x 18 μm, 1.6 nC bunch
Signal
Background
Background image with no signal

## Slide 45
EOS BPM Signal with Time Delay
0.8 ps delay
1.4 ps delay
Nominal t0

## Slide 46
EOS BPM Radially Binned Data

## Slide 47
EOS BPM Time Window
Total Time Window: 2 ps
Slope: ~0.51 ps/mm
Resolution: 5.6 fs/pixel
Calculated Slope: 0.27 ps/mm