## Slide 1: Compact electron spectrometry for tens-of-GeV laser-plasma accelerators
Xiantao Cheng, Jose Franco Altamirano, Aaron Bernstein,
Rafal Zgadzaj, Michael Downer
Xiantao Cheng
PhD 2024
~ 10 GeV LPA electron results
Anaculaesei, Matt. Rad.
Extremes 9, 014001 (2024)
Texas
Picksley, PRL 133, 255001 (2024)
Shrock, PRL 133, 045002 (2024)
U. Md.
LBNL
LBNL
Gonsalves,
angle
(mrad)
0
1
-1
PRL 122, 084801 (2019)
pC mrad-1
GeV-1
Ee (GeV)
6 pC
Pathways to 20-30 GeV (SLAC-like) LPA electron bunches:
• hybrid LWFA/PWFA
• multi-staging
Advanced Accelerator Concepts Workshop – UCLA - July 26-31, 2026
ELI-NP
~17GeV
Calin Hojbota
This Workshop
Plenary WG1 Mon 8:45am

## Slide 2: SLAC 30 GeV e-spectrometer
Multi-meter, multi-ton magnetic e-spectrometers
compromise the unique virtues of LPAs:
• compactness
• low expense

## Slide 3
Dual-screen magnetic spectrometers help to deconvolve
electron launch angle variations from energy variations...
...but are limited in accuracy because they rely on
diffuse spectral features of the e-beam.
Shape of peaks
changes between
screens.
Resulting errors
in electron energy
determination
increase rapidly
for Ee ≳ 3 GeV.
B. Pollock et al., “Two-screen method for determining electron beam energy and deflection from laser wakefield acceleration,”
Particle Accelerator Conference (2009)

## Slide 4: Fiducial wire e-spectrometer at Texas PW Laser
electron signal
1
1st prototype deployed:
Wang et al., Nat. Commun. 4, 1988 (2013)
diameter: 127 µm
T
4×4cm
“compact”
(ISO 2768 CNC machining precision)
7 cm
e-/X-ray
source
location:
150 J,
150 fs
f/45
-tron X-rays
Fujifilm
BAS—IP
50 or 100µm
pixels
2 wire
arrays
× 0.1

## Slide 5: Magnetic field structure can be simplified at high e-energy
intra-magnet
e- trajectories
[cm]
[cm]
directly
measured
Errors , from use of
equivalent fields...
...are negligible
[GeV]
[GeV]
* pivot point
,
*
compared to contributions to , from variances in:
(source location)
• (magnetic field)
• (fiducial wire positions)
• (shadow positions)

## Slide 6: Estimation of measurement error
Error ( in electron energy propagates from errors in 4 independent variables:
shadow
positions:
(≲ IP pixel)
fiducial wire
positions:
(CNC machin-
ing precision)
source
position:
(X-ray
triangulation)
magnetic
field:
(Hall probe
accuracy)
3
Electron trajectory formulas:

## Slide 7: Measurement error contours )
4
D=5 cm
L=3 m
D=5 cm
L=3 m
D=10 cm
L=10 m
D=10 cm
L=10 m
Such contours are
essential for evaluating
spectrometer designs
for a specified energy
range (E), space (L),
and accuracy (∆E/E).

## Slide 8
Energy-dependent shape & contrast of wire shadows provide
an independent (albeit coarse) e-spectrometer.
Design of fiducial wires
• cylindrical
(𝜃-independent)
• diameter ~ pixel width
(minimize
• high Z
(high-contrast shadow)
GEANT4 simulations reveal energy-dependent shadow shapes:
(arb. units)
Relative recorded
brightness at detector

## Slide 9
Everyday analysis of real data (Ex. 1)
x [pixel no.]
Step 1) examine raw data
• perfectly straight
electron “streak”
with E-independent
width ∆ ~ 1 mrad.
• Shadows e1-e9 on
e-streak; X1 – X6
on 𝛽-tron X-rays.
~ 2 GeV
Ross filters for 𝛽-tron
X-ray spectrum analysis

## Slide 10
Everyday analysis of real data (Ex. 1)
Step 2) X-ray triangulation of source location
z [mm]
0
500
𝛽-tron X-ray data from IP
row 1
row 2
IP shifted ~1mm
due to loose fit in
holder.
1 mm IP shift
row 1
shadows
row 2
shadows
corrected for
IP shift

## Slide 11
Everyday analysis of real data (Ex. 1)
Step 3) Reconstruction of launch angle and e-spectrum
Energy-dependent ,
inconsistent with straight
electron “streak”.
after correcting magnet
position error.

## Slide 12: Ex. 2a: E > 3 GeV features + energy-dependent
8
Anaculaesei et al. “Acceleration of high-charge electron bunch to10 GeV ni a 10cm nanoparticle-assisted wakefield accelerator,”
Matt. Rad. Extremes 9, 014001 (2024)
preliminary piece-wise
analysis of features with
different ∆y deflections.
comprehensive
reconstruction
> 3 GeV

## Slide 13: Ex. 2b: Comparison with simulated dual-screen measurement
9
E-beam from Ex. 2a, replicated in GEANT4, exposes 2 virtual screens.
screen 1
screen 2
electron energy obtained
from fiducial wire analysis
: electron energy obtained
from dual-screen analysis
≤ 5%
discrepancy
large discrepancies for E ≳ 3 GeV
Reason for discrepancy:
• Peak locations determined
jointly by and .
• The mixture depends on
screen location.
• Fiducial wire method is not
subject to this ambiguity.

## Slide 14
Ex. 3: “Zeeman” splitting of fiducial wire shadows
plasma
retro-
reflector
Studies of inverse Compton-scattered X-rays:
Hannasch et al., Sci. Rpts.11, 14368 (2021)
4 split shadows
close-up of split e6 shadow

## Slide 15: “Zeeman” splitting analysis
7
Schumaker et al., “Ultrafast electron radiography of magnetic fields in high-intensity laser-solid interactions,” PRL 110, 015003 (2013)
depth ~ 10 µm
⊗
Separated spectra of HE and LE electron bunches
total angle entering magnet
(launch + PM deflection)
Total angle [mrad]
Raj et al., “Probing ultrafast magnetic-field generation by current filamentation instability in fs relativistic laser-matter interactions,” PR Res 2, 023123 (2020)

## Slide 16
C. I. Hojbota (PI), O. Z. Labun, D. D. Phan, R. Sain, D. Stec, B.M. Hegelich
A. J. Franco-Altamirano, T. Araujo, Ross Rudzynsky, Afsana Mimi Raka, Erick Martinez, Bryan Dinh, R. Zgadzaj, M. C. Downer
B. Diaconescu (Co-PI/LER), L. Tudor, V. Nastasa, L. Neagu, S. Balascuta, K. Sphor, P.G. Bleotu, V. Phung Ling Jen, C. Ticos, S. Ter-Avetisyan, M. Patrascoiu, M. Gugiu, S. Popa, V. Popescu, D. Ursescu, A. Dumitru, L. Vasescu, A. Nazaru, D. Nistor, S. Rusnac, A. Lazar, R. Schmelz, I. Dancus, P. Ghenuche (Co-PI), V. Malka, C.A. Ur, et al.
B.Hidding, C. Aniculaesei, M. Cerchez, M. Osenberg, O. Bilen, T. Heinemann, E. Hartmann, N. Thomas, T. Wilson, A. Pukhov
P. Franke, T. Ha, E. McCary, S. Milton, L. Labun, J. Paye, C. Neacsu, G. Tiwari, G. Plateau, H. Smith
V. Winter
Fiducial wire spectrometer at ELI-NP

## Slide 17
Double dipole configurations
Two dipoles - enhanced dispersion
Fiducial wire array
B
B
Half chicane – lower dispersion
Fiducial wire array
B
B

## Slide 18: Fiducial wire shadow visability
Monoenergetic e-bunches, 1mrad beam divergence, 100 μm spatial resolution of detctor, for 250 μm W wires
10GeV
15GeV
25GeV
125 μm diam. W wire
250 μm diam. W wire

## Slide 19: Fiducial wire shadow with finite bandwidth and divergence are shallower but still detectable
Uniform spectral density 10GeV – 14GeV
250 μm diam. W wire
125 μm diam. W wire
Fiducial wire shadow projection

## Slide 20
SUMMARY
• The fiducial wire spectrometer provides a low-cost, versatile, small-laboratory method for
determining energy and launch angle of electrons from pointing-unstable LPAs with few-%
accuracy up to tens of GeV.
• Avoids reference to diffuse spectral features of the beam itself.
• We have developed principles for...
...designing fiducial wires
...analyzing spectral errors
...optimizing accuracy
• We have shown that individual fiducial wire shadows independently measure...
...electron energy
...pointing angle
...angular divergence
...electron density
...magnetic fields from intense
laser-matter interactions
X. Cheng et al., “Compact high-resolution multi-GeV electron spectrometer for PW-laser-driven plasma accelerators,” - forthcoming.
X. Cheng,”Advanced diagnostics of laser-driven plasma accelerators,” PhD Thesis, U. Texas at Austin.
Financial support: US Dept. of Energy grant DE-SC0011617.

## Slide 21: Shadow visibility vs. E and L2 (GEANT4 simulations)
5
d = 0.127 mm, tungsten (W)