## Slide 1: FPGA-Based Autonomous Data Logging for Real-Time Beam Monitor Signal Processing in Proton Therapy
Christian Groh – Michael Eichin
RealTime2026, 29 May 2026

## Slide 2: Agenda
Proton Therapy @ PSI
Introduction Pencil Beam Scanning
Beamline variations
Motivation for autonomous logging
Technical Details of autonomous logging
Results and performance

## Slide 3
Proton Therapy @ PSI

## Slide 4: Proton Therapy @ PSI
PSI facility to treat tumour patients with protons
COMET Superconducting accelerator. One accelerator for all treatment areas
Gantry1 PSI development. In operation since 1996. Worldwide 1st gantry with spot scanning pencil beam technology.
Gantry2 PSI development. Performance optimized Gantry design for continuous scanning technologies.
OPTIS2 PSI development. Horizontal fixed beamline based on scattering technology.
Gantry3 Commercial gantry from VARIAN Medical Systems. Based on raster scanning technology
COMET Superconducting accelerator. One accelerator for all treatment areas
Gantry1 PSI development. In operation since 1996. Worldwide 1st gantry with spot scanning pencil beam technology.
Gantry2 PSI development. Performance optimized Gantry design for continuous scanning technologies.
OPTIS2 PSI development. Horizontal fixed beamline based on scattering technology
Gantry3 Commercial gantry from VARIAN Medical Systems. Based on raster scanning technology

## Slide 5: Proton Therapy @ PSI
OPTIS2 Gantry 2 Gantry 3

## Slide 6
Pencil Beam Scanning

## Slide 7: Proton Therapy – Pencil Beam Scanning – Technology
Beam leaves the particle accelerator (cyclotron) with
Beam is degraded to specific energy ()
Transported through a beamline to a gantry – length
Adapt setup for patient specific requirements on gantry
Single pencil beam leaves nozzle towards patient
3 mm sigma
D. Meer

## Slide 8: Proton Therapy – Pencil Beam Scanning – Technology
Deposit energy in a tumor by scanning its volume
Change scanning magnet currents for lateral positioning
Change beam energy for depth variations (degrader)
D. Meer
Pencil beam
Depth scanning
Lateral scanning
D. Meer
Gantry2

## Slide 9: Proton Therapy – Pencil Beam Scanning – Summary
Spot scanning inside energy layer orthogonal to beam direction
Scanning depth is modulated by energy changes
Beam current has technical importance (detectors, treatment time)

## Slide 10
Beamline (transmission) variations

## Slide 11: Transmission
Transmission is the ratio of:
: Proton current at coupling point
: Proton current at accelerator
Large variations in transmission from day to day
Small variations in transmission during a day
200 MeV:
Control over is important

## Slide 12: Transmission Compensation
Measure transmission
Generate scaling factor
Accelarator current can be freely chosen

## Slide 13
Motivation for autonomous logging

## Slide 14: PSI Site Setup
Gantry 1 and 2 are fully controlled by PSI
Energy changes
Position changes
Beam-current changes
Beam on/off
Gantry 3
Commercial system
    Responsible for treatment
    Decides on parameters (energy, current, …)
    Must verify treatment parameters
Needs access to PSI beam line
coupling point

## Slide 15: Gantry 3 Setup
G3 system delegates energy and beam-current changes to PSI IOC (TcsAdapter)
TcsAdapter is interface to beamline
TcsAdapter receives commands from G3
    Open/CloseBeamStoppers
    SetBeamEnergy
    SetBeamCurrent
Commands must be executed as fast as possible
TcsAdapter unaware of spot position changes
TcsAdapter unaware of beam on/off

## Slide 16: Gantry 3 Setup
TcsAdapter must measure transmission
TcsAdapter must scale beam currents
TcsAdapter must not negatively interfere with treatment
Solution:
Autonomously log data
Process during idle times

## Slide 17
Technical Details of
Autonomous Logging

## Slide 18: Hardware
IOC:
    VME – MVME8105 with VxWorks BSP
    hosts TcsAdapter software
FPGA:
    XMC board from Alpha Data
    AMD Kintex Ultrascale FPGA
    Sits on IOC
    implements autonomous logging
DSE: (detector specific electronics)
    PSI development
    AMD Kintex-7 FPGA
    Sits along beamline near detectors
    Sends detector data every
CH1
CH3+4
CH2
DSEs

## Slide 19: Naïve Logging
naïve: log all incoming data
Example below:
    10’000 data points recorded
    4’300 useless (beam was off)
Only 4MB available; 16 Bytes per sample
CH1
CH3+4
CH2
DSEs

## Slide 20: Autonomous Threshold Logging
Threshold TH1: start logging if CH1 > TH1
Threshold TH2: stop logging if CH1 < TH2
TH2 < TH1
Beam off periods are not recorded
CH1
CH3+4
CH2
DSEs

## Slide 21: Autonomous Threshold Logging
What is logged:
    timestamp
    CH1, CH2, CH3, CH4
    var CH1, var CH2
Compression :
    Average over samples
    Variance (std. dev.) calculated for ch1 and ch2
    Longer logging
No cost to IOC
byte

## Slide 22: Interaction with IOC
FPGA: logs into RAM until full
IOC: access logged data on commands
    setBeamEnergy
    setBeamCurrent
    store it in IOC memory and reset FPGA logging
    Should take less than160 ms

## Slide 23: Post Processing on IOC
CloseBeamStoppers (>4seconds)
Persist logged data
Calculate statistics from data
Calculate scaling factors

## Slide 24
Results and performance

## Slide 25: Single Energy Layer
Coupling point current for
43 spots in this layer; 107989 samples
Spot durations
Rising and falling signal need to be removed

## Slide 26: Single Energy Layer
Coupling point current for
43 spots in this layer; 107989 samples
Spot durations
Rising and falling signal need to be removed
Average and standard deviation are easy to calculate
This must be done per channel and per energy layer

## Slide 27: Memory Readout from FPGA
During
    setBeamEnergy
    setBeamCurrent
Memory full in 3% of cases
Average read out time
Median read out time

## Slide 28: Calculate statistics on IOC
During
    closeBeamStoppers
    We have
2 million samples are processes on average
Average calc. time
Median calc. time

## Slide 29: Application: Transmission Compensation
Scaling active since 2025-10-13
~16’500 transmission curves
Logging active since 2025-04-29

## Slide 30: Energy Dependent Scaling
113 references between 2025-10-13 and 2026-03-19

## Slide 31: Also Useful for …
characterizing detectors
debugging detectors

## Slide 32: Where to go …
Logging:
G2 and G3 working
O2 coming soon
Possible future applications:
Beam current regulation on G2
UHDR on G2

## Slide 33
Thank you.
Questions?
Project Team:
Michael Eichin – FPGA
Christian Groh – Software
Machine Operation:
Vitaly Finogeshin
Advisors:
David Meer
Urs Rechsteiner
Zema Chowdhuri