## Slide 1: Detector Development & High-Z Sensor Characterisation at PSI
University of Birmingham, 15th July 2026
J. Mulvey, D. Greiffenberg, M. Anastasi, B. Schmitt, A. Bergamaschi,
M. Brückner, M. Carulla, R. Dinapoli, S. Ebner, K. Ferjaoui,
E. Fröjdh, V. Gautam, S. Hasanaj, V. Hinger, M. Hürst, V. Kedych,
T. King, S. Li, L. Lunin C. Lopez Cuenca, A. Mazzoleni, D. Mezza, K. Moustakas,
A. Mozzanica, M. Müller, C. Ruder, S. Silletta, D. Thattil, X. Xie, J. Zhang
PSI Center for Photon Science, 5232 Villigen PSI, Switzerland

## Slide 2
Introduction to PSI
Science at PSI
User Facilities – SLS, SwissFEL, etc.
Photon Science Detector Group – How we support user facilities
High-Z Characterisation at PSI
Overview of High-Z materials
GaAs
CZT
High-Z Future & Applications
Photon Science
Medical Physics
Astronomy

## Slide 3: The Paul Scherrer Institut
PSI - Swiss federal research facility
Largest Swiss research institute for natural and engineering sciences
Hosts large scale research infrastructure including:
SINQ – Neutron scattering
SμS – Muons
COMET – Proton therapy
CHRISP – Particle Physics
SLS2.0
SLS since 2001
SLS2.0 since 2025
SwissFEL
Since 2016
First user operation in 2019
~ 2200 Employees ~320 PhD Students ~200 Postdocs

## Slide 4: Swiss Light Source 1.0 (SLS)
Electron accelerator a
3rd generation synchrotron source
Utilizes Bremsstrahlung to generate X-rays
Multiple electron bunches in circulation a
Not typically coherent
X-ray energy selected through a monochromator
2.4 GeV electrons in 12 ps bunches
288m circumference
5.4 keV median energy a a
400 mA total current (< 10 mA per bunch)
partially with undulators
100 eV up to 40 keV (> 100 keV for SLS2.0)
M. Böge [1]
Mitchell et al [2]

## Slide 5: From SLS1.0 to SLS2.0
SLS2.0
4th Generation a
2.7 GeV electrons
Still 288m
Remains 400 mA
> 100 keV (50% higher flux at these energies)
Mitchell et al [2]
SLS1.0
3rd Generation Synchrotron light source
2.4 GeV electrons
288m circumference
400 mA total current
100 eV up to 40 keV a
M. Böge [1, 3]

## Slide 6: Swiss Fast Electron Laser (SwissFEL)
High spatial and temporal coherence (laser-like)
2.1 - 5.8 GeV electrons
Extremely short pulse duration (1 – 20 fs)
100 Hz repetition rate
Photon energy between 0.25 and 12 keV
Order of 1010 photons per 12 keV
[4]

## Slide 7: Photon Science Applications
Synchrotron
High-res imaging (ptychography down to nm range)
Structure determination (crystallography, X-ray diffraction)
Absorption imaging (tomography)
Spectroscopy (element specific)
Macromolecular Crystallography Leonarski et al.
Operando diffraction Hocine et al.
Ptychography
Holler et al.
Magnetic contrast ptychography – Butcher et al.
Atomic Dynamics
Chapman et al.
FEL
Ideal for probing dynamic systems on the scale of fs
Damage-free “diffraction before destruction”
Single-shot imaging
Similarly to particle physics
Use the diffraction pattern to work back to understand structure of crystal/protein/etc. under test

## Slide 8: Detector Requirements
Atomic Dynamics
Chapman et al.
Small pixels < 100 μm (even 5 μm)
Energy range: eV → ~ 100 keV
Count rate per pixel:
  SLS 1.0: Few MHz
  SLS 2.0: 10s MHz
  SwissFEL: 100 THz
Large area: 16Mpixel (32x32 cm2)
High frame rates: ~10 kHz

## Slide 9: We’ve already come a long way…
W. Röntgen, 1895
P. G. Jung [5]

## Slide 10: Hybrid detectors – A step forward
Hybrid sensors fit our requirements
Semiconductor sensor
  Direct conversion
Bump bonding
  Minimum pixel pitch ~ 20 μm
  Noise due to input capacitance
Versatile CMOS readout chip
  Fast readout
  Radiation hardness

## Slide 11: Readout Modes
Photon counting
Charge integrating
vth
Number of
photons
Single photon counting (SLS)
Total number of
electron-hole pairs
Charge integrating detectors (XFEL)
Count rate limited
Dynamic range limited

## Slide 12: Readout Modes
Noiseless, very large dynamic range
Stable operation, well established calibration
Fluorescence suppression
Frame rate scales with counter depth
Pile-up ≳1MHz/pixel
Soft X-ray detection prevented by noise
Compromise between noise and dynamic range
Works also at pulsed sources
Detection of multiple soft X-rays
Saturation for long integration times
fluorescence
Total
5
fluorescence
Total
5
+1
+1
+1
+1
Total
4
Total
5.5
Photon counting
Charge integrating

## Slide 13: PSI Detector portfolio
 | Microstrip 1D | Pixel 2D | 
 | 25/50 µm pitch | 75 µm pitch | 25 µm pitch
Single photon counting |  |  | 
Charge integrating |  |  | 
MYTHEN
GOTTHARD
MÖNCH 0.3
JUNGFRAU
EIGER

## Slide 14: (Almost) Everything in-house

## Slide 15: The JUNGFRAU ASIC
A. Mozzanica, 2014 [6]
High-Flux CZT
Si
SMA for HV
Spectroscopic CZT
Charge integrating
Low noise
Dynamic per-pixel gain switching
75 μm pitch
256 x 256 pixels per 2 x 2 cm2 chip
Can resolve single photons in highest gain
JUNGFRAU’s charge integrating architecture allows us to probe sensor properties, such as the afterglow effect, in greater detail

## Slide 16: The JUNGFRAU ASIC
Charge integrating
Low noise
Dynamic per-pixel gain switching
75 μm pitch
256 x 256 pixels per 2 x 2 cm2 chip
Can resolve single photons in highest gain
JUNGFRAU’s charge integrating architecture allows us to probe sensor properties, such as the afterglow effect, in greater detail
Large areas (16Mpixels) can be achieved by tiling modules

## Slide 17: Upgraded Light Sources: Many More Photons!
fluorescence
Total
5
+1
+1
+1
+1
Total
4
Total
5.5
fluorescence
Total
5
Photon counting
Charge integrating with adaptive gain
Noiseless, very large dynamic range
Stable operation, well established calibration
Fluorescence suppression
Frame rate scales with counter depth
Pile-up ≳1MHz/pixel
Soft X-ray detection prevented by noise
Single photon resolution for hard/tender X-rays
  Detection of multiple soft X-ray
Works also at pulsed sources
Dynamic range up to ~104 photons/pixel/s
Supported flux scales with frame rate
Saturations for long integration times
Challenging calibration and cumbersome data processing needs high throughput!
Charge integrating

## Slide 18: Upgraded Light Sources: Many More Photons!
Photon counting
Charge integrating with adaptive gain
Noiseless, very large dynamic range
Stable operation, well established calibration
Fluorescence suppression
Frame rate scales with counter depth
Faster shaping, pile-up tracking
Pile-up ≳1MHz/pixel
Soft X-ray detection prevented by noise
Single photon resolution for hard/tender X-rays
  Detection of multiple soft X-ray
Works also at pulsed sources
Dynamic range up to ~104 photons/pixel/s
Supported flux scales with frame rate
Saturations for long integration times
Challenging calibration and cumbersome data processing needs high throughput!

## Slide 19: PSI Detector portfolio
 | Microstrip 1D | Pixel 2D | 
 | 25/50 µm pitch | 75 µm pitch | 25 µm pitch
Single photon counting |  |  | 
Charge integrating |  |  | 
MATTERHORN
MYTHEN
GOTTHARD
MÖNCH1.0
JUNGFRAU2

## Slide 20: Silicon has limits

## Slide 21: Frontiers of Photon Detectors
Silicon
LGAD
High-Z
Low Energy (< 1 keV)
High Energy (> 40 keV)
n+
n-
Al
p+
Al
x x x x x x x x x x x x x
ASIC
Standard
n+
Al
p+
Al
x
x x
ASICS
SiO/N
Zhang, et al. JINST 2022; 17(11): C11011.

## Slide 22
Introduction to PSI
Science at PSI
User Facilities – SLS, SwissFEL, etc.
Photon Science Detector Group – How we support user facilities
High-Z Characterisation at PSI
Overview of High-Z materials
GaAs
CZT
High-Z Future & Applications
Photon Science
Medical Physics
Astronomy

## Slide 23: The Need For A New Material
SLS 2.0 (4th gen. synchrotron)
  Increased flux for photon energies > 20 keV
  Silicon (low-Z) inefficient in this range
High-Z
  Higher atomic number → Stops more hard X-rays
  Three varieties we have worked with: GaAs:Cr, CdTe, CdZnTe
  Thinner is better for charge sharing, but stops less

## Slide 24
High-Z Sensors
Courtesy of D. Greiffenberg

## Slide 25: Material Properties
Courtesy of D. Greiffenberg
 | Atomic Number, Z | Auto-fluorescence [keV] | Band Gap Energy (@RT) [eV] | e—h+ Pair Creation Energy [eV] | Resistivity [Ω∙cm] | Electron μ∙τ [cm2/V]
Silicon | 14 | 1.7 | 1.12 | 3.6 | < 104 | > 1
Cd0.9Zn0.1Te | 48, 30, 52 | 23, 27 | 1.57 | 4.6 | 1010 | 10-3..10-2
CdTe | 48, 52 | 23, 27 | 1.44 | 4.4 | 109 | 3 x 10-3
GaAs:Cr | 31, 33 | 10 | 1.43 | 4.2 | ~ 109 | < 5 x 10-4

## Slide 26: CZT

## Slide 27: CZT Variants
High-Flux (HF)
  Optimised for high flux environments (10-10 ph s-1 mm-2) in medical and security applications
  Increased hole lifetime
  Afterglow performance is also improved
CZT manufactured by Redlen Technologies
1.5 mm thick
2 x 2 cm2 single chip samples
Operated at -500 V (But rated for -900 V)
Two variants
By STFC:
B. Thomas, 2017 [7]
Spectroscopic (SP)
  The original (Standard) variety.
  Efficient for energies upto 120 keV
10 mm Thick
2 mm Thick
Thicker material tested by STFC

## Slide 28: Dark Frame
Lots of scratches in both variants
Increased ADU
From polishing on pixelated surface
No major defects in High-Flux
Spectroscopic has many defects by comparison
Significant elevation in ADU
But only one device of each variant tested
Average dark frame (raw ADU values) as integration time increases
5 μs baseline (Pedestal Correction)

## Slide 29: Leakage Current
Similar to dark frames
  But now we can quantify the low current
  Even in defect region < 8 pA per pixel
  All measurements at 15 °C
2-3 orders of magnitude improvement over GaAs
Better than silicon
Material | MPV Pixel Leakage Current
High-Flux CZT | ~ 0.18 pA
Standard CZT | ~ 0.49 pA
Silicon | ~ 2 pA
GaAs | ~ 190 pA [8]
-500 V Bias
15 °C

## Slide 30: Noise
Generally flat distribution across sensor area
  5 μs integration time
  Reasonably consistent across sensor material
Noisy pixels everywhere
  But generally corresponds to hot spots
  Otherwise, better than GaAs and comparable to silicon
Material | Average Pixel Noise
High-Flux CZT | ~ 0.28 keV
Standard CZT | ~ 0.29 keV
Silicon | ~ 0.3 keV [9]
GaAs | ~ 0.41 keV [8]
-500 V Bias
15 °C
5 μs Int. Time

## Slide 31: Am241 Spectra
Cluster Size (Virtual Pitch) | FWHM [keV] (Energy Resolution) | 
 | High-Flux | Spectroscopic
3x3 (225 μm) | 5.8% | 6.5%
5x5 (375 μm) | 8.2% | 9.0%
7x7 (525 μm) | 11.9% | 13.7%
9x9 (675 μm) | 16.7% | 20.0%
Similar spectra for High-Flux and Spectroscopic
  5x5 cluster required to collect all charge
  Slight shift in photopeak between 3x3 and 5x5
  Spectral resolution is the ratio of FWHM and peak position (59.5 keV)
  Comparable with GaAs (especially earlier generations) [8]
  Would likely be improved by a pixel-by-pixel calibration
Cluster Size | Photopeak Position
3x3 | 58.7 keV
5x5 | 59.5 keV*
7x7 | 59.9 keV
9x9 | 60.1 keV
  *Assumed to be Am241’s 59.5 keV photopeak
Cluster Size | Photopeak Position
3x3 | 58.7 keV
5x5 | 59.5 keV*
7x7 | 59.9 keV
9x9 | 60.2 keV
-500 V Bias
15 °C
5 μs Int. Time

## Slide 32: Hecht Scans
Mo (17.5 keV) is ideal for Hecht
Absorption depth only ~ 15 μm
Extract μe·τe by fitting with 7x7
Charge loss in 3x3 is not purely trapping → Charge sharing dominates for low sensor bias and small cluster sizes
μe·τe for High-Flux CZT is better than GaAs and CdTe (~ 0.4 x 10-3 cm2/V [8] and 1 x 10-3 cm2/V [10])
Cluster Size (Virtual Pitch) | μe·τe [x10-3 cm2/V] | 
 | High-Flux | Spectroscopic
3x3 (225 μm) | 9.7 ± 0.8 | 4.41 ± 0.13
5x5 (375 μm) | 9.6 ± 1.2 | 14.0 ± 0.6
7x7 (525 μm) | 10.0 ± 1.1 | 13.96 ± 0.6
9x9 (675 μm) | 11.9 ± 0.9 | 13.1 ± 0.7
Preliminary
Preliminary
15 °C
5 μs Int. Time
Mo Peak (17.5 keV)
No peak below 100V
No longer depleted?
μe·τe 3x3 is untrustworthy
(Curve dominated by charge loss)

## Slide 33: Charge Collection Efficiency
Charge sharing lowers 3x3 peak even at -500 V
7x7 is affected at -20 V
Only trapping contributes to the 9x9 peak (lower at -20 V compared to -500 V)
Spectroscopic can still collect a majority of charge (> 97% CCE) at -20 V!
(Where High-Flux needs over 100 V to start collecting charge)
Charge collection efficiency (CCE) at 500 V is over 99% for both variants
Charge loss due to charge sharing
Charge loss due to charge sharing
Cluster Size (Virtual Pitch) | Charge collection efficiency | 
 | Spectroscopic @ -20V | Spectroscopic @ -500 V
3x3 (225 μm) | 55.8% | 97.1%
5x5 (375 μm) | 86.6% | 99.4%
7x7 (525 μm) | 94.4% | 99.9%
9x9 (675 μm) | 97.0% | 99.9%
15 °C
5 μs Int. Time
Mo Peak (17.5 keV)

## Slide 34: GaAs:Cr
Results by Kirsty A. Paton

## Slide 35: Dark Current, Noise & Dynamic Range
 | Mean Dark Current (pA) | Mean Noise (rms e- E.N.C.)
TSU (2016) | 391.0 (σ = 16.9 %) | 115.93 (σ = 5.5 %)
TSU (2017) | 201.6 (σ = 17.3 %) | 101.65 (σ = 9.4 %)
TSU Dectris

## Slide 36: Mo Fluorescence Flat Field
5 µs integration time
Chiller set to 15°C
Bias – 300V

## Slide 37: Energy Resolution
2016 TSU - 6.9%
2017 TSU - 4.3%

## Slide 38: Hecht Measurements
2016 TSU - 1.83 × 10-4 cm2/V
2017 TSU - 4.73 × 10-4 cm2/V

## Slide 39: Interpolation with GaAs
JUNGFRAU
75 µm physical pixel pitch
Interpolation requires charge sharing on the order of a pixel pitch
Challenging for thinner Si sensors
MÖNCH
25 µm Physical Pixel Pitch
2.5 µm Interpolated Pixel
Swiss Flag 25 µm Wide, Cross 7µm Wide
S. Chiriotti et al., JInst, (2022)
Notes: 2 µm gold on 200 µm silicon sample fabricated at LXN, PSI

## Slide 40: High-Z Discussion

## Slide 41: Afterglow
0 photons
1 photons
2 photons
+1
+1
+1
+1
+1
+1
Time (many seconds)
0 photons
1 photons
2 photons
Afterglow
Auto-fluorescence
Noise
Time (many seconds)
0 photons
1× 102 photons
1× 104 photons
+100
+10000
+1
10 ms
10 ms
0 photons
1× 102 photons
1× 104 photons
10 ms
10 ms
0 photons
1× 102 photons
1× 104 photons
Reset
Integration window:
few μs
Reset
Integration window:
few μs
Integration window:
few μs
Synchrotron
XFEL
CZT has very minimal afterglow
  Would work at synchrotrons with photon counting
  or XFELs with resets
GaAs & CdTe suffer much more from this problem
  Photon counting might still be the way forward

## Slide 42: Which Material Wins?
CZT is the obvious choice
Leakage current/afterglow, signal stability all comparable with silicon
Should work well with charge integrating detectors at FELs
Yield (including defect rate) needs to be further studied)
< 1.5mm thick not currently possible
  Charge sharing results in larger pixel sizes
  Not suitable for photon counting unless pixels are grouped together in 2x2 150 μm super pixels
GaAs:Cr could be an alternative for photon counting
High dark current is less of a concern
Better commercial availability (full 500 μm thick modules)
Uniformity is still a concern however
Courtesy of D. Greiffenberg

## Slide 43
Introduction to PSI
Science at PSI
User Facilities – SLS, SwissFEL, etc.
Photon Science Detector Group – How we support user facilities
High-Z Characterisation at PSI
Overview of High-Z materials
GaAs
CZT
High-Z Future & Applications
Photon Science
Medical Physics
Astronomy

## Slide 44: New CZT Assemblies
CZT + JF1.2
Testing dynamic gain switching with electrons
Full CZT Module (2x Quads) + JF1.2
0.5M uniformity and pilot experiments at beamlines
Expected later this year
CZT + Matterhorn
Photon counting with high-Z
MH optimised for high count rates (20 MHz/Pixel)
Aims to mitigate afterglow with thresholding and fast shaping
Higher bias and thinner sensors are preferred (GaAs?)
Initially larger pixels for CZT (150 μm pitch)
First MH1.0 wafers expected at the end of the year
E. Fröjdh, 2023 [12]
CZT Quad
CZT Quad

## Slide 45: Further Characterisation Planned
Planer CZT Crystals
From Redlen (CA) and JX Advanced Metals (JP)
More in-depth characterisation
  Electrical properties (depletion)
  Decoupling of μ and τ
  Edge scanning with a laser
GaAs is considered for electron microscopy
Electrons are absorbed in their entirety (not MIP-like)
  But not before they scatter within the sensor
  Higher absorption efficiency is preferred to reduce track length
Short track lengths mean charge spread is dominated by diffusion
  Simpler to interpolate
  (For longer track lengths, interpolation would require machine learning)

## Slide 46: Medical Applications
WidePIX (Based on Medipix)
  1mm CdTe
  Multichannel “color” radiographs
  Correction required for “veiny” structure
Photon Counting Computed Tomography (PCCT)
  CdTe
  Offers better temporal discrimination between photons
CdTe is often favoured over CZT
  Cheaper and easier to make large area detectors
  CZT also relatively new to the industry
Z. Melníková, et al [13]
F. Maddaloni et al [15]
T. Flohr, et al [14]

## Slide 47: Medical Applications
SANTIS and EIGER from Dectris (PSI spinoff)
  GaAs (~500 μm thick)
  High resolution/sensitivity phase contrast imaging
  Small pixel size and large sensitive area is desirable
  Can provide the same SNR but with a higher resolution (compared to scintillation-based detectors)
CZT “Gamma Camera” for SPECT
  Imaging with radioisotopes: 99mTc and 123I
  Better contrast and energy resolution than previous CdTe detectors
J. Scholz, et al [16]
T. Ito, et al [17]

## Slide 48: Astronomy Applications
Many CZT detectors used in space/upper atmosphere
Low leakage current of CZT is ideal when power requirements and cooling are key
    Rules out GaAs (Also has a lower absorption efficiency)
  CdTe performs worse than CZT in most metrics
    Veiny/defect structure of CdTe is not ideal
    Can also afford CZT’s smaller active area
D. Shy, et al [18]

## Slide 49: Astronomy Applications
India’s AstroSat launched in 2015
CZT based detector
Measured the polarisation of gamma ray bursts [19]
TERI is another example launched in 2025
  Goal is test CZT’s long term viability and space [18]
  Also aimed at the MeV range, key for the study of astrophysical jets
  15 mm CZT from Redlen [21]
  Up to 50 mm has also been characterised. 25 mm shows < 1% energy resolution at 662 keV [22]
D. Shy, et al [18]
ESA [20]

## Slide 50: Outlook and Summary

## Slide 51: Summary
PSI continues to develop cutting edge detectors for photon science applications
Jungfrau & Matterhorn represent the latest in charge-integrating and photon-counting technology
High-Z has a bright and brilliant future in photon science
Already well established for high energy photon detection in medical physics and astronomy
Thin, large area sensors are key and more work is still be to done!
Characterisation work continues at PSI, aiming to unlock high-Z’s full potential
Full module pilot experiments expected soon to demonstrated feasibility to users

## Slide 52: Thank You!
K.A.P. has received funding from the EU’s Horizon 2020 MSCA grant. No. 884104 (PSI-FELLOW-III-3i)
V. Kedych, M. Müller, P. Sieberer, K. Ferjaoui, M. Brückner, E. Fröjdh, J. Heymes, C. Lopez-Cuenca,
K. Moustakas, C. Ruder, J. Mulvey, S. Silletta, V. Hinger, A. Mozzanica, M. Carulla, X. Xie,
R. Dinapoli, A. Bergamaschi, S. Li, D. Thattil, J. Zhang, A. Mazzoleni, V. Gautam, D. Mezza, T. King, B. Schmitt.
Not pictured: S. Ebner, D. Greiffenberg, S. Hasanaj, K.A. Paton, C. Posada Soto, M Hürst.

## Slide 53: References
[1] M. Böge, ‘First Operation of the Swiss Light Source’, Proceedings of EPAC 2002, Paris, France, 2002
[2] E P. Mithcell, et al, “Demystifying the Synchrotron Trip: A First-Time User’s Guide.” Structure, 7, no. 5, May 1999, R111–R121. https://doi.org/10.1016/S0969-2126(99)80063-X
[3] M. Böge, et al, ‘Commissioning and First Operation of SLS 2.0, the New Swiss Light Source’, in Proceedings of IBIC 2025, paper MOCI01
[4] Paul Scherrer Institute. "SwissFEL Accelerator." SwissFEL – Swiss X-ray Free Electron Laser. Accessed October 6, 2025. https://www.psi.ch/en/swissfel/accelerator
[5] P. G. Jung, et al, "Microdome-Grooved Gd₂O₂S:Tb Scintillator for Flexible and High Resolution Digital Radiography.“ Optics Express, 18, no. 14 (2010): 14850–14858. https://doi.org/10.1364/OE.18.014850
[6] A. Mozzanica, et al, ‘Prototype characterization of the JUNGFRAU pixel detector for SwissFEL’, JINST, 9, (2014), C05010
[7] B. Thomas et al, ‘Characterisation of Redlen high-flux CdZnTe’, 2017, JINST, 12, C12045
[8] K. A. Paton, ‘Studying a new variety of GaAs:Cr with the JUNGFRAU charge integrating detector’, HiZPAD Workshop, London, United Kingdom, 2025
[9] A. Mozzanica et al, ‘Characterization results of the JUNGFRAU full scale readout ASIC’, 2016, JINST, 11, C02047
[10] M. Veale, “Solid-State Radiation Detectors: Chapter 3” (1st ed.), 2015, CRC Press
[11] D Greiffenberg et al, Energy resolution and transport properties of CdTe Timepix-Assemblies 2011 JINST 6 C01058
[12] E. Fröjdh, et al, ‘A look at single photon counting detectors for SLS2.0’, The 13th International Conference on Position Sensitive Detectors (PSD13), Oxford, United Kingdom, 2023

## Slide 54: References
[13] Z. Melníková, et al, “Large-Area CdTe Detector for X-ray Imaging and Overcoming Crystal Imperfections in CdTe Sensors”, International Workshop On Radiation Imaging Detectors (iWorid), Bratislava, Slovakia, 2025
[14] T. Flohr, et al, “Photon-counting CT review”, Physica Medica, 79, 126–136, 2020
[15] F. Maddaloni et al, “Spatial resolution characteristics of a clinical Photon-Counting Computed Tomography scanner”, International Workshop On Radiation Imaging Detectors (iWorid), Bratislava, Slovakia, 2025
[16] J. Scholz, et al, “Biomedical x-ray imaging with a GaAs photon-counting detector”, APL Photonics, 5(10), 106108, 2020, https://doi.org/10.1063/5.0023819
[17] T. Ito, et al, “Experimental evaluation of the GE Discovery 870 CZT clinical SPECT/CT system”, Journal of Applied Clinical Medical Physics, 22(1), 1–13, 2021, https://doi.org/10.1002/acm2.13158
[18] D. Shy, et al, “Development of the cadmium zinc TElluride Radiation Imager (TERI)”. arXiv preprint arXiv:2408.04559, 2024, https://doi.org/10.48550/arXiv.2408.04559
[19] T. Chattopadhyay, et al, “Hard X-Ray Polarization Catalog for a Five-year Sample of Gamma-Ray Bursts Using AstroSat CZT Imager”, The Astrophysical Journal, 936:12 (13pp), 2022 September 1
[20] European Space Agency, 2023, Gamma-ray Burst Illustration, March 28, 2023, https://www.esa.int/ESA_Multimedia/Images/2023/03/Gamma-ray_burst_illustration
[21] Y. Zhu, Z. He, “Performance of Larger-Volume 40x40x10 and 40x40x15 mm3 CdZnTe Detectors”, IEEE Transactions on Nuclear Science, VOL. 68, NO. 2, February 2021
[22] A. E. Bolotnikov et al, “CdZnTe position-sensitive drift detectors with thicknesses up to 5 cm”, Appl. Phys. Lett. 108.093504, 2016

## Slide 55: Extra Slides

## Slide 56: Photon Science Applications
Fourier ptychography
Wakonig et al.
Full field fluorescence
imaging
Marone et al.
Macromolecular Crystallography Leonarski et al.
Operando diffraction Hocine et al.
Magnetic contrast ptychography – Butcher et al.
Atomic Dynamics
Chapman et al.
Ptychography
Holler et al.
Tender X-ray
Emission Spectrometer
Huthwelker et al.
Hard X-rays
Soft
X-rays
Photon-starved
High flux
High Speed
Courtesy of A. Bergamaschi

## Slide 57: More Material Properties
Courtesy of D. Greiffenberg
 | Atomic Number, Z | Auto-fluorescence [keV] | Density [g/cm3] | Band Gap Energy (@RT) [eV] | e—h+ Pair Creation Energy [eV] | Resistivity [Ω∙cm] | Electron μ∙τ [cm2/V] | Hole μ∙τ [cm2/V] | Afterglow Performance
Silicon | 14 | 1.7 | 2.33 | 1.12 | 3.6 | < 104 | > 1 | > 1 | Near Perfect
Cd0.9Zn0.1Te | 48, 30, 52 | 23, 27 | 5.78 | 1.57 | 4.6 | 1010 | 10-3..10-2 | 10-5 | Pretty Good
CdTe | 48, 52 | 23, 27 | 5.85 | 1.44 | 4.4 | 109 | 3 x 10-3 | 2 x 10 -4 | Not Ideal
GaAs:Cr | 31, 33 | 10 | 5.32 | 1.43 | 4.2 | ~ 109 | < 5 x 10-4 | ~ 5 x 10-7 | Not Ideal

## Slide 58: I-V Characteristics
  Surprise:
  Each contact combintation (Pt/CdZnTe/Au or Au/CdZnTe/Au) is recifying!
  In CdTe, Pt/CdTe/Au or Au/CdTe/Au make quasi-ohmic contacts.
  Comment on yield:
  5x Assemblies received (JF1.1 + 3x CZT High Flux/2x CZT spectroscopic)
  2x Assemblies unusable (Both put JF1.1 in saturation in G2)
  1x CZT High Flux died after applying -900 V (according to specs)
  2x Assemblies alive (m401 CZT Spectroscopic + m343 CZT High Flux)
-500 V
-900 V
-500 V
-900 V
Spectroscopic CZT (JF1.1 m401)
High Flux CZT (JF1.1 m343)

## Slide 59: Leakage Current
Integration Time (IT) Scan per pixel
  Measure the raw ADC values (ADU) a different integration times.
Apply gain calibration
  ADU → keV → e- → Q
  Average e-h pair creation energy: ~ 4.43 eV
  100 ADU (~ 3 keV) at 475 μs IT
Fit linear region (avoiding initial dip)
  Extract the leakage current (dQ/dt) from the gradient (E/t)
Note: Initial fall in ADU before 25 μs where it becomes linear
  Believed to be dominance of electronic effects
  Usually operated at a fixed integration time
Target (e.g. Zr)
X-ray Generator
Detector
Characteristic X-rays (e.g. 15.8 keV)
-500 V Bias
15 °C

## Slide 60: Scratches
Microscope scan
  Different Redlen CZT sensor (96x96 pixels)
  Scratches visible over the entire pixel array
  More scratches visible when zoomed in
# clusters map
  Shows many more scratches
  More clusters found along scratches

## Slide 61: # Clusters From Calibration
Healthy # of clusters
Details more apparent
  Defect structure in SP
  Many more scratches than in IT scan
Reduced # due to:
  Scratch on backside
  HV Solder
HV Solder

## Slide 62: Energy Resolution
Energy Resolution with 5x5 clusters
High-Flux is better than Spectroscopic
Spectroscopic is also less uniform
Corners likely due to low statistics
-500 V Bias
15 °C
5 μs Int. Time

## Slide 63: Interpolation with CZT
75 µm
25 µm
15 µm
5 µm
Spectroscopic CZT
Courtesy of D. Greiffenberg

## Slide 64: A Minor Blemish…

## Slide 65: IT Scan
Courtesy of K. A. Paton

## Slide 66: Hecht Measurements
2016 TSU - 1.83 × 10-4 cm2/V
2017 TSU - 4.73 × 10-4 cm2/V

## Slide 67: Interpolation with GaAs
Physical Pixel Pitch 75 µm
Physical Pixel Pitch 75 µm
Physical 75µm Pixel Pitch
Interpolated 25µm Pixel Pitch

## Slide 68: Imaging
Hamamatsu microfocused X-ray tube
40 kV peak voltage
30 uA current
Flat field corrected (~ 2 h for each dataset)

## Slide 69: Imaging
See S. Cartier et al, JInst (2014) for more details

## Slide 70: Results Summary
 | Leakage Current Per Pixel [pA] | Noise [keV] | μe∙τe [x10-3 cm2/V] | Energy Resolution at 60 keV [%]
Silicon | ~ 2 | 0.23 | > 1 ** | Can’t detect 60 keV
CZT High Flux | 0.18 | 0.28 | 10 | 5.8
CZT Spectroscopic | 0.49 | 0.29 | 14 | 6.5
CdTe Schottkey | ~ 0.76 | ~ 0.5 | 1.9 * | 5.6 *
GaAs:Cr | 190 | 0.41 | 0.39 | 4.68
* Measured with Time-Pix rather than JUNGFRAU [11]
** For both electrons and holes

## Slide 71: Papers
GaAs:Cr
D. Greiffenberg et al., JInst, (2017)
D. Greiffenberg et al., Sensors (2021)
CdTe
M . Markus et al., JInst (2022)
Signal Stability
D. Greiffenberg et al. Jinst (2025)
Afterglow
Coming SoonTM
CZT
Coming SoonTM

## Slide 72
1st Joint X-ray Sensor Workshop From Soft to Hard X-rays
ETH Zurich (Switzerland) 12.-16. January 2026
Bridging Science and Detection Technology for Soft X-ray Applications
High-Z Pixel Array Detectors for Photon Science
Webpage: https://indico.psi.ch/e/XRaySensors2026
Notes: Change QR code