## Slide 1: Positron emission tomography imaging harnessing polarization correlations Mihael Makek1 T. Bokulić1, D. Grošev2, A. M. Kožuljević3, S. Parashari4, L. Pavelić3, M. Žuvić2 1Department of Physics, Faculty of Science, University of Zagreb, Croatia 2University Hospital Centre Zagreb, Croatia 3Institute for Medical Research and Occupational Health, Zagreb, Croatia 4Instituto de Física Corpuscular (IFIC), Valencia, Spain
Exotic atoms: fundamental aspects, applications and advances in radiation detectors Zagreb, Croatia, Dec 3-5 2025

## Slide 2: Standard PET detector
Injection and metabolisation of positron labeled radiopharmaceutical
Positron annihilation:
e+e- -> g g
The pair of gs identified based on:
  coincidence time
  Energy of each 511 keV
  Back-to-back geometry assumed
M. Makek
e++e- -> g+g
RING OF DETECTORS
Positron-labeled radiopharmaceutical
Positron annihilation in tissue

## Slide 3: Standard PET detector
Injection and metabolisation of positron labeled radiopharmaceutical
Positron annihilation:
e+e- -> g g
The pair of gs identified based on:
  coincidence time
  Energy of each 511 keV
  Back-to-back geometry assumed
LORs --> reconstructed image
M. Makek
RING OF DETECTORS

## Slide 4: Background in PET
Backgrounds mostly due to:
  Compton scatter in patient's body
M. Makek
g’
g
Scatter

## Slide 5: Background in PET
Backgrounds mostly due to:
  Compton scatter in patient's body
  Random coincidence
  -> reduced image contrast
M. Makek
Random coincidence

## Slide 6
Additional information available: Polarization of annihilation quanta
M. Makek

## Slide 7: Polarization of annihilation quanta
In positron annihilation back-to-back gammas generated
They are quantum-entangled and polarized in orthogonal directions [P. Dirac, Proc. Camb. Philos. Soc. 26 (1930) 361; R. Hanna, Nature 162 (1948) 332]
How to measure polarization correlations?
Use of polarization correlations in PET imaging?
Cost-efficient implementation?
g (511 keV)
g (511 keV)
M. Makek

## Slide 8: How to measure polarization correlations
When both linarly polarized annihilation photons undergo Compton scattering:
Polarimeric modulation:
[Pryce and Ward, Nature 160 435; Snyder, Pasternak, Hornbostel, Phys. Rev. 73 (1948) 440-8]
with:
M. Makek

## Slide 9: How to measure polarization correlations
When both linarly polarized annihilation photons undergo Compton scattering:
Polarimeric modulation:
Δφ (deg)
a.u.
[Pryce and Ward, Nature 160 435; Snyder, Pasternak, Hornbostel, Phys. Rev. 73 (1948) 440-8]
with:
M. Makek

## Slide 10: Polarization correlations in PET
Potential of polarization correlations for S/B improvement in PET?
  The polarization correlations are expected only for signal pairs
  Discriminate true coincidences from randoms
Δφ (deg)
a.u.
g
g
g
g
Scatter
Random conicidences
Standard event ID in PET: - back to back, coincidence
- 511 keV
Additional event ID in PET:
[Kuncic et al., NIM A648 (2011), McNamara et al., Phys. Med. Biol 59 (2014), M. Toghyani et al., Phys. Med. Biol 61 5803(2016)]
M. Makek

## Slide 11: How to measure gamma polarization in PET?
Modern PET detectors: - highly segmented, - large axial coverage
M. Makek
Typical Compton polarimeters:
  2 sensitive layers: 1st for measuring the recoil electron, 2nd for the scattered gamma
-> Measurement of polarization viable only using single-layer detectors​
A PET system based on 2-layer detectors would dramaticaly increase the cost of the apparatus

## Slide 12
Development of PET demonstrator with polarization measurement capability
M. Makek

## Slide 13: Detector modules: single-layer polarimeters
M. Makek
GAGG matrix (2.9x2.9x20 mm3) Single-sided readout by SiPM!
S. Parashari et al. (2022), NIMA, 1040:167186
Compton event ID
2 pixels fired
Animation by E. Piškorić

## Slide 14
M. Makek
GAGG matrix (2.9x2.9x20 mm3) Single-sided readout by SiPM!
S. Parashari et al. (2022), NIMA, 1040:167186
Detector modules: single-layer polarimeters
Compton event ID
2 pixels fired

## Slide 15
M. Makek
GAGG matrix (2.9x2.9x20 mm3) Single-sided readout by SiPM!
S. Parashari et al. (2022), NIMA, 1040:167186
Detector modules: single-layer polarimeters

## Slide 16: The PET demonstrator
Four modules with 16 x 16 scintillator pixels
  2 modules 3.2 mm pitch, GAGG
  2 modules 2.2 mm pitch, LYSO
Readout by TOFPET2
Annular structure allowing diameters 420 – 700 mm
Rotation around the scanner axis to emulate the full ring
EXOTIC ATOMS ZAGREB 2025
M. Makek
Tests with high activity sourced at University Hospital Centre Zagreb

## Slide 17
PET imaging with polarization-correlated events
M. Makek

## Slide 18: The spatial resolution
Line sources Ge-68 ~45 MBq each, in epoxy phantom, 2 cm distance
Ring diameter: 430 mm; GAGG-3mm modules
1 pixel events, Epx=511 keV
FWHM ~2.5 mm
M. Makek
Pol.-cor. Events, |f1-f2|=90°±20°
FWHM ~5.1 mm

## Slide 19: The spatial resolution
Line sources Ge-68 ~45 MBq each, in epoxy phantom, 2 cm distance
Ring diameter: 430 mm; GAGG-3mm modules
1 pixel events, Epx=511 keV
FWHM ~2.5 mm
M. Makek
Pol.-cor. Events, |f1-f2|=90°±20°
FWHM ~5.1 mm
LOR ambiguity

## Slide 20: The spatial resolution
Line sources Ge-68 ~45 MBq each, in epoxy phantom, 2 cm distance
Ring diameter: 430 mm; GAGG-3mm modules
1 pixel events, Epx=511 keV
M. Makek
FWHM ~5.1 mm
LOR ambiguity
Pol.-cor. Events, |f1-f2|=90°±20°
Pol-cor events
1-px events
Spatial resolution vs. LOR gain

## Slide 21: Random background rejection
S/Brand from coincidence time spectrum, 5 ns window
EXOTIC ATOMS ZAGREB 2025
M. Makek
Single pixel events
A. M. Kožuljević, PhD Thesis, University of Zagreb, 2025
Pol-cor events, 72°<q<90°, 70°<|Df|<110°
Pol-cor events, 72°<q<90°, -20°<|Df|<20°

## Slide 22: Random background rejection
EXOTIC ATOMS ZAGREB 2025
M. Makek
Pol-cor events, 72°<q<90°, 70°<|Df|<110°
Pol-cor events, 72°<q<90°, -20°<|Df|<20°
Single pixel events
NEMA-NU4, Ga-68, 400 MBq initial activity
SBR improvement up to 20% for Pol.-Cor. events
Pol.cor events |Df|=90°±20° Single pixel events Pol.cor events |Df|=0°±20°
[ArXiv: 2510.11504]

## Slide 23: Random background rejection
EXOTIC ATOMS ZAGREB 2025
M. Makek
NEMA-NU4, Ga-68, 400 MBq initial activity
SBR improvement up to 20% for Pol.-Cor. events
Pol.cor events |Df|=90°±20° Single pixel events Pol.cor events |Df|=0°±20°
[ArXiv: 2510.11504]
[ArXiv: 2510.11504]
Azimuthal modulation (m) - SBR estimator

## Slide 24: Standard vs. Pol-Cor Image
1-Pixel events
Pol. correl. events
Ga-68
400 MBq
Ring diameter: 420 mm; GAGG-3mm modules
EXOTIC ATOMS ZAGREB 2025
M. Makek

## Slide 25: Standard and Pol-Cor PET synergy
EXOTIC ATOMS ZAGREB 2025
M. Makek
Single-pixel (standard) + Pol-Cor(~10% extra) data
Spatial resolution

## Slide 26: Standard and Pol-Cor PET synergy
EXOTIC ATOMS ZAGREB 2025
M. Makek
10% data increase, resolution and image quality preserved
NEMA NU-4 Phantom
Single-pixel (standard) + Pol-Cor(10% extra) data

## Slide 27: Further image improvement?
EXOTIC ATOMS ZAGREB 2025
M. Makek
Merged image contrast vs. pol-cor. image resolution
Pol-cor events
1-px events
Pol-cor events
1-px events
Merged image SNR vs. LOR gain
Contrast improvement observed with polarization -correlated event resolutions < 4 mm
SNR improvement observed with LOR gain > 8%
-> Fine-balance between contrast and sensitivity with current methods

## Slide 28: Conclusions
PET using polarization-correlated events
First imaging with clinically relevant activities
Improved random suppression
Additional signal to improve sensitivity
Balance contrast vs. sensitivity
M. Makek

## Slide 29: ACKNOWLEDGEMENTS
This work was supported by Croatian Science Foundation under the project IP-2022-10-3878 and in part by the “Research Cooperability” Program of the Croatian Science Foundation, funded by the European Union from the European Social Fund under the Operational Programme Efficient Human Resources 2014–2020, grant number PZS-2019-02-5829