## Slide 1: Hybrid system
2026 CAP CONGRESS, June 23rd
Radon-222 trapping with liquid nitrogen

## Slide 2: Outline
Motivation
Commissioning
System description
Calibration step
Characterization
Conclusion

## Slide 3: Motivation
Gas assay: (low pressure)
Rn trap : activated carbon as sole material
Excellent adsorption capacity at ⇣ T∘
Hybrid system
Increase sensitivity of radon measurement in large volume systems: PICO (PV 10 m3)
Rn trap : Bronze wool + Activated carbon
Water trap absorb the water before it reaches the charcoal
Wool: Not prone to blockage + limited adsorption of O2 and N2
Wool: To absorb residual moisture upstream
Wool: Allowing the system to breathe + limits the drop in P & Q.
Prone to clogging at ⇣ T∘  moisture, O2 & N2 condensation

## Slide 4: Commissioning
Charcoal:
Purification & Cleaning
Nitric acid (5 % HNO3)  leach out and remove contaminants (238-U & 226-Ra)
Increase surface roughness and the number of pores
Rinsed with UPW
Bake  150°C for 24 to 48 hours, with a nitrogen flow
Materials: Pipe, Fittings, Flexible tube & Wool
USC: nuclear radioactive decontamination solution
Methanol cleaning : valve, ss & bronze wool
Passivation and pickling

## Slide 5: System description
Water trap (short U-shaped ½” pipe)
SS wool
3/16" AISI 316 stainless steel bearing balls
3X FVCR Bellows sealed valves
2X VCR Particle filter, 0.5 µ, 316 SS
11.22-inch, Type T Thermocouple, Probe
Carbon trap (long U-shapped ½” pipe)
Carbo-Act International (0.2mBq/kg)
Bronze wool

## Slide 6: Calibration step
Rn emanation (exhalation)
  Radon atoms generated by the decay of radium in a material and escape through its interstices, mainly due to recoil effects
Radon source : # of Rn when reaching secular equilibrium
Rn capture by the trap system from the calibration source

## Slide 7: Calibration & Counting
Radon source
  Use to calibrate the hybrid system  determine trapping efficiency
  Detector (component) : Counting the Rn rate release (radon assay)
SNOLAB radon board
  Contains two traps
  Calibrated : known efficiency
  Not designed to handle large volumes of N2  efficiency decreases

## Slide 8: Characterization
Moisture :
  Reduces the dynamic adsorption coefficient
  Can take the places of Rn on the microscope pores of the adsorbent.
  Increase the probability that Rn escapes the trap
  LC  reduces & distorts the counting efficiency.
    Can condensate on scintillating material

## Slide 9: Characterization
Dynamic Adsorption (Ka)
  Is the retention (breakthrough) time
  is the mass of the absorbent
  Increases exponentially as the temperature decreases
Wang et al., Applied Radiation and Isotopes 191 (2023) 110564

## Slide 10: Characterization
Retention (Breakthrough) time 
  Designates the period during which carbon reaches saturation and can no longer reliably retain radon.
  The breakthrough time must always be greater than the assay (exposure) time.

## Slide 11: Characterization
Understanding the behavior of the system under vacuum
  Flow rate on T, P
  P and flow rate drop
  Homogeneity of T in the charcoal column
    Thermocouple directly in the carbon
    Cryofluid level in the dewar
    T around the probe region: internal material vs. external fluid
Inside
Outside

## Slide 12: Internal Temperature
Cryogenic fluid level in dewar dependent
T of the external fluid is not the same as that of the internal system.
Without heating, the temperature of liquid nitrogen (LN2) would have to be reached over a long period.
  However, the analysis is carried out within a limited time frame

## Slide 13: Flow rate & Pressure drop
amplify with flow rate
Should be monitored during an assay : assess clogging.
Trapping efficiency
  more time to interact and adhere with adsorbent material

## Slide 14: Conclusion
Has been done
  Hybrid system to trap Rn
  No blockage when LN2 is used as a cryo-fluid
  Characterize : P & T change in the system with the input flow
  Improve procedure, calibrate initial T as an example
Need to be done
  Add an CO2 trap
    ↑ trapping efficiency: competes for the same adsorption sites
    Prevents and reduces cryogenic blockages
  Run thermal model to have a better understanding
    Compared with recorded data: T, P & Q

## Slide 15: Annex
Reynolds
   laminar flow
Laminar flow
  Uniform
  Predictable residence time

## Slide 16: Annex
Parameters :
  𝜆 (activity=N𝜆 Bq or disintegration / day)
  LC in count alpha
  Please note : the expression in (parentheses) is the number of radon atoms

## Slide 17: Annex
Preliminary result: effectiveness of the activated carbon trapping system

## Slide 18: Annex