## Slide 1
CRIS Collaboration meeting 2026
Osama Ahmad
9 months evaluation
Field Ionization Unit (FIU) @ CRIS

## Slide 2: Field Ionization
Department of Physics and Astronomy (IKS)
Hyperfine Structure (HFS) :
F1
F2
F3
F4
Rydberg
IP
AI
Non-res
01
Reducing the background contributions from:
Collisional ionization
Laser related background, especially from high power non-resonant step
FIU via Rydberg state makes high power laser obsolete
Electric field ionization of Rydberg states in a collinear geometry
A. Vernon et al., Sci. Rep. 10, 12306 (2020).

## Slide 3: Timeline
Department of Physics and Astronomy (IKS)
02
Installing FIU at CRIS
FIU commissioning I
March 2024
Understood the dips
April 2024
October 2024
High-lying rydbergs (n=24) affected strongly by Stark shift
Atoms at locations with strong field along beamline not rydberg excited  field ionization not effectful
Resolved with lower CEC deflector potentials and lower rydbergs (n=18, K)
First online implementation
November 2024
New Rydberg series observed in Fr
Which series? Ask Pierre
Improvements in FIU
February 2025

## Slide 4: Previous design (1.0)
Department of Physics and Astronomy (IKS)
03
A. Vernon et al., Sci. Rep. 10, 12306 (2020).
Acceleration lenses are redundant with our bunch width.
Limitation in vacuum in IR
Grids shortening issues.
Deflectors
Grids
Acceleration lenses

## Slide 5: Problems
Department of Physics and Astronomy (IKS)
GRIDS SHORTENING
ASSEMBLY
04

## Slide 6: New proposed design
Field Ionizer Schematics.Rev. Sci. Instrum. 93, 065004 (2022); doi: 10.1063/5.0061736 Department of Physics and Astronomy (IKS)
Based on GUNILLA setup in Gothenberg.
Instead of GRIDS, using a set of parallel coaxial circular plates.
First plate is at ground potential, ensuring a field free interaction region.
All the other plates can be individually biased.
05
Designing the deflectors like plates, similar to the CEC deflectors (instead of rings with PCB)
Position for new deflectors

## Slide 7: New design (2.0)
Figure thanks to C.M Fajardo - Zambrano Department of Physics and Astronomy (IKS)
06
Assembled and installed in the beamline in the first week of March
Tests performed during ISCOOL commissioning.
No field ions even went up to 2500V on PLATE 2.
Bender electrodes did work
The diameter of the hole was 28 mm, electric field gradient not strong enough to ionize.
Back to grids model with copper plates (instead of PCBs) with thicker wire.
Currently using only two grids (three grids cannot fit)
The new deflectors (plates) installed.
Commissioning with offline source (K) done in May before Ar beamtime.
ALTERNATE OPTION (2.1)

## Slide 8: FIU vs 1064
Department of Physics and Astronomy (IKS)
07

## Slide 9: FIU vs 1064
Department of Physics and Astronomy (IKS)
08

## Slide 10: Outlook
Department of Physics and Astronomy (IKS)
Signal/Bck needs to be looked into.
Depending upon the potential applied on the grids, bender needs to be tuned to recover the full beam.
Ionization of rydbergs 300 cm-1 below the IP with FIU.
Was planned to be used for IP measurement in RaF and also for AcF.
Versatile tool to search for Rydberg’s and atomic studies.
Plans for LS3 offline campaign.
09

## Slide 11
THANK YOU