## Slide 1
XGRADE 2026
CA24131-ENRICH
Advances in X-ray & Gamma Ray Detectors and Applications
GridPix Detector as an Advanced Alternative to IXPE GPD: Experimental Characterization and Validation with GEANT4 Simulations.
Presented by,
Saba Imtiaz , PHD Scholar, INAF-IAPS, Roma.
Co-Authors: Dr Paolo Soffitta, Dr Fabio Muleri, Dr Alessandro Di Marco, Dr Dawoon kim, Dr Markus Gruber, Dr Vladislavs Plesanovs
GridPix

## Slide 2
Introduction to IXPE
3
1
The Gridpix Detector and the InGrid
4
Experimental Study and Results
Conclusions and Future Directions
5
Overview of the IXPE Mission.
Scientific Motivation for X ray Polarimetry.
GridPix Detector Architecture and ASIC Integration.
Operational Principle of Detector.
Comparative Analysis: GridPix vs GPD.
Overview of Experimental Setup.
Results Obtained from Experimental Data.
Results Obtained from Simulated Data.
Comparative Analysis of Experimental and Simulated results.
Summary of Key Findings.
Potential Improvements and Future Research Directions.
2
How Photoelectric Polarimeters Work
Gas pixel detector (GPD) Architecture.
How Gas Pixel Detector Work
Limitations of GPD ASIC
PRESENTATION ROADMAP
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA

## Slide 3
The Imaging X-ray Polarimetry Explorer (IXPE), launched on 9 December 2021, carries three telescopes equipped with detectors capable of measuring the polarization of X-rays from astrophysical sources.
The detectors were developed by a team of INFN and INAF researchers.
Our Work
IXPE during environmental testing
at Ball Aerospace
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
INTRODUCTION
ABOUT SPACE OBSERVATORY...
Weisskopf, M. C., Soffitta, P., Baldini, L., Ramsey, B. D., O’Dell, S. L., Romani, R. W., ... & Thomas, N. E. (2022).
3

## Slide 4
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
INTRODUCTION
ABOUT SPACE OBSERVATORY...

## Slide 5
Why We Do Polarimetry?
X-ray polarimetry is performed to measure the polarization of X-rays emitted by astrophysical sources.
It provide unique information about the geometry, magnetic fields, and emission mechanisms in extreme environments such as black holes, neutron stars, and supernova remnants.
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
WHY POLARIMETRY?
EXAMPLE TARGETS STUDIED BY IXPE
Magnetic Field Geometry, overlaid to the Intensity.
As Measured by IXPE for the Crab Nebula.
Bucciantini, N., Romani, R. W., Xie, F., & Wong, J. (2024).
How are Particles accelerated in Blazar jets?
Kirk, J. G., Rieger, F. M., & Mastichiadis, A. (1998). Particle acceleration and synchrotron emission in blazar jets. arXiv preprint astro-ph/9801265.

## Slide 6
↑
First X-ray polarimeters flown on sounding rockets.
Angel et al. 1969 Wolff et al. 1970
Novick 1970
Tindo et al. 1970
Novick et al. 1972
Dark age: many X-ray polarimeters developed, none flown…
e.g. Kaaret et al. 1989
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
INTRODUCTION of IXPE
HISTORY OF X-RAY POLARIMETRY
1969 - 1972 1976
...
GPD
Costa et al. 2001 Bellazzini et al. 2006
2000’ 2017
2021
OSO-8: First significative detection of X- ray polarization from the Crab Nebula.
Weisskopf et al. 1976 Weisskopf et al. 1978
↓
↓
↑
↓
↑
↑

## Slide 7
Why We Do Polarimetry?
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
ONBOARD DETECTOR
THE GAS PIXEL DETECTOR (GPD)
Concept from Costa et al. (2001)
Primary Ionization
Electron Drift
Gas Electron Multiplier
Readout Chip
Diffusion
1
2
3
4
The incident photon enters
in the gas cell through a
Berylium Window.
Incident photon emit photoelectron in gas. These photoelectron travel in gas medium.
Electrons reach out at GEM and then get multiplied.
At the end photoelectron track record at ASIC
DME

## Slide 8
Why We Do Polarimetry?
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
ONBOARD Detector
Principle of the Gas Pixel Detector (GPD)
GPD Assembly
Muleri, Fabio, et al.
A new design for the gas pixel detector." Space Telescopes and Instrumentation 2012: Ultraviolet to Gamma Ray. Vol. 8443. SPIE, 2012.
Photoelectron emission follows a cos(2φ) modulation, showing a clear anisotropic distribution for polarized photon.
Photoelectron emission is uniform (flat distribution) with no preferred direction for unpolarized photon.
Photoelectron Emission and Angular Distribution for Polarized X-rays

## Slide 9
Why We Do Polarimetry?
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
DETECTOR LIMITATIONS
Factors Affecting the Detector Performance
GPD operates effectively only in a limited X-ray energy range (2–8 keV).
Detector introduces dead time(1.1ms), reducing efficiency at high event rates.
The deposition of ions onto the dielectric GEM leads towards the decrease in gas gain.
The ASIC of Gass pixel detector offers two-dimensional photoelectron track.
Gas pixel detector has limited spatial resolution because of charge diffusion during drift.
GPD has relatively low energy resolution due to fluctuations in gas ionization and amplification process.
Its effective area is also limited because the gas absorption probability decreases with increasing photon energy.
Bellazzini, Ronaldo, et al. "Photoelectric X-ray polarimetry with gas pixel detectors." Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment 720 (2013): 173-177.

## Slide 10
A
Next-Generation
Approach
For
X-ray
Polarimetry
GRIDPIX
DETECTOR

## Slide 11
Why We Do Polarimetry?
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
GRIDPIX DETECTOR
DETECTOR INTRODUCTION AND ITS ASIC
The GridPix detector combines a micromesh amplification grid with a Timepix3 pixel readout chip
Kaminski, J., et al. "GridPix detectors–introduction and applications." Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment 845 (2017): 233-235.
SEM Images of GridPix Strcture:
Top images explain the pixel readout and micromesh Grid with zoom in.
Bottom Images show how the mesh is supported by pillars and aligned with the pixels.

## Slide 12
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
GRIDPIX DETECTOR
Comparison of Detector Structures: GridPix vs GPD
2.Drift Region
1.
3.Micromesh Grid
4.
6.TimePix
5.
7.Anode
1.Cathode
2.Drift Region
3.
4.Induction Gap
5.
The micromegas grid supported by pillars in the GridPix detector creates a very small and uniform amplification gap directly above each pixel, enabling precise single-electron detection and better spatial resolution compared to the GEM-based amplification in the Gas Pixel Detector.
Gas Pixel Detector
Ligtenberg, C., Heijhoff, K., Bilevych, Y., Desch, K., van der Graaf, H., Hartjes, F., ... & Timmermans, J. (2018). Performance of a GridPix detector based on the Timepix3 chip. Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, 908, 18-23.
GridPix Detector
Ionization electrons produced in the gas drift region move toward the micromesh under the applied electric field.
The drift field transports electrons while preserving the structure of the photoelectron track.
The collected charge is then recorded pixel-by-pixel for precise track reconstruction.

## Slide 13
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
GRIDPIX DETECTOR
Comparison of Detector Readout ASICs: GridPix vs GPD
Derived from Medipix
256*256 pixels
14.08 * 14.08 mm² sensitive area
Two mode of operation
3D photoelectron Track
Negligible dead time
TOA
TOT
Developed for X ray Polarimetry
GPD
GridPix
300 * 352 pixels
15 * 15 mm² sensitive area
Self-Triggering Mode
2D photoelectron Track
Dead time(1.1ms)
Heijhoff, K., et al. "Timing measurements with a 3D silicon sensor on Timepix3 in a 180 GeV/c hadron beam."

## Slide 14
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
GRIDPIX DETECTOR
Comparison of Detector Readout ASICs: GridPix vs GPD
55 µm pitch
Common clock 680 MHz & 40 MHz to measure TOA & TOT for each pixel.
Chip outputs digital information
GridPix
GPD
Chip outputs analog information
50 µm pitch
Each pixel integrates the collected charge from the GEM
Diagram shows how the time of arrival (ToA) and time of threshold (ToT) measured by Timepix3 for two different signal Amplitude.
Heijhoff, K., et al. "Timing measurements with a 3D silicon sensor on Timepix3 in a 180 GeV/c hadron beam."

## Slide 15
Why We Do Polarimetry?
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
EXPERIMENTAL SETUP
With GRIDPIX DETECTOR
Prototype of GridPix Detector with Polarized X-ray Source at IAPS-INAF Calibration Facility.
GridPix
Assembly
This image shows detector mounted on the mechnaical satge beneath the polarized x ray source for measurement.
Close-up of detector housing:
- Entrance Window
- Gas Chamber
- Electrical Connection
Complete
Experimental
Setup
Its include X ray
Source, Detector
mounted stage, Positing stage, electronics
and cables

## Slide 16
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
EXPERIMENTAL SETUP
Key Experimental Parameters
Gas Mixture:
Argon(Ar): 80%
Dimethyl ether(C2H6O) : 20%
X ray Source:
Molybdenum(Mo) Target for both:
Polarized Source
Unpolarized Source
Photon Energy:
For Polarized Source : 8.7 KeV & 17.4 KeV
For Un-Polarized Source : 8.05 KeV
Detector Position:
ε = 0
ε= 90
Pressure and Temprature:
Pressure: 1.05 bar
Temperature: at Room Temprature
Run Duration:
Run duration: 30 mins
Number of triggers per run:
Drift Values:
Edrift : 280 V/cm
Vgrid : 340 V
Data Format & Analyzed:
Data Format: HDF5
Analyzed using Python based Software.

## Slide 17
EXPERIMENT RESULTS
Real Data Results
For 8.7 KeV
1.
PHA: Pulse height amplitude . PHA is proportional to charge collected event by event in detector.
Blue curve: Measured energy distribution.
Orange curve: Fitted distribution with Gaussian fit to determine the peak position and resolution.
2.
This plot shows the azimuthal angle distribution of reconstructed photoelectron tracks, which is used to measure X-ray polarization.
Blue histogram
Represent the measured distribution of emission angles from the detector.
Orange curve
A cos² modulation fit applied to the data.
3.TRK_Size
This histogram shows the distribution of reconstructed photoelectron track sizes in pixels.
with most tracks occupying around 170 pixels, reflecting the typical charge spread of photoelectron tracks.
Energy Spectrum
Modulation Curve
3. TRK_SIZE
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
R = 17.5%

## Slide 18
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
EXPERIMENTAL RESULTS
Real Data Results
For 8.7 KeV
4.
This is a 2D correlation plot between deposited energy and track size in the detector.
The bright yellow region represent the most events. corresponding to the main energy peak of the source
This plot shows the spatial distribution of reconstructed events on the detector plane.
Yellow → many events
Purple → fewer events
This histogram shows the distribution of reconstructed photoelectron track sizes, with most tracks have about 150 pixels, corresponding to the typical charge spread of photoelectron tracks in the detector.
5.
6.
Normalized Track Size Distribution

## Slide 19
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
EXPERIMENT RESULTS
Real Data Results
Time Trend Plots for 8.7 KeV
These plots show the stability of detector and polarization parameters as a function of time.

## Slide 20
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
EXPERIMENT RESULTS
Real Data Results
For 17.4 KeV
Top-Left: Track size distribution
Top-Right: Spatial event distribution
Bottom-Left: Azimuthal angle (polarization) distribution
Bottom-Middle: Full track size spectrum
Bottom-Right: Energy vs track size correlation

## Slide 21
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
EXPERIMENT RESULTS
Real Data Results
Energy Calibration Results:
Shows the linear relationship between PHA and photon energy, with small offset.
Shows how the energy resolution improves with increasing photon energy, better detector performance at higher Energies.
Energy Resolution Results:

## Slide 23
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
SIMULATION RESULTS
Energy Spectrum & Resolution
EnRes = 14.92 ± 0.30 % (FWHM/mu)
EnRes = 16.26 ± 0.26 % (FWHM/mu)
Energy Spectrum of 8.7 KeV
Energy Spectrum of 17.4 KeV

## Slide 24
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
SIMULATION RESULTS
Energy Calibration & Track Size Distribution

## Slide 25
Comparison of Real Data And Simulation
Comparison of Real and Simulated Data

## Slide 26
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
COMPARISON RESULTS
Real Data Vs Simulation
Track Size Distribution Comparison Between Experimental and Simulated Data

## Slide 27
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
COMPARISON RESULTS
Energy Calibration and Resolution
Validation of Energy Calibration and Resolution Using Simulated and Experimental Data

## Slide 28
INAF
ISTITUTO
NAZIONALE
DI ASTROFISICA
Conclusion