## Slide 1
Advances in HPGe detector tecnologies for next generation high resolution gamma spectroscopy
Walter Raniero
INFN Legnaro National Laboratory, Italy
walter.raniero@lnl.infn.it
Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026)
24-26/03/2026, Palermo

## Slide 2
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Outline
HPGe gamma detectors
PLM (Pulse Laser Melting): Next generation of contact on HPGe detectors
PLM and Photolitography on Planar and Coxial HPGe segmented detectors
Neutron damage in PLM planar HPGe segmented detectors
HPGe Double-Sided Strip Detector (DSSD)

## Slide 3
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
HPGe detector
B doped = 0.3 µm (implantation)
Li doped = 0.9 mm (diffusion)
HPGe
Charge carrier density: 1010 cm-3
Surface passivation:
Diameter: 40mm
Height: 20mm
HPGe planar detector

## Slide 4
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
HPGe detector state of the art: AGATA
ATC (AGATA triple cluster detector)
2Kg weight
80 mm diameter
90mm long
1 core inner contact
36 segments
J. Eberth et al. Eur. Phys. J. A (2023) 59: 179
3 asymmetrical HPGe detector
Installation at LNL-INFN
J.J Valiente al. NIM Phys. Res. A (2023) 1049

## Slide 5
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
n-type detector «AGATA»
n-type HPGe bulk
n+(Li) [V-]
Passivation
E
h
e
ϒ
The boron junction is thin, and easily segmentable

## Slide 6
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
p-type detector « future AGATA»
n+(Sb)
p+(B) [V+]
n+ (Li)
The lithium junction is thick and not stable under annealing
treatments
p-HPGe with segmentable n+ junction collecting electrons.
Polarity inversion demanded to test higher damage resistivity (hole trapping by neutron damage)
p-HPGe
bulk
E
e
h
ϒ
Thin and thermally stable n-type dopants (Pulsed Laser Melting).
Passivation

## Slide 7
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Outline
HPGe gamma detectors
PLM (Pulse Laser Melting): Next generation of contact on HPGe detectors
PLM and Photolitography on Planar and Coxial HPGe segmented detectors
Neutron damage in PLM planar HPGe segmented detectors
HPGe Double-Sides Strip Detector (DSSD)

## Slide 8
Deposition
Doping layer
pulse < 100ns
Doped layer
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
PLM (Pulse Laser Melting)
Melting temperature is reached - short time (<100 ns)
Only the surface (< 200 nm) is melted, the bulk is at Tamb
High dopant concentrations with very sharp dopant profile
Very clean process suitable for preserving the Ge hyperpurity
Suitable for complex contact geometries (segmentation)
Autors:
D. De Salvador, G. Maggioni, D. Napoli, E. Napolitani

## Slide 9
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
PLM laser Technology
𝜆=248 nm, 22 ns
Frequency: 1-10 Hz
ED= 50-1300 mJ/cm2
Square 5x5mm2 spot
Homogeneity: < 2%
lateral resolution <30 µm
Motorized XYZ stage
5 mm
Excimer KrF laser
Notes: KrF (Krypton Fluoride)

## Slide 10
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
PLM on HPGe: Crystal preparation
3 µm Al2O3 slurry
Polished surface (CMP)
(Low waviness,
low roughness)
rotating disc rinsed with H2O2 1%
(pH 12 with KOH)
Rate: ≈200 nm/min
… + Colloidal Silica (30nm size)
Rate: ≈ 400 nm/min
2025
Grinded surface
Etched surface
(High waviness,
low roughness)
Grinding
Polishing
3 µm Al2O3 slurry

## Slide 11
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
HPGe gamma detector (Passivation)
Passivation techniques: study the evolution of Ge – GeO – GeO2
HNO3/HF (3:1) etching and quenching bath:
Water (Ge_W)
Methanol (Ge_QM)
Sulfide termination (Ge_S)
Hydride termination (Ge_H10)
X-ray Photoelectron Spectroscopy (XPS)
S. Carturan et al., Mater. Chem. Phys. 2015
Ge (2p)
GeO
National Laboratories of Legnaro (LNL)
University of Padua (UNIPD) & IKP Cologne
I-V Curve (leakage current)
Crystal is depleted (reverse bias) Crystal at 80 ÷ 90K

## Slide 12
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
PLM on HPGe: material characterization
Impurities concentration in Ge bulk
Narrow Contact
SIMS (Secondary
Ions Mass Spectrometry)
Van Der Pauw Hall
Al
V. Boldrini et al., Journal of Physics D: Applied Physics (2018) volume 52, 3
1 pulse
- - - 1 pls activation
4 pulses
- - - 4 pls activation
Sb

## Slide 13
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Outline
HPGe gamma detectors
PLM (Pulse Laser Melting): Next generation of contact on HPGe detectors
PLM and Photolitography on Planar and Coxial HPGe segmented detectors
Neutron damage in PLM planar HPGe segmented detectors
HPGe Double-Sided Strip Detector (DSSD)

## Slide 14
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Photolitograpgy on 2D - PLM: 1° type Segmentation
Au deposition
100 nm PVD deposition of Au in Ar plasma with ultrapure target in vacuum (10-6 mbar)
Photolithography
Photoresist deposition, baking, exposure and development, followed by Au stripping and resist removal.
Intercontact gaps passivation
(3:1) HNO3 : HF etching followed by chemical quenching passivation.
Full area PLM

## Slide 15
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
PLM contact: 2° type Segmentation
Selective etching
Removal of untreated dopant using selective etchants to protect the near junction.
Chemical passivation
Passivation of undoped surfaces with suitable solutions.
Lithographic process using selective etching solutions:
Hot pure H2SO4 for Sb deposition (preserve Sb junction)
H2O2 for GeP deposition (slowly etches everywhere)
Kern solutions (H2O2, H3PO4, Ethanol) for Al-Ge deposition (preserve Al junction)

## Slide 16
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Sb/p-HPGE/Al,
N=6x109 cm-3
L=28mm, t=10mm
Soft contact + Al plated (segmented)
Sb/p-HPGE/Al,
N=1.6x1010 cm-3 L=33mm, t=8.5mm
S. Bertoldo et al., MSSP(2025) volume 200
2D – Planar p-type PLM segmented detector

## Slide 17
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Sb/p-HPGE/Al, L=35mm, t=10mm
C-V test to estrapolate HPGe purity
R(Ω) segmentation gap (0.1 ÷ 0.8mm gap)
This confirms that the PLM process is highly safe against bulk contamination of HPGe material
2D – Planar p-type PLM segmented detector

## Slide 18
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
S. Bertoldo et al., MSSP(2025) volume 200
https://doi.org/10.1016/j.mssp.2025.109967
2D – Planar p-type PLM segmented detector

## Slide 19
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Scanning, reconstruction, and pattern
Photolitography on 3D – PLM detector

## Slide 20
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Al plated
After photolithography
Al plated + photoresist spry coating
50mm d, 50mm h,
n-type crystal,
AlGe PLM junction, Li core
Photolitography on 3D – PLM detector

## Slide 21
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Outline
HPGe gamma detectors
PLM (Pulse Laser Melting): Next generation of contact on HPGe detectors
PLM and Photolitography on Planar and Coxial HPGe segmented detectors
Neutron damage in PLM planar HPGe segmented detectors
HPGe Double-Sided Strip Detector (DSSD)

## Slide 22
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Neutrons damage on planar PLM segmented detector
Reaction: 7Li (p,n) 7Be
Prototype detector is located at 30° 9.5 cm
Neutrons are directly measured with
- CLYC7 scintillators, 30° 2 m
- GASP HPGe γ detector, 90° 1 m
7Be + e- 7Li 477.6 keV
Target
N3G Planar HPGe
CLYC7
HPGe GASP
Beam
R. Escudeiro at all. “Neutron radiation damage on a planar segmented germanium detector” proceeding Presented at the XXXVII Mazurian Lakes Conference on Physics, Piaski, Poland, September 3-9, 2023
380nA 4MeV proton beam
7Li target, 100µm

## Slide 23
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Neutrons damage on planar PLM segmented detector
Detector prototype:
Sb/p-HPGE/Al, L= 32mm, t= 8.6mm
4 contacts + guard ring
241Am
241Am
E = 59,5 keV
FWHM < 1,1 keV
60Co
E = 1332 keV
FWHM < 1,8 keV
Resolution at 80V operational bias before the neutron damage run
Energy resolution in the lab

## Slide 24
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Annealing procedure: 7 days at 105°C continuously pumped inside the cryostat
60Co
E = 1332 keV
FWHM < 2 keV
241Am
E = 59,5 keV
FWHM < 1,6 keV
Neutrons damage on planar PLM segmented detector
Drastic drop in resolution after
≈ 3∙109 neutrons/cm2 irradiation fluence
241Am
E = 59,5 keV
FWHM = 3,2 – 4,2 keV

## Slide 25
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Outline
HPGe gamma detectors
PLM (Pulse Laser Melting): Next generation of contact on HPGe detectors
PLM and Photolitography on Planar and Coxial HPGe segmented detectors
Neutron damage in PLM planar HPGe segmented detectors
HPGe Double-Sided Strip Detector (DSSD)

## Slide 26
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
HPGe Double-Sided Strip detector (DSSD)
LAUE lens telescope
HRIS (high resolution imager and spectrometer for space astronomy)
PI: N. Auricchio
Cadmium Zinc Telluride (CZT)
HPGe n-type: high resolution, position sensitive detector
Silicon DSSSD detector
HPGe n-type DSSD detector:
Rad-hard detector,
efficiency for detecting beta-decay electrons
Radioactive ion beams (RIB) enable exploration of nuclei far from stability, with high‑energy ions implanted into an active stopper in in‑flight decay‑spectroscopy experiments. Use HPGe as an active stopper in decay-spectroscopy
PI: S. Capra

## Slide 27
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
HPGe p-type Double-Sided Strip detector (DSSD)
Al/Ge p+
Al
Al
Sb-n+ junction
p-type HPGe
6 top Al/Ge PLM P+
6 bottom Sb PLM N+
Thin 2mm prototype detector
Cross-section:
2026

## Slide 28
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
HPGe p-type Double-Sided Strip detector (DSSD)
Single diode with continuous junctions and metalization
Photoresit and bake on both sides
Exposure and development
Trenching and resist removal
2026

## Slide 29
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
HPGe p-type Double-Sided Strip detector (DSSD)
I-V Test (Sb side)
C-V Test
2026
I-V Test (Al side)

## Slide 30
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Summury
The research carried out on PLM contacts has produced promising results:
PLM n+/p+ contacts technology preserve hyperpurity of HPGe crystal
PLM n+/p+ contacts are rad-hard on neutron damage (annealing recovery)
PLM n+/p+ contacts are thermally stable, narrow with high dopant concentration profile
PLM n+/p+ contacts are thinner (150÷250nm) enabling for segmented HPGe detectors; this technology could be valuable for developing new types of detectors, such as DSSD planar devices for nuclear or astrophysics application.
ò

## Slide 31
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
R&D Gamma ray detector Team
Davide De Salvador
Stefano Bertoldo
Enrico Napolitani
Francesco Sgarbossa
Chiara Carraro
Sara Carturan
Gianluigi Maggioni
Francesco Recchia
Dino Bazzacco
Filippo Nicolasi
Walter Raniero
Daniel Napoli
Andrea Mazzolari
Lorenzo Malagutti
Andres Gadea
Stefano Capra
Giacomo Secci
Alberto Pullia
Bénédicte Million
Luciano Manara

## Slide 32
Thanks for
your Attention !
walter.raniero@lnl.infn.it

## Slide 34
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Vdepl. : 1300V
n+ junction
(spring contact)
n+ junction (!) / (!)(!)
(indium pad)
n+ junction (!)(!)(!)
(elastic tabs)
Sb/p-HPGE/Al,
N=6x109 cm-3 D=40mm, t=20mm
Sb/p-HPGE/Al,
N=1.6x1010 cm-3 L=35mm, t=10mm
Contact force
Breakdown Voltage
2D – Planar PLM segmented detector : I-V test

## Slide 35
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
241Am
E = 59,5 keV
FWHM = < 2 keV until threshold
Operational Voltage 80-120-160V
Neutron irradiation for 20 and 2 min to 5 min runs with 241Am to better characterize resolution worsening
Drastic drop in resolution after
≈ 3∙109 neutrons/cm2 irradiation fluence
Neutrons damage on planar PLM segmented detector

## Slide 36
W. Raniero - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
PLM laser Technology
Z-axis
Rotor
XY-axes
Z-axis
Rotor
2D - Planar detector
3D - Coaxial detector
Laser beam