## Slide 1
Integrated Front-End Electronics and Flexible Interconnection Systems for Segmented High-Purity Germanium Detectors
Giacomo Secci
INFN Milano, Italy
Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026)
24-26/03/2026, Palermo

## Slide 2
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Content
State-of-the-Art HPGe Detectors
The N3G Projects (Next Generation Germanium Gamma detectors)
Innovative Electrical Connections
Integrated Charge Sensitive Preamplifiers for HPGe Detectors: Design and Experimental Validation

## Slide 3
High Purity Germanium (HPGe) detectors: impurity concentration of 1010 cm-3
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
State-of-the-Art HPGe Detectors
Remarkable energy resolution: ≈ 2 keV @ 1.33 MeV
Interaction points identification through detector segmentation and Pulse Shape Anaysis (PSA)
γ trajectories reconstruction: tracking algorithms
AGATA (Avanced GAmma Tracking Array)
GRETA (Gamma-Ray Energy Tracking Array)

## Slide 4
n-type HPGe germanium crystals, featuring an external p-type boron implantation
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
State-of-the-Art HPGe Detectors
Heavily n-doped cavity realized by lithium evaporation and diffusion
Passivation layer to prevent electrical discharges
Tapered hexagonal geometry with asymmetric shape: 4π configuration

## Slide 5
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
State-of-the-Art HPGe Detectors
Cold stage: input FET (BF862-20 mW power consumption), feedback capacitor (1 pF) and feedback resistor (1 GΩ) → "block-fet" Faraday cage
Warm stage: trans-impedance preamplifier with fast reset circuit

## Slide 6
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
State-of-the-Art HPGe Detectors
Each block-fet includes the input FET and the feedback components of three preamplifier channels
Each warm PCB implements three preamplifier channels

## Slide 7
State-of-the-art HPGe detectors are supplied to scientific collaborations by industry, and their technical and manufacturing specifications are not known.
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Next Generation HPGe Detectors
N3G Project: Next Generation Germanium Gamma Detector (INFN)
GOAL: developing new HPGe detectors while acquiring the know-how about their management and maintenance
Physics of Semiconductors: doping technologies photolithography
Mechanical Advances: vacuum technology enclosures
Electronics: detector connection charge sensitive preamplifiers

## Slide 8
Innovative doping technique: Pulsed Laser Melting (PLM)
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Junction depth on the order of 100 nm – 300 nm: easier segmentation and dead-layer minimization, BUT…
difficult to be connected to the front-end electronics
Surface scratches and excessive pressure may cause irreversible damage to the junction
Next Generation HPGe Detectors

## Slide 9
GOAL: reliable electrical interface for signals transmission
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Mechanical and thermal integrity without any structural or functional damage
Flexible PCBs to be wrapped around the detector
Detector Electrical Connections

## Slide 10
Three models developed: one meets the stringent project requirements
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Digital microscope inspection: no visible damage
Test with a planar HPGe detector: significant reduction of leakage current compared to standard contact methodologies
Detector Electrical Connections
micro-spring
connection
flexible-PCB
connection

## Slide 11
We consider any electrode as a current generator with a parallel capacitance
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Read-Out: Charge Sensitive Preamplifiers
We can describe the induced current with a delta-like pulse:
The CSP is the active circuit that integrates the current pulse into a voltage signal:
The amplitude of the signal is proportional to the charge released by the studied interaction

## Slide 12
We consider any electrode as a current generator with a parallel capacitance
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Read-Out: Charge Sensitive Preamplifiers
We can describe the induced current with a delta-like pulse:
The CSP is the active circuit that integrates the current pulse into a voltage signal:
The signal can then be sent to a shaping amplifier or directly digitized

## Slide 13
Integrated charge sensitive preamplifiers for holes-collecting electrodes: developed using the AMSC35B4C3 350 nm process
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Read-Out: Charge Sensitive Preamplifiers

## Slide 14
The CSP architecture employs a four-stage configuration, comprising three common-source gain stages followed by a source follower buffer
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP
input transistor optimization: W and L influence noise in opposing ways:
Larger area = larger capacitance = more white series noise
A trade-off is necessary: W/L = 2500 µm / 0.7 µm
Larger area = lower Af= lower flicker noise
The gate must be folded multiple times!

## Slide 15
The CSP architecture employs a four-stage configuration, comprising three common-source gain stages followed by a source follower buffer
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP
input transistor optimization: W and L influence noise in opposing ways: W = 2500 µm and L = 0.7 µm optimized values
RC Miller Compensation
Line-driver architecture based on Operational Transcionductance Amplifier
2-stages line-driver configured as non-inverting amplifier: operating point externally adjustable
50 Ω terminated line driving capability

## Slide 16
Wire bond are transmission lines (r = 0.1 Ω mm−1, l = 1 nH mm−1, c = 0.1 pF mm−1): they introduce high-frequency resonance on the line-driver loop gain
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP
Local RC load reduce this high-frequency resonance
Decoupling capacitors near power pads prevent resonance on supply lines

## Slide 17
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
Unfortunately, we do not yet have a fully operational detector available and disassembling one would be too risky.
ASIC directly bonded on a test PCB featuring a PTFE substrate
900 µm x 475 µm area consumption
25.2 mW power consumption
21 MeV equivalent dynamic range
The detector capacitance was emulated by connecting a capacitor between the input-node and ground

## Slide 18
1 MeV-15 MeV range of interest for gamma spectroscopy experiments
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
22.8 ns rise time @ Cd = 15 pF

## Slide 19
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
Non-linearity analysis: CSP + shaping amplifier (ORTEC 572) + multi-channel analyzer
Integral non-linearity factor < 0.3‰ on the 15 MeV equivalent dynamic range

## Slide 20
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
Noise analysis: pulser + CSP + shaping amplifier (ORTEC 572)
Equivalent Noise Charge (ENC):
FWHM Resolution:
)
Fano statistic limit: 209 e−rms (@ 1 MeV)

## Slide 21
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
Fano statistic limit: 209 e−rms (@ 1 MeV)
Noise analysis: pulser + CSP + shaping amplifier (ORTEC 572)

## Slide 22
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
Fano statistic limit: 1.46 keV FWHM @ 1 MeV
Noise analysis: comparison with state-of-the art discrete-component CSPs

## Slide 23
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
Frequency by frequency measurement of
Noise power spectral density analysis
Frequency by frequency measurement of the output spectral density of noise: Sout(f)
Input-referred power spectral density of noise:

## Slide 24
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
Noise power spectral density analysis for Cd = 0 pF, 4.7 pF, 15 pF, 22 pF
… from theory:
… to reality:

## Slide 25
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Results
PCB substrate study: FR4 and PTFE boards compared under identical conditions at zero detector capacitance to isolate substrate noise

## Slide 26
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Integrated CSP: Mounting Strategies
ASIC advantages: low power consumption and compact area enable alternative mounting strategies inside the cryostat

## Slide 27
Integrated charge sensitive preamplifiers for electrons-collecting electrodes
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Read-Out of Electrons-Collecting Electrodes
Experimental validation:
35 mW power consumption
21 MeV dynamic range
30 ns rise-time with Cd = 15 pF (stronger compensation)
0.3 ‰ integral non-linearity factor (15 MeV dynamic range)

## Slide 28
All preamplifiers are also available with differential output
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
CSP with differential output

## Slide 29
Next generation HPGe detectors introduce new challenges
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Conclusion
New flexible PCBs were engineered to realize a robust connection while ensuring mechanical and thermal integrity
New ASIC CSPs were designed and successfully tested
The new circuit architecture has demonstrated high reliability, excellent linearity and very low noise.

## Slide 30
Thank You for
Your Attention !
giacomo.secci@mi.infn.it