## Slide 1
Design, Fabrication and Qualification of Cryogenic High-Vacuum Enclosures for HPGe Detectors
Giacomo Secci
INFN Milano, Italy
Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026)
24-26/03/2026, Palermo

## Slide 2
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Content
State-of-the-Art HPGe Detectors
The N3G Projects (Next Generation Germanium Gamma detectors)
Mechanical Solutions for HPGe Detectors: Vacuum-Tight Capsules Design and Validation
Inside the Capsule: Innovative Detector Support Architecture and Mounting Strategies

## Slide 3
High Purity Germanium (HPGe) detectors: impurity concentration of 1010 cm-3
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
State-of-the-Art HPGe Detectors
Remarkable energy resolution: ≈ 2 keV @ 1.33 MeV
Interaction points identification through detector segmentation and Pulse Shape Anaysis (PSA)
γ trajectories reconstruction: tracking algorithms
AGATA (Avanced GAmma Tracking Array)
GRETA (Gamma-Ray Energy Tracking Array)

## Slide 4
n-type HPGe germanium crystals, featuring an external p-type boron implantation
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
State-of-the-Art HPGe Detectors
Heavily n-doped cavity realized by lithium evaporation and diffusion
Passivation layer to prevent electrical discharges
Tapered hexagonal geometry with asymmetric shape: 4π configuration

## Slide 5
State-of-the-art HPGe detectors are supplied to scientific collaborations by industry, and their technical and manufacturing specifications are not known.
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Next Generation HPGe Detectors
N3G Project: Next Generation Germanium Gamma Detector (INFN)
GOAL: developing new HPGe detectors while acquiring the know-how about their management and maintenance
Physics of Semiconductors: doping technologies photolithography
Mechanical Advances: vacuum technology enclosures
Electronics: detector connection charge sensitive preamplifiers

## Slide 6
...once upon a time, in a faraway factory, a hyper-pure germanium crystal was born...
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detectors Fabrication
cleaning and polishing
RFMS Radio Frequency Magnetron Sputtering: dopant precursor film deposition (e.g. aluminum, antimony...)
PLM process: the laser pulse is directed onto the surface of the rotating sample
UV laser makes the dopant diffuse in germanium
rapid regrowth rebuilds the crystalline structure that tightly incorporates the dopant
layer thickness: 100 nm-300 nm

## Slide 7
100 nm gold deposition through RF sputtering: lower contact resistance and junction protection
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detectors Fabrication
Photolithography
(positive) photo-resist deposition and soft bake
light exposure and photo-resist development
baking and drying
trench passivation in methanol baths
etching: electrodes isolation
cleaning: photo-resist stripping with acetone

## Slide 8
HPGe detectors require operation in:
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detectors Operation Conditions
High-vacuum environment
Cryogenic environment
Moisture and oxygen can react with the detector surface and degrade its electrical properties
Germanium narrow band-gap (0.66 eV)
Cryostat: vacuum-insulated structure designed to provide both mechanical support and efficient thermal coupling to a cryogenic reservoir of liquid nitrogen (77 K)

## Slide 9
The vacuum-tight capsule contains and protects the HPGe crystal:
  cylindrical aluminum case
  stainless steel flange
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Design
Electrical feed-trough connectors
Vacuum port
Helicoflex Gasket
Porous Getter

## Slide 10
High-performance metal seal designed for extreme environments (aerospace, vacuum, nuclear)
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Helicoflex Gaskets
Soft outer metallic layer: intentionally softer than flange material to ensure tight sealing
Viscoplastic deformation: material flows into surface micro-asperities
Long-term reliability: accommodates tolerances, surface imperfections, and operational variations
Internal helical spring: provides elasticity and maintains consistent contact pressure

## Slide 11
Non-evaporable getters: reactive metallic materials (Zr, Ti, V) that remove gases via chemisorption
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Porous Getter
Thermal activation required (≈450–900 °C): removes oxide layer and exposes reactive surface
Selective gas sorption: effective for O₂, N₂, H₂, CO, CO₂ , ineffective for noble gases and light hydrocarbons
Regenerable performance: pumping speed decreases with use but can be restored through thermal reactivation
ST 172 porous getters: high surface area, enhanced pum-ping efficiency and low-temperature activation (~450–500 °C)

## Slide 12
Tracer gas leak detection (mass spectrometry): high sensitivity and fast respons; independent, self-contained system
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Leak Detection
Helium as tracer gas: sensitivity down to ~10⁻¹¹ mbar·L/s; limitations include memory effects and material permeation

## Slide 13
Components cleaned with isopropyl alcohol before sealing
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Leak Test
Leak-test configuration optimized to minimize parasitic leaks
Helicoflex® gaskets installed according to datasheet specifications: analytical tightening torque calculation

## Slide 14
Components cleaned with isopropyl alcohol before sealing
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Leak Test
Leak-test configuration optimized to minimize parasitic leaks
+10% tightening torque applied to flange screws:

## Slide 15
Getter activation is essential
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Getter Activation
Dedicated manipulator: capsule evacuation and sealing plug opening/closing
PT100 temperature sensor installation
Getter activation: 6 A current over ∆t = 10 min

## Slide 16
Getter activation is essential
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Getter Activation
Dedicated manipulator: capsule evacuation and sealing plug opening/closing
PT100 temperature sensor installation
Temperature monitoring
Getter activation: 6 A current over ∆t = 10 min

## Slide 17
Getter activation is essential
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Getter Activation
Dedicated manipulator: capsule evacuation and sealing plug opening/closing
PT100 temperature sensor installation
Pressure monitoring
Getter activation: 6 A current over ∆t = 10 min

## Slide 18
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Capsule: Validation
Pressure measured at the cap tightening:
Pressure measured at the cap reopening (∆t = 5d):
Effective capsule pressure:
Over 5 days: (from 10−6 mbar to 10−4 mbar in one month)
Cryogenic vacuum operation improves long-term vacuum stability (capsule in cryostat)

## Slide 19
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Inside the Capsule: Support Structure
Coaxial HPGe detector mounted within enclosure using a dedicated mechanical support system
Alumina hollow conical support: provides structural stability and precise alignment of the crystal
Indium: conductive contact while preventing mechanical stress and friction
Central electrical rod: high negative bias voltage (3–5 kV) for detector polarization

## Slide 20
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Inside the Capsule: Support Structure
Spring-supported base: alumina cone rests on a single-wave annular spring (Inconel) housed in an aluminum holder fixed to the flange
Cryogenic compatibility: Inconel spring ensures elasticity and mechanical stability at liquid-nitrogen temperatures

## Slide 21
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Inside the Capsule: Support Structure
Detector, cone, spring and spring-holder closed by a PEEK cylindrical case
Clamping system: aluminum ring + M2.5 steel screws fix the PEEK case to the spring-holder
Controlled tightening force: 4.4 N mm tightening torque (half the torque to unwrap a candy!)
Crystal position independent from the container absolute orientation

## Slide 22
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Assembly

## Slide 23
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Assembly

## Slide 24
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Assembly

## Slide 25
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Assembly and Test
Liquid nitrogen cooling down and test
I-V characterization: leakage current measured with a precision ammeter
Upper segment: high leakage current … issues during etching
Bottom segments: high leakage current … gold on passivation
Middle segments: leakage current in the picoampere range

## Slide 26
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Detector Assembly and Test
Resistance measurements between adjacent segments
Measurements performed in the middle region to minimize leakage current
10 TΩ between adjacent lateral segments
Few TΩ between middle and top/bottom segments
High resistance confirms effective electrical isolation from trenching
Low breakdown voltage and surface damage prevented charge collection measurements
Connection system and mechanical support proved reliable and efficient

## Slide 27
A reliable design for cryogenic vacuum enclosures has been developed and validated for operation at liquid nitrogen temperatures
G. Secci - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Conclusion
Sealing and vacuum stability are effectively ensured through HelicoFlex technology and NEG-based chemical pumping
Leak detection methodologies based on mass spectrometry and helium tracing proved essential for system validation
Experimental results confirm excellent performance, with minimal pressure increase
Mechanical integration of the detector guarantees stable operation, proper biasing and reproducible assembly

## Slide 28
Thank You for
Your Attention !
giacomo.secci@mi.infn.it