## Slide 1
Direct pre-amplifier sampling ADC for high-count rate and high-resolution X-ray spectroscopy
A. Abba, V. Arosio, P. Barcella, F. Caponio, E. Carlotti, M. Carminati, S. Carsi, C. Fiorini, B. Pedretti, E. Proserpio, C. Tintori
Nuclear Instruments
CAEN SPA
Politecnico di Milano – DEIB
XGRADE – Palermo 2026 – 26-03-2026

## Slide 2
Comparison between resistive feedback and transistor reset charge amplifier
RESISTIVE FEEDBACK CHARGE PRE-AMPLIFIER
TRANSISTOR RESET CHARGE PRE-AMPLIFIER

## Slide 3
Pileup and saturation on continuous reset amplifier

## Slide 4
Pileup and saturation on continuous reset amplifier

## Slide 5
Transistor reset amplifier for high rate detector
RATE INCREASE

## Slide 6
Transistor reset amplifier for high rate detector

## Slide 7
Transistor reset amplifier for high rate detector

## Slide 8
Sampling the transistor reset detector output

## Slide 9
Reshape and amplify TRP output signal
Additional noise introduced by amplifier
Pileup effect  trade-of between shaping time and gain
Detector output

## Slide 10
Reshape and amplify TRP output signal
Deconvolution
M=t/DT

## Slide 11
Reshape and amplify TRP output signal
Low M=t/DT ratio
High M=t/DT ratio
To introduce Low sampling error M should be high
  High M means:
Ts = 10ns  100Msps
Simulation 5 Mcps
Fe55 Ka = 500mV
(2.2% = 130eV on Fe55 Ka)
        High sample rate (fast ADC)
        Long shaping time (pileup)

## Slide 12
Reshape and amplify TRP output signal
Low M=t/DT ratio
High M=t/DT ratio
To introduce Low sampling error M should be high
  High M means:
Ts = 10ns  100Msps
Simulation 5 Mcps
Fe55 Ka = 170mV
(2.2% = 130eV on Fe55 Ka)
        High sample rate (fast ADC)
        Long shaping time (pileup)

## Slide 13
Reshape and amplify TRP output signal
Deconvolution

## Slide 14
Reshape and amplify TRP output signal
Deconvolution

## Slide 15
Reshape and amplify TRP output signal
Deconvolution
Signal Sharpening
Converts each exponential decay into short impulse of amplitude ΔI
Improves temporal resolution and separates overlapping events
High-Pass Action on Noise
Acts as a digital high-pass filter: attenuates low-frequency drift but amplifies mid-/high-frequency components
White noise spectrum becomes “colored,” with increased power at higher frequencies
Noise Variance Increase If input noise has variance σ², output variance ≈ (1 + α²) σ²RMS noise grows by factor √(1 + α²)
SNR Degradation
Peak amplitude remains ΔI, but noise floor rises resulting

## Slide 16
Reshape and amplify TRP output signal
Deconvolution
+
Digital Shaping
Deconvolution
Trigger (Fast Shaper)
Slow Shaper
Baseline Restorer
-
Energy Sample
Spectrum

## Slide 17
Reshape and amplify TRP output signal
Deconvolution
+
Digital Shaping
Over Compensated
Correctly Compensated
Under Compensated

## Slide 18
Direct sampling with 20 bit ADC
Trigger (Fast Shaper)
Slow Shaper
Energy Sample
Spectrum
TRP
Output

## Slide 19
Direct sampling with 20 bit ADC

## Slide 20
20 bit ADC effective resolution
5.9KeV = 25mV
3V
Resolution | LSB (mV) | LSB per keV
14 bit | 3 V/2¹⁴ ≃ 0.1831 | 4.237 mV/keV / 0.1831 mV/LSB ≃ 23.2 LSB/keV
16 bit | 3 V/2¹⁶ ≃ 0.04578 | 4.237 / 0.04578 ≃ 92.6 LSB/keV
20 bit | 3 V/2²⁰ ≃ 0.002861 | 4.237 / 0.002861 ≃ 1 480 LSB/keV

## Slide 21
Interleaving 160 MHz 20 bit digitizer using 4x40 MHz SAR ADC
40 MHz 20 bit SAR ADC

## Slide 22
Gain and offset calibration
Calibration Signal
Acquire Calibration Ramps For each ADC channel, apply a precise linear ramp 𝑥(𝑡)x(t) (e.g. 0 → 𝑉​ over 2 µs) from a 20-bit DAC.
Fit a Straight Line
Correct Offsets Subtract each channel’s own offset so that each line passes through the origin:
Equalize Gains
mref is channel 1
mi
-qi

## Slide 23
Gain and offset calibration
Calibration Signal
PROGRAMMABLE PHASE SHIFT CLOCK GENERATOR
SAMPLE
SIN FIT (LM FIT)
PHASE SHIFT EXTIMATOR
< 100ps
ADJUST CLOCK DELAY LINES
PHASE LOCKED

## Slide 24
System overview
FPGA (Zynq 7020)
4xADC 20 bit
ADC DRIVER
Low Noise Front End
Calibration DAC
Clock generator with phase shifter
Signal Input
Fast GPIO (3 in 3 Out TTL)
Slow GPIO (16 in/Out TTL)
+-250V BIAS SDD/SiPIN/PIP/SiPM 1mA
+-5KV BIAS HpGE/CZT/CdTE 0.5mA
Standard AMPTEK/MOXTEK connector
2A 2V Peltier driver

## Slide 25
System overview

## Slide 26
Characterization setup
KETEK AXAS 3.0 SDD Module
V000-KTN0-H020-MA2C
Analog output of the detector
Ultraspectra
Ultraspectra web interface
55Fe 25uCi / 10mCi
Detector courtesy of KETEK

## Slide 27
Characterization setup
Mn Ka Resolution=125 eV

## Slide 28
Low rate (10 kcps OCR) resolution vs shaping time
Chart: Resolution (eV) vs Peaking Time (ns)
| 18.75 | 25 | 31.25 | 37.5 | 50 | 62.5 | 100 | 125 | 150 | 187.5 | 250 | 375 | 500 | 625 | 750 | 875 | 937.5 | 1000 | 1062.5 | 1125 | 1187.5 | 1250 | 1312.5 | 1375 | 1437.5 | 1500 | 1625 | 1750 | 1875 | 2125 | 2375 | 2625 | 3000 | 3500 | 4000
 | 209 | 181.7 | 175.5 | 168.6 | 160.8 | 155.8 | 148.2 | 145 | 142.9 | 139.2 | 133.5 | 129.3 | 128.4 | 128.3 | 126.6 | 125.4 | 125.6 | 125.5 | 126.1 | 125.4 | 126.2 | 125 | 125.3 | 125.5 | 125.6 | 125.6 | 125.7 | 125.7 | 126 | 126.2 | 126.5 | 126.6 | 127 | 128.9 | 129.6
Chart: Resolution (eV) vs Peaking Time (ns)
| 250 | 375 | 500 | 625 | 750 | 875 | 937.5 | 1000 | 1062.5 | 1125 | 1187.5 | 1250 | 1312.5 | 1375 | 1437.5 | 1500 | 1625 | 1750
 | 133.5 | 129.3 | 128.4 | 128.3 | 126.6 | 125.4 | 125.6 | 125.5 | 126.1 | 125.4 | 126.2 | 125 | 125.3 | 125.5 | 125.6 | 125.6 | 125.7 | 125.7
Chart: Resolution (eV) vs Peaking Time (ns)
| 18.75 | 25 | 31.25 | 37.5 | 50 | 62.5 | 100 | 125 | 150 | 187.5 | 250 | 375
 | 209 | 181.7 | 175.5 | 168.6 | 160.8 | 155.8 | 148.2 | 145 | 142.9 | 139.2 | 133.5 | 129.3

## Slide 29
Trapezoidal filter-based throughput (OCR vs ICR)

## Slide 30
High rate (> 1 Mcps OCR) resolution characterization
ICR (cps) | OCR (cps) | Resolution (eV) | Live Time (%)
507k | 358k | 141 | 71 %
510k | 409k | 144 | 80 %
1558k | 1049k | 155 | 67 %
1558k | 1236k | 162 | 79 %
1846k | 1358k | 171 | 74 %
2730k | 2018k | 200 | 74 %
Live Time (%)
Chart: Resolution / Live-time vs OCR
| 358 | 409 | 1049 | 1236 | 1358 | 2019
Resolution | 141 | 144 | 155 | 162 | 171 | 200
Live Time | 0.70611 | 0.80354 | 0.6733 | 0.79332 | 0.73604 | 0.73956

## Slide 31
55Fe Spectrum – ICR 1.6 Mcps – 74% Live Time 171 eV resolution
Mn Ka Resolution=171 eV

## Slide 32
Ardesia Project ARray of DEtectors for Synchrotron radIation Applications
Synchrotron Radiation Source
Focusing
Optic
Incident X-rays
Emitted X-ray
Fluorescence
Energy Dispersive
Detector
Sample
Monochromator
XRF
XAFS
XFM
Detector requirements:
X-ray energy range: 0.2 keV – 20 keV (Si detection region)
Target energy resolution (~125 eV of FHWM Mn-Kα) at optimum shaping time and low rate, good resolution (≤ 200 eV) at short shaping time and high throughput (> 1 Mcps/ch.)
Modular and scalable design, to easily increase sensitive area and adapt to different experiment configurations
Motivation :
Develop a low-noise and versatile detector based on monolithic arrays of Silicon Drift Detectors for high-rate synchrotron applications
(mainly XRF, XAFS, and XFM)

## Slide 33
Ardesia Project ARray of DEtectors for Synchrotron radIation Applications
SDD array
4 × 4 5-mm-side
24 mm
Mo collimator
Detector Jig
Module PCB
Al Frame
Hosts the 16/64 channel Monolithic SDD array (450 µs, 80 µm, or 1 mm thicknesses)
500 µm thick Molybdenum Mask to prevent charge sharing
Active area after collimation: 324 mm2 (81% of window)
Four custom-designed 4-channel CUBE Preamplifiers
Preamplifiers and Detector are glued to a High-density PCB
Increase high-rate capability of SDD-based X-ray spectrometers
4 × 4 5-mm-side
24 mm
24 mm
8 × 8 2.5-mm-side

## Slide 34
Resolution vs Shaping Time (ARDESIA)
Chart: Resolution vs Shaping Time
| 13 | 20 | 50 | 80 | 100 | 120 | 150 | 200 | 300 | 500 | 800 | 1000 | 1200 | 1500 | 2000 | 3000 | 4000 | 5000 | 8000
2 Kcps | 202 | 189 | 154.7 | 145.5 | 140 | 135.2 | 132 | 128 | 129.7 | 129.9 | 131.7 | 134.6
250 Kcps | 160 | 151 | 145 | 141 | 136 | 132
500 Kcps | 153 | 147 | 145 | 144 | 140 | 134 | 131
1500 Kcps | 205 | 176 | 165 | 150 | 146 | 142 | 136
ICR:

## Slide 35
High rate (> 1 Mcps OCR) resolution characterization (ARDESIA)
ICR (cps) | OCR (cps) | Resolution (eV) | (Live Time %)
19.5 | 19k | 140 | 97.0%
1230k | 740k | 163 | 60.0%
1230k | 806k | 162 | 65.0%
1230k | 985k | 176 | 80.0%
1210k | 1020k | 220 | 95.0%
Live Time (%)
Chart: Resolution and Live time vs. OCR
| 19000 | 740000 | 805167 | 985250 | 1020315
Resolution (eV) | 140 | 163 | 162 | 176 | 220
Live Time (%) | 0.97 | 0.6 | 0.65 | 0.8 | 0.95

## Slide 36
55Fe Spectrum – ICR 1.2 Mcps – 85% Live Time 190 eV resolution (ARDESIA)
Mn Ka Resolution=190 eV

## Slide 37
Realtime MCA processing algorithm
Z-k
Z-m
ACC
_
_
Leakage current subtraction
-
Z-k
Z-m
ACC
_
_
DERIVATE
ENERGY FILTER
TRIGGER FILTER
SAMPLE
>
DETECTOR
SPECTRUM
THRESHOLD

## Slide 38
Trigger algorithm signals
Detector output

## Slide 39
Energy Filter
Energy Sampling
samples

## Slide 40
Leakage current compensation
Linear leakage extimator
Trigger
ADC
-
Leakage ramp generator
Leakage free input signal

## Slide 41
High rate model-based fitting filter

## Slide 42
High rate model-based fitting filter
Digital CFD with sub-sample regression
Linear Regression (AFTER)
Linear Regression (BEFORE)
CFD:

## Slide 43
High rate model-based fitting filter
Minimum acceptable distance between events
ACCEPTED
REJECTED

## Slide 44
High rate model-based fitting filter
Chart: Resolution vs OCR (ICR = 1.62 Mcps)
| 1520 | 1489 | 1455 | 1394 | 1361 | 1334 | 1279 | 1226 | 1147 | 1099 | 1052 | 1007 | 882 | 797 | 642 | 517 | 336 | 262 | 212
 | 185 | 184 | 182 | 179 | 177 | 176 | 172 | 170 | 166 | 164 | 162 | 160 | 157 | 155 | 151 | 148 | 143 | 142 | 141
Chart: Live Time vs Resolution - (ICR = 1.62Mcps)
| 185 | 184 | 182 | 179 | 177 | 176 | 172 | 170 | 166 | 164 | 162 | 160 | 157 | 155 | 151 | 148 | 143 | 142 | 141
 | 0.935 | 0.915 | 0.895 | 0.859 | 0.837 | 0.82 | 0.79 | 0.754 | 0.705 | 0.67 | 0.647 | 0.629 | 0.545 | 0.49 | 0.395 | 0.318 | 0.2 | 0.16 | 0.13

## Slide 45
Comparison between trapezoidal shaper and regression
Chart: Fast Fit/Slow Trapezoidal Resolution comparison (ICR = 1.6Mcps)
| 1520 | 1505 | 1489 | 1455 | 1419 | 1394 | 1375 | 1361 | 1334 | 1312 | 1279 | 1255 | 1226 | 1190 | 1147 | 1139 | 1099 | 1052 | 1018 | 1007 | 951 | 882 | 852 | 797 | 794 | 672 | 642 | 554 | 517 | 383 | 336 | 262 | 232 | 212 | 128
FIT (eV) | 185 | 184 | 182 | 179 | 177 | 176 | 172 | 170 | 166 | 164 | 162 | 160 | 157 | 155 | 151 | 148 | 143 | 142 | 141
Trap (eV) | 289 | 225 | 209 | 188 | 173 | 171 | 162 | 155 | 152 | 147 | 145 | 141 | 138 | 134 | 133 | 133
At constant ICR, the trapezoidal and fit methods are compared in terms of resolution using the same dead time

## Slide 46
Ultraspectra & Web Interface
1 channels 160 Msps, 20 bit digitizer
Realtime TRP optimized trapezoidal filter and model-based regression for high rate operation
Up to 32 million point memory for waveform digitalization
SmartMCA web based software with realtime fitting and calibration
C/C++/C#/Python/Java/NodeJS/Labview … SDK and REST API for slow control and readout
Onboard ARM CPU with Ubuntu Linux
Easy integration in SCADA system (TANGO and EPICS integration is ongoing). Onboard Grafana connector.
Fully supported in Sci-Compiler

## Slide 47
Future Development: 16 channels Ultraspectra
16 channels, 160 Msps
±6, ±4, ±2 input dynamic range
Trapezoidal and real-time fitting
1 Gbps + 10 Gbps readout interfaces
Tango and Epics support
Web Interface
Rack 2U format

## Slide 48
Firmware customization
Allows for customizing the firmware of digitizers and logic units.
Simple and intuitive graphical programming language to learn.
Nearly 200 IP cores designed for nuclear physics applications.
Enables any researcher to develop their own acquisition and processing setup without FPGA knowledge.
Firmware portability across different hardware.
Extends the system's lifecycle by maintaining firmware compatibility with future hardware.

## Slide 49
Firmware customization
DAQ121 – 1 Gsps 32 channels differential – 12 bit
DAQ141 – 1 Gsps 32 channels differential – 14 bit
(COMPATIBLE WITH ESS & ISIS READOUT)
DT5571
DT5560
V2495
DT550W
R5560
20 bit !!
V2730 V2740/5
DT1260
SCICOMPILER
OPENFPGA
HARDWARE PLATFORMS

## Slide 50
Firmware customization
FIRMWARE
READOUT GUI
SIMULATE
WINDOWS AND LINUX SDK

## Slide 51
Open source SDK compatible with all Nuclear Instruments and CAEN OpenFPGA Digitizer
Available open source on GitHUB
https://github.com/NuclearInstruments/SCISDK
Works on: Windows, Linux
Compatible with Raspberry PI
Available on: PIP, NPM, and APT (for Ubuntu) and ANT (java)
Documentation and examples available here:
https://nuclearinstruments.github.io/SCISDK/
Works inside Jupyter Lab, pre installed on new instruments