## Slide 1
High precision X-ray measurements with the VOXES spectrometer at INFN-LNF
Alessandro Scordo
INFN Frascati National Laboratory, Italy
Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026)
24-26/03/2026, Palermo

## Slide 2
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Spoil & Wake Up
In this talk you’ll hear about a X-ray detection system that:
Reaches ~ 1330 DE/E - 0.5% E/DE (7 eV FWHM) @ 8 keV
Works @ Room Temperature
Does not need to be operated in a facility
It is, actually, a facility by itself
It’s (more or less) compact and movable
…and wine!!!

## Slide 3
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Original Motivations
Kaonic atoms allow to test strong interaction with strangeness at threshold (PK- ~ 0 MeV/c)
Still open questions in the field require new and more precise measurements of low energy and sub-eV width transitions

## Slide 4
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Original Motivations
nl = 2dsinqB
FWHM ~ 1-10 eV can be achieved depending on the quality of the crystal and the dimensions of the detectors
Natural background reduction from geometry
But….
- Small solid angles can be covered
- Typical efficiencies : 10-5 – 10-8
- Typical d (Si) ≈ 5.5 Å (good for E < 6 keV)
- Typical source size 10-100 mm
Bragg spectroscopy is usually exploited by XAS, XES, synchrotron light users, etc…
This means:
Point-like sources
High yields (no need to increase them)
1 eV (Si, Mica, etc) or 2-3 eV (HAPG) resolution is very easily achievable
But what if we really need wider sources for higher statistics?
Or what if our photons come from a diffused isotropic source?

## Slide 5
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Original Motivations
A set of two slits can be used to define the geometrical parameters of the X-ray beam:
∆θ , S0 (solid lines)
∆θ’ , S0’ (dashed lines)
Is it all? Are the red lines really (only) background?
How much can we increase the effective source size keeping FWHM as low as required?
S0’ represents the effective source from which the emitted X-rays ‘’see’’ the crystal
S0 is the portion of effective source with the highest number of emitted X-rays ‘’seeing’’ the crystal
X ray source
Slits
HAPG
Signal photons
Background (?)

## Slide 6
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Pyrolitic Graphite Mosaic Crystals
Mosaic crystal consist in a large number of nearly perfect small crystallites.
Mosaicity makes it possible that even for a fixed incidence angle on the crystal surface, an energetic distribution of photons can be reflected
Increase of efficiency
(focusing) ~ 50
Loss in resolution
Pyrolitic Graphite (d = 3.354 Å):
Bending does not influence resolution and intensity
Mosaic spread down to 0.05 degree
Integral reflectivity ~ 102 higher than for other crystals
Variable thickness (efficiency)
Excellent thermal and radiation stability
Highly Annealed Pyrolitic Graphite (HAPG):
Lower mosaicity (Dq≈0.07°)
Lower reflectivity
Better resolution (FWHM)
Highly Oriented Pyrolitic Graphite (HOPG):
Higher mosaicity (Dq≈1°)
Higher reflectivity
Worse resolution (FWHM)

## Slide 7
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Von Hamos configuration
VH configuration can further improve the signal collection efficiency.
In this configuration, also the vertical dimension of the X-ray source can be exploited with no (theoretical) limits

## Slide 8
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The VOXES spectrometer @ LNF
0-10 mm opening
26 mm circular frame
< 1 mm sensitivity
Slits
(STANDA 10MAOS10-1)
Dectris Ltd
MYTHEN2 detector:
32 x 8 mm surface
640 channels  50 mm resolution
4-40 keV range
Working @ room temperature

## Slide 9
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The Semi Von Hamos configuration
Position detector rotation allows for a wider dynamic range, higher detection efficiency and lower cross talk or charge sharing
Position detector rotation does not worsen energy resolution

## Slide 10
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Calibration procedure and function
For small q values (sinq ≈ q) or small ranges, is it also possible to calibrate with a linear function

## Slide 11
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Effects of the HAPG mosaicity
Mosaic spread is worsening resolution (as expected…)
…but why not increasing the rate???

## Slide 12
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Effects of the crystal mosaicity
Mosaicity should be chosen depending on the application
Number of crystallites per mm3 is of the same order of magnitude in every crystal
The number of effectively used crystallites is given by the integral of the mosaicity spread within the energy range of interest:

## Slide 13
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Effects of the crystal thickness
The resolution worsening effect induced by the crystal thickness is not anymore predominant; on the contrary, the resolution broadening induced by the source size is more important.
Which is the effect of HAPG thickness?

## Slide 14
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
VOXES Characterization: results
Given the energy and rc it is alwasy possible to find the optimal configuration to obtain the best peak position precision
Valid for all energies
(tested for 6-20 keV)
In the limit of a background free pure gaussian peak, the precision is related to the resolution via:
Dq (°)

## Slide 15
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
VOXES Characterization: results
700 eV dynamic range (one shot)
Resolution still at 6 eV level (s)
>1 keV dynamic range
Resolution still at 1,2 ‰
Almost 2 mm source size
Did we understand the geometry properly?
How can we test the actual source size?
Do we get what we should?
X-ray tracing simulations

## Slide 16
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Ray Tracing Simulations: spectra reconstruction
Peak position and s well reproduced
Source dimensions in agreement with the horizontal and vertical expected values
RT
Exp.
Shadow + XOP code embedded in OASYS framework
(EOS26 school)

## Slide 17
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Ray Tracing Simulations: reflection efficiencies
Reflection efficiencies have been measured for Cu (up) and Fe (down) with a Silicon PinDiode as reference detector
RT simulation have been realized with accurate Cu and Fe lineshapes
Reflection efficiencies have been then computed with RT simulations and compared to measured ones
The combination of these two studies fully validates the XOP and SHADOW based ray tracing simulations for HAPG mosaic crystals.

## Slide 18
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The VOXES facility @ LNF
High resolution (FWHM ~ 0,1%) Bragg spectroscopy line (4-20 keV)
Large acceptance and efficiency line with a PinDiode detector (2-40 keV)
New background-free source box made of Teflon
Set of HAPG/HOPG and Si(220) crystals with various curvature radii
ANDOR DY934P-BR-DD-T2 1024x1024 pixels CCD
Pixel size: 13x13 mm
Timepix 256x256 pixels
Pixel size: 55x55 mm
Amptek X123 SiPin Detector
6 mm2 active area
Mo/W X-ray tubes

## Slide 19
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Bayesian Auto-alignment procedure
Developed by S. Manti @ LNF

## Slide 20
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The VOXES facility @ LNF
Fe + Co spectrum obtained with the high resolution Bragg spectrometer line
FeKb1,3
FeKb’
CoKa1
CoKa2
Multi-target spectrum obtained with the high efficiency line

## Slide 21
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The MITIQO project

## Slide 22
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The MITIQO project
Metal’s State (Oxidation) reduces the quality and it is driven by Metals

## Slide 23
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The MITIQO project
FERROFOLIN 40 mg/15 ml + 0,185 mg/15 ml
Goal: distinguish oxidized liquid Fe from Fe0

## Slide 24
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The MITIQO project

## Slide 25
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The MITIQO project
According to several studies, Principal Components Analysis applied to relative metal concentrations provide hints on geographical origin of wines
MITIQO high efficiency line detection capabilities have been calibrated with diluted samples of FERROFOLIN

## Slide 26
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
The MITIQO project
Hints of clusterization! -> Reduce unwanted contamination in the setup

## Slide 27
A. Scordo - Advances in X-ray and Gamma RAy DEtectors and applications (XGRADE 2026) 24-26/03/2026, Palermo
Conclusions
VOXES is now a well established facility at the INFN Laboratories of Frascati
We are also open to external applications/collaborations
We also welcome suggestions and ideas for possible applications
Come visit us in Frascati