## Slide 1
Embedding Position Sensitivity in the Detector with ANNA: an Analog Neural Network ASIC
M.Amadori, M.Ronchi, G.Borghi, M.Carminati and C.Fiorini
Politecnico di Milano and INFN, Milano, Italy
PSD14: The 14th International Conference on Position Sensitive Detectors
London (UK), 03-09-26

## Slide 2: Outline
Motivation: edge-processing in large detection systems (Total-Body PET)
Neural networks for event reconstruction in monolithic scintillators
The ANNA ASIC: V1 and V2
Experimental measurements
Future application for LGADs readout
Conclusions

## Slide 3: Motivation: edge-processing in large detection systems (TB PET)
(Katal S. et al. - Advantages and Applications of Total-Body PET Scanning. Diagnostics, 2022)
Advantages
Full-body coverage
Higher sensitivity
Faster scans
Lower radiation dose
Disadvantages
High number of channels
Complex system implementation
High power
Trend towards Total Body PET systems
Total time 20min
Total time 2-20min
Low-power, edge-processing of detector signals required
(also for other applications with large number of detector modules)

## Slide 4: Monolithic Scintillator Detector and Neural Network Reconstruction
Y
X
Monolithic Scintillator: simplifies the scintillator structure (vs. pixelated), facilitate DOI, at cost of possible pile-up
SiPMs Readout: now quite standard vs. PMTs as more compact and insensitive to magnetic fields
Neural Network Algorithm: better and faster estimation of g-ray interaction in the crystal, vs. Anger and other methods
Position Reconstruction
Algorithm (Neural Network)
Photodetectors + front-end
Scintillator
crystal

## Slide 5: Traditional Digital Processing vs Analog Processing (ANNA)
Analog Neural Network ASIC (ANNA)
FPGA
64 ADCs
Traditional Processing
Computer
Image Reconstruction Algorithm
Image Reconstruction Algorithm
Analog Processing
Computer
Image Reconstruction Algorithm
Image Reconstruction Algorithm
x 64
digital
signals
Only 2
analog
signals
Y
X
Y
X
2 ADCs

## Slide 6: Traditional Digital Processing vs Analog Processing (ANNA)
Analog Neural Network ASIC (ANNA)
Analog Processing
Computer
Image Reconstruction Algorithm
Image Reconstruction Algorithm
Only 2
analog
signals
Y
X
2 ADCs
Analog operations performed directly on analog signals coming from photodetectors
No need for ADC and FPGA for embedded processing
Interaction coordinates (X,Y) directly at the output of the ASIC
Compatibility with Front-end ASIC for SiPMs current processing (Amplitude, Timing)
Same approach possible also for other segmented solid-state detectors
Analog approach (as memristors) to in-memory computing (vs. digital)

## Slide 7: ANNA: Analog Multiply and Accumulate (MAC) principle
Features:
Capacitive cross-bar array
Charge sum by an integrator
Activation function implemented in the integrator
Cf

## Slide 8: ANNA (Analog Neural Network ASIC): V1 version (proof-of-concept)
weight
MAC:
Switches: allows to circulate the charge in two directions (positive and negative weights)
Activation function: clipped ReLU implemented by the op-amp bias limits
Quantization-Aware Training
4 bits weight
(+ 1 bit sign)
Cell i,j

## Slide 9: Characterization of ASIC V1 (CMOS 0.35 µm, 24 mm2)
Microcontroller Board
Carrier Board
Mother Board
X
Y
Layer 1
Layer 2
3
Digital (SPI)
Integrators
Graphic interface:
ASIC programming
Inputs delivery
Data acquisition
Inner nodes

## Slide 10: ANNA V1 Performance Characterization
Grid of 121 simulated points (ANTS2) provided as ASIC voltage inputs
8x8 SiPMs signal [V]
x [mm]
y [mm]
X
Y
Hardware-aware model, used during training, takes into account parasitics, offsets and charge injections, measured during stand-alone test of the ASIC
Original, simplest model
Hardware-aware model

## Slide 11: ANNA V2
New improved architecture (including recursive network)
Parasitics: more repeatable and easier to characterize, access to raw analog outputs of each layer
Complexity: simplified neuron design, fewer steps needed for MAC operations
 | ANNA V1 | ANNA V2
Process node | 350 nm | 110 nm
Supply voltage | 3.3 V | 1.2 V
Area | 23.5 mm2 | 9.5 mm2
Neural network config | Fixed | Programmable
Neural network | 64 inputs, 2 outputs 64 x 20 x 20 x 2 (3 weighted layers) | 69 inputs, 32 outputs 69 x 32 x 32 x 32 x 32 x 32 (5 weighted layers)
Parameters (weights and biases) | 1762 (5 bits – 31 analog levels) | 6496 (6 bits – 63 analog levels)
Inference execution time | 4.5 µs | < 1.5 µs
Throughput | 0.77 GOP/s | > 10 GOP/s
Efficiency | 0.05 TOPs/W | > 5 TOPs/W

## Slide 12: ASIC V2 design: new neuron schematic
New Multiplying Element (C-2C DAC)
Output capacitance always 2C
Charge injection reduction by design
Power reduction due to C reduction (30fF)
New Integrator (inverter-based)
6-bit weights precision
Variable feedback capactitor (15 values)
Programmable power consumption
More compact layout
Tanh-like activation function

## Slide 13: ASIC V2 architecture: recursive network
Standard network (as in V1)
Recursive network
High Area occupation
No Flexibility
Simple Design
Low Area
High Flexibility
Difficult Design
32 outputs are feed back to the inputs
Weights (stored in SRAM) are iteratively updated each layer

## Slide 14: ANNA V2 – Architecture and Layout
Recursive
Non
Recursive
35 inputs
35 inputs
ANNA V2
Single reusable layer
32 outputs
Fully programmable neural network (layers and neurons) and parameters (weight, biases, activation functions)
Configurable power consumption and power cycling
Programmable layer duration and synchronization
Trigger based input processing and options to improve throughput

## Slide 15: 64x20x20x2 Network - Results
Simulated events on a 51mm x 51mm LYSO scintillator
Linear mapping between mm and Voltages
Grid of gamma-ray events emulated with ANTS3
Model vs ASIC Results
Mean Absolute Error (MAE)
Tensorflow Model | ~ 0.9 mm
ANNA | ~ 1.1 mm
Scintillator
51mm
51mm
10mm

## Slide 16: 64x20x20x2 Network – Noise Analysis
Total SiPM Area
51mm
51mm
Noise of 10000 events
Cf = 120fF for every layer
Spatial Precision of

## Slide 17: Future Developments – LGADs signals processing with ANNA processing
1
Resistive LGADs (DC-RSD) exploits charge sharing to achieve excellent spatial resolution, much smaller than the pixel size, reducing number of readout channels, power and material budget (N.Cartiglia, NIMA, 2022)
2
Neural networks (NNs) allows a fast and energy-efficient reconstruction of events in DC-RSD
3
ANNA implements the NN directly on-chip
We aim to develop a demonstrator for a new, low-power 4D tracker based on LGADs with embedded reconstruction of the event features (position, timing) by ANNA
Activity in collaboration with INFN-TO/FI, UNI-TO, UNI-Pe

## Slide 18: Conclusions and future work
ANNA brings neural-network processing directly into the analog detector readout chain, avoiding prior signal digitalization.
The new ANNA V2 architecture combines high programmability, compact implementation and ultra-low-power analog computation.
Experimental results demonstrate accurate position reconstruction, not too far from the TensorFlow model.
This approach enables true edge processing in large-scale detector systems, reducing data throughput and system-level complexity.
Beyond monolithic scintillators, ANNA opens the way to “intelligent detectors” with embedded reconstruction, including future low-power LGAD trackers.
Next steps:
Construction (ongoing) and test of PETRARCA: a PET detector with embedded, ANNA-based position recontruction
Development of a proof-of-concept prototype of resistive LGAD detector with ANNA processing

## Slide 19: Relevant Publications
S. Di Giacomo, et al., "Implementing an Integrated Neural Network for Real-Time Position Reconstruction in Emission Tomography with Monolithic Scintillators“, IEEE TRANSACTIONS ON RADIATION AND PLASMA MEDICAL SCIENCES, VOL. 8, NO. 5, MAY 2024.
S. Di Giacomo, et al., «Experimental Validation of ANNA: Analog Neural Network ASIC for Event Positioning in Monolithic Scintillation Detectors», IEEE Transactions on Radiation and Plasma Medical Sciences, 2025.
M.Amadori, et al., «A Compact Neuron Implementation for a Capacitive Analog Neural Network for In-Sensor Processing», 2024 31st IEEE International Conference on Electronics, Circuits and Systems (ICECS).
M. Ronchi, et al., «Design, Implementation and Analysis of an Integrated Switched Capacitor Analog Neuron for Edge Computing AI Accelerators», IEEE Transactions on Circuits and Systems I: Regular Papers, 2025.
Acknowledgements This work was supported by the IANNET project (Italian MUR) and by the ANNA project (Italian INFN)

## Slide 20: The RADLAB team
Thank you for your attention!

## Slide 21
Backup

## Slide 22: Analog Neuron Operation
Charge-redistribution approach
Charge only C with Vin to minimize energy consumption
Charge is redistributed among all capacitors closing Ss
Only the capacitors corresponding to the desired weight are connected to the integrator, while the others are disconnected by means of their respective switch.
Positive weight
Negative weight
Step 1
Step 2
Step 3
The output is

## Slide 23: ANNA V2 – Multiply and accumulate
The multiplication phase is implemented through a 2-2C architecture, providing a weighted charge to the integrator.
Compared to the previous implementation:
Faster (less switching steps)
Much more compact (less capacitors: above 4 bits the C-2C is more area-efficient than binary weighted)
More accurate (consistent parasitics, constant output impedance)

## Slide 24: ANNA V2 – new architecture
Two computation regions:
Full matrix (69+bias)×32: first layer
Recursive region ((32+bias)×32): subsequent layers
The recursive region is isolated from inputs during operation
Weights are iteratively updated each layer
All weights are stored in SRAMs in the ASIC

## Slide 25: Detector simulations and ANNA test
Detector simulation
Detector simulated in ANTS: LYSO scintillator 51 × 51 × 10 mm3 coupled to an 8 × 8 SiPM array.
Generated a training dataset (27000 events) and test dataset (4500 events).
Labeled datasets: the input data are the SiPM signals and the associated labels are the coordinates of the interaction point in the crystal.
Chosen neural network topology: 64 × 20 × 20 × 2.
Model of the network trained and tested in Tensorflow.
Expected (x,y)
LYSO
8x8
SiPMs
10 mm
51 mm
51 mm
TEST PHASE
ANTS simulation
Untrained model in TENSORFLOW
TRAINING PHASE
Training
dataset
ANTS simulation
Trained model in TENSORFLOW
Test dataset:
real (x,y)
Measured (x,y)
Real network implemented by ANNA
Weightsf

## Slide 26: Results: Edge Detection
Edge detection:
Image provided on the analog domain, 3x3 patches
Analog edge filter on the ASIC
Analog output sampled
0
-7
0
-7
-7
0
-7
0
28
32
W =
.
Convolutional Filter
ANNA
Input Image
Output Image
1

## Slide 27: ANNA V2 Summary
Power Consumption [W]
Computation Speed [GOP/s]
ANNA V2
Comparison of Neural Accelerators