## Slide 1
Results from the
HEPTO Combined Function Permanent Magnet
Alex Hinton, Ben Shepherd (STFC Daresbury Laboratory)
Dušan Topalović (Kyma S.p.A.) Abolfazl Shahveh (Diamond Light Source)
Particle Accelerators and Beams Conference
Lancaster
01/07/2026

## Slide 2: Introduction
Upgrade
2.3 kW main coils + water cooling
80 W trim coils
I.FAST: EU funded project to develop new sustainable accelerator technologies and favour partnerships with industry.
Fourth generation synchrotron light sources make use of multi-bend achromat lattices to reduce beam emittance and increase radiation brightness.
Diamond-II: upgrade from double bend to 6-bend achromat.
Lattice includes 48 transverse gradient (DQ) magnets: ~108 kW.
Task 11.3 – Manufacture and test prototype permanent magnet (PM) based DQ magnet as energy saving alternative to Diamond-II design.

## Slide 3: HEPTO (Hybrid Electromagnet-Permanent Magnet Tuneable Optics)
Parameter | Value | Units
Central dipole field | -0.6951 | T
Central gradient | 33.179 | T/m
Effective dipole magnetic length | 870 | mm
Effective gradient magnetic length | 850 | mm
Integrated dipole field | -0.6047 | T.m
Integrated gradient | 28.1857 | T
Bending radius | 16.795 | m
Good field radius (GFR) | 7 | mm
Integrated field quality (ΔB/B) in GFR | 0.05 | %
Gradient tuning range | ±2.5 | %
Temperature stability (3°C range) | ±0.01 | %
Main trim coil pair
FeNi thermal shunt alloy
Auxiliary trim coil pair
45M6 NdFeB permanent magnet blocks
Grade 1010 steel yoke and poles

## Slide 4: Mechanical Design Completed by Kyma
2. Two screws for opening gap between auxiliary pole and aluminium plate. Brass shims can be added/removed to adjust vertical position of auxiliary pole.
Gap is pre-shimmed with 300 µm thickness.
Shim thicknesses in range 10-200 µm.
Vertical gap between main poles defined by pre-shimmed spacer.
In-built mechanical shimming methods allow adjustment of field strength and homogeneity.

## Slide 5
Magnet Assembly
60 kg peak force
2000 kg peak force

## Slide 6: Magnetic Measurements
Senis Type-H Hall probe
3-axis motion stage
Field measurements performed in dedicated magnet measurement laboratory at Daresbury.

## Slide 7: Field Multipoles
Magnetic fields described in terms of “multipoles”.
Complex representation of magnetic field
Complex representation of point in magnet aperture
Reference radius
Complex multipole coefficients
n | Name | Function
1 | Dipole | Bending
2 | Quadrupole | Focussing
3 | Sextupole | Chromatic aberrations
≥4 | Higher order | Field errors
HEPTO pole tips curved – dipole constant along trajectory.
Multipoles evaluated by measuring points around circular good field region and integrating over curved nominal electron beam trajectory.

## Slide 8
Nominal beam trajectory
Br
Hall probe point measurement
Fourier transform

## Slide 9: Initial Measurements
DIPOLE
QUADRUPOLE
SEXTUPOLE
 | Integrated Gradient / T | Integrated gradient error / % | Integrated Homogeneity ΔB/B / %
Requirement | 28.1857 |  | 0.05
Measured | 28.070 ± 0.002 | -0.41 | 0.082±0.005

## Slide 10: Shimming
Vertical gaps between main and auxiliary poles adjusted via addition / removal of brass shims.

## Slide 11: Post-shimming
 | Integrated Gradient / T | Integrated gradient error / % | Integrated Homogeneity ΔB/B / %
Requirement | 28.1857 |  | 0.05
Pre-shimming | 28.070 ± 0.002 | -0.41 | 0.082±0.005
Post-shimming | 28.196 ± 0.001 | +0.04 | 0.029±0.001

## Slide 12
Trim coils
Target = independent tuning of dipole and quadrupole by ±2.5 %
19.5 W
29 W

## Slide 13
Tuning Range
Trim coil currents limited to ±5 A.
Quadrupole tuned whilst keeping dipole field constant – emittance measurements.
Maximum achievable tuning range is ±1.1 %.
Less than target range of ±2.5 %.
Electromagnet tuning is inefficient due to low magnetic permeability of permanent magnets in the magnetic circuit.

## Slide 14
Thermal Stability
+0.17 % / degree
-0.06 % / degree
FeNi thermal shunt alloy glued to PM blocks

## Slide 15
HEPTO magnet built and measured as part of I.FAST WP11.3.
Demonstrator of sustainable alternative to electromagnets.
Image © STFC John Dawson
Conclusion
Parameter | Target | Achieved | Units
Integrated gradient | 28.1857 | 28.196 ± 0.001 | T
Integrated field quality (ΔB/B) in GFR | 5E-4 | (2.9±0.1)E-4 | 
Gradient tuning range | ±2.5 | ±1.1 | %
Trim coil power | <2.3 | 0.05 | kW
Temperature stability (3°C range) | ±0.01 | +0.51 | %
Next steps: further collaboration between STFC and Kyma through EPITA.
Development of scalable ZEPTO tuneable permanent magnet technology.

## Slide 16
Notes: The abstract pattern can be removed or repositioned if required. Be careful to ‘Send to Back’ so that it does not obscure any important information.

## Slide 17
@STFC_matters
Science and Technology Facilities Council
Science and Technology Facilities Council
Acknowledgements:
Ben Shepherd (STFC)
Alfie Shahveh (Diamond Light Source)
Dušan Topalović, Tadej Milharcic, Mirko Kokole, Jure Počkar (Kyma S.p.A.)
Notes: The abstract pattern can be removed or repositioned if required. Be careful to ‘Send to Back’ so that it does not obscure any important information.