## Slide 1: Characterizing Vibrations in a Dry Dilution Refrigerator for Low Noise Experiments
Jackson Fisher, Wouter Van De Pontseele, Lucia Steinke

## Slide 2: Importance of Vibrations
Causes errors in superconducting qubits and degrades performance of sensitive particle detectors
Affected experiments can include neutrino-nucleus scattering (RICOCHET), neutrinoless double beta-decay search (CUORE), dark matter search (SuperCDMS)...
1
K. J. Vetter et al., “Improving the performance of cryogenic calorimeters with nonlinear multivariate noise cancellation algorithms,” The European Physical Journal C, vol. 84, no. 3, Mar. 2024.
Augier, C., Beaulieu, G., Belov, V. et al. Ricochet Progress and Status. J Low Temp Phys 212, 127–137 (2023).
H. P. Binney et al., “Distinguishing Types of Correlated Errors in Superconducting Qubits,” Physical Review Applied, vol. 26, no. 2, Art. no. 024025, Aug. 2026, doi: 10.1103/j235-47t6.

## Slide 3: Colorado Underground Research Institute (CURIE)
Future physics lab located under ~200m of rock
Measured 700x muon attenuation
Vibrations, electromagnetic, and radon noise also important
2
Talk at CPAD: "The Colorado Underground Research Institute: Status and Site Characterization", Luke Wanner

## Slide 4: Maybell Quantum Fridge Plus
~300 K
~50 K
~9 K
~700 mK
~200 mK
Mixing Chamber, 10 mK
1, 2, 3... Hz Vibrations
Internal Suspension, 3-Point Kevlar Rods
Minus-K Isolators
3
A4k Mu-Metal Magnetic Shield
Pulse Tubes
~3 K
Notes: Talk about what dilution refrigerators are, say that they are refrigerators capable of hitting millikelvin temperatures at the mixing chamber, and as shown in this diagram, willl have intermediate stages which progressively get colder as you go down.
They are used for operating quantum devices which require low temperature environments, dry dilution refrigerators in specific have become popular due to the use of the pulse tube cryocooler.
This pulse tube creates a closed loop cooling system which makes the much easier and cheaper to operate, but induces significant vibrational noise which can hinder the performance of any sensitive experiments in the fridge.
The purpose of my project is to characterize the vibrational noise of a maybell quantum fridge plus at room and low temperature using several accelerometers.
These vibrations are important for mitigating vibrations for future low noise experiments, or understanding what contribution of false events the vibratoins contribute in these experiments. Affected technologie scould be superconducting qubits or sensitive detectors, such as cryogenic bolometers for neutrion sensing or sensors used for the search for dark matter.

## Slide 5
3
Minus-K
Maybell Quantum Fridge Plus
~300 K
~50 K
~9 K
~700 mK
~200 mK
Mixing Chamber, 10 mK
1, 2, 3... Hz Vibrations
Internal Suspension, 3-Point Kevlar Rods
Minus-K Isolators
A4k Mu-Metal Magnetic Shield
Pulse Tubes
~3 K
Notes: Talk about what dilution refrigerators are, say that they are refrigerators capable of hitting millikelvin temperatures at the mixing chamber, and as shown in this diagram, willl have intermediate stages which progressively get colder as you go down.
They are used for operating quantum devices which require low temperature environments, dry dilution refrigerators in specific have become popular due to the use of the pulse tube cryocooler.
This pulse tube creates a closed loop cooling system which makes the much easier and cheaper to operate, but induces significant vibrational noise which can hinder the performance of any sensitive experiments in the fridge.
The purpose of my project is to characterize the vibrational noise of a maybell quantum fridge plus at room and low temperature using several accelerometers.
These vibrations are important for mitigating vibrations for future low noise experiments, or understanding what contribution of false events the vibratoins contribute in these experiments. Affected technologie scould be superconducting qubits or sensitive detectors, such as cryogenic bolometers for neutrion sensing or sensors used for the search for dark matter.

## Slide 6
3
Kevlar Rods
Maybell Quantum Fridge Plus
~300 K
~50 K
~9 K
~700 mK
~200 mK
Mixing Chamber, 10 mK
1, 2, 3... Hz Vibrations
Internal Suspension, 3-Point Kevlar Rods
Minus-K Isolators
A4k Mu-Metal Magnetic Shield
Pulse Tubes
~3 K
Notes: Talk about what dilution refrigerators are, say that they are refrigerators capable of hitting millikelvin temperatures at the mixing chamber, and as shown in this diagram, willl have intermediate stages which progressively get colder as you go down.
They are used for operating quantum devices which require low temperature environments, dry dilution refrigerators in specific have become popular due to the use of the pulse tube cryocooler.
This pulse tube creates a closed loop cooling system which makes the much easier and cheaper to operate, but induces significant vibrational noise which can hinder the performance of any sensitive experiments in the fridge.
The purpose of my project is to characterize the vibrational noise of a maybell quantum fridge plus at room and low temperature using several accelerometers.
These vibrations are important for mitigating vibrations for future low noise experiments, or understanding what contribution of false events the vibratoins contribute in these experiments. Affected technologie scould be superconducting qubits or sensitive detectors, such as cryogenic bolometers for neutrion sensing or sensors used for the search for dark matter.

## Slide 7: Piezoelectric Sensors and Geophones
2271A (Lowtemp)
PCB Piezotronics
393B05 (Roomtemp) PCB Piezotronics
HG-2 Geophones
HGS Products
4

## Slide 8: Sensor Operation and Specifications
5
Seismic
mass
Piezo crystal
Rigid base
Charge
amplifier
Piezoelectric
High Impedance
Charge

## Slide 9: Sensor Operation and Specifications
5
Seismic
mass
Piezo crystal
Rigid base
Charge
amplifier
N
S
Spring
Magnet
Case
Signal out
Piezoelectric
Geophone
High Impedance
Charge
Coil

## Slide 10: Sensor Operation and Specifications
Sensor | Lowtemp | Roomtemp | Geophone
Frequency Range | DC to 4 kHz | DC to 500 Hz | DC to 100 Hz
Sensitivity Deviation due to Temperature | -15 % at ~5 K | Not cryo-compatible | Unknown
5
Seismic
mass
Piezo crystal
Rigid base
Charge
amplifier
N
S
Spring
Magnet
Case
Signal out
Piezoelectric
Geophone
High Impedance
Charge
Coil

## Slide 11: Accelerometer Measurement Chain
Thermalized
wiring tree
Geophones
Lowtemp Sensor
Mixing Chamber
Roomtemp
Sensor
6
Notes: So the measurement chain I created for this experiment includes sensor to signal conditioner to digitizer. A charge converter is required for the 2271A, like I previously stated. I use the low noise coaxial cabling from PCB Piezotronics, which is necessary for decreasing electromagnetic noise from entering the setup. In the fridge, the 2271A and the geophones are mounted at the mixing chamber using a stud that the sensor screws onto, and their signals run up the fridge's thermalized wiring tree. The room-temperature sensor is attached to the side of the fridge. All other electronics are outside the fridge.

## Slide 12: Accelerometer Measurement Chain
Thermalized
wiring tree
Geophones
Lowtemp Sensor
Mixing Chamber
Roomtemp
Sensor
Low-noise
Amplifier
LNA10
Charge
Converter
422E11
Signal Conditioner
480E09
Signal
Conditioner
480E09
Digitizer
(Spectrum
M2p.5911)
Low-noise
chassis
1 kHz Lowpass Filter
2 kHz Lowpass Filter
6
Notes: So the measurement chain I created for this experiment includes sensor to signal conditioner to digitizer. A charge converter is required for the 2271A, like I previously stated. I use the low noise coaxial cabling from PCB Piezotronics, which is necessary for decreasing electromagnetic noise from entering the setup. In the fridge, the 2271A and the geophones are mounted at the mixing chamber using a stud that the sensor screws onto, and their signals run up the fridge's thermalized wiring tree. The room-temperature sensor is attached to the side of the fridge. All other electronics are outside the fridge.

## Slide 13: Data Analysis
Lowtemp Flat Sensitivity: 0.1249 V / (m/s2)
Roomtemp Flat Sensitivity: 1.042 V / (m/s2)
Sensor Calibration
Amplitude Spectral Density (Welch's Method)
RMS Displacement
7

## Slide 14: Data Analysis
Lowtemp Flat Sensitivity: 0.1249 V / (m/s2)
Roomtemp Flat Sensitivity: 1.042 V / (m/s2)
Sensor Calibration
7
Amplitude Spectral Density (Welch's Method)
RMS Displacement
Interested in this range

## Slide 15: Data Analysis
Lowtemp Flat Sensitivity: 0.1249 V / (m/s2)
Roomtemp Flat Sensitivity: 1.042 V / (m/s2)
Sensor Calibration
7
Amplitude Spectral Density (Welch's Method)
RMS Displacement
Interested in this range
1
Average FFTs, apply Hanning windowing (U)
2
Calculate ASD

## Slide 16: Data Analysis
Lowtemp Flat Sensitivity: 0.1249 V / (m/s2)
Roomtemp Flat Sensitivity: 1.042 V / (m/s2)
Sensor Calibration
1
Average FFTs, apply Hanning windowing (U)
2
Calculate ASD
3
4
Calculate Displacement
ASD
Calculate RMS Displacement
7
Amplitude Spectral Density (Welch's Method)
RMS Displacement
Interested in this range

## Slide 17: Low Temperature Measurement
8
Notes: These are some initial measurements of our sensors on the floor of our lab. The 2271A and 393B05 agree on peak locations quite well, but disagree at lower frequencies. Like I said before, the 2271A is going to be much more susceptible to noise, and the lab is a very noisy environment so this may be raising the noise floor. I have found that triboelectric noise, capacitancce from the sensor and cabling, and the 1/f electronic noise of the charge converter are big contributors for the 2271A. So, at lower frequencies the 2271A is hitting a noise floor and starts to match better with the 393B05 at higher frequencies. Most of the largest peaks are at 60 Hz and the harmonic sof 60 Hz, so this is likely just electric noise, which makes sense because there is nothing really shielding the equipment. If the 2271A sensor is put in a fridge which is meant to mitigate vibrations and shield from electromagnetic noise, then I think the noise floor would decrease even more. The pulse tube mainly induces vibrations at 1 Hz and then harmonics of 1 Hz, so its unfortunate that the low frequency noise has the most noise, but this could change once the sensor is put in a fridge.

## Slide 18: Low Temperature Measurement
8
From Geophones, RMS displacement from 10 Hz to 200 Hz:
Z Axis = 86.7 nm
X Axis = 209.4 nm
Y Axis = 70.2 nm
Notes: These are some initial measurements of our sensors on the floor of our lab. The 2271A and 393B05 agree on peak locations quite well, but disagree at lower frequencies. Like I said before, the 2271A is going to be much more susceptible to noise, and the lab is a very noisy environment so this may be raising the noise floor. I have found that triboelectric noise, capacitancce from the sensor and cabling, and the 1/f electronic noise of the charge converter are big contributors for the 2271A. So, at lower frequencies the 2271A is hitting a noise floor and starts to match better with the 393B05 at higher frequencies. Most of the largest peaks are at 60 Hz and the harmonic sof 60 Hz, so this is likely just electric noise, which makes sense because there is nothing really shielding the equipment. If the 2271A sensor is put in a fridge which is meant to mitigate vibrations and shield from electromagnetic noise, then I think the noise floor would decrease even more. The pulse tube mainly induces vibrations at 1 Hz and then harmonics of 1 Hz, so its unfortunate that the low frequency noise has the most noise, but this could change once the sensor is put in a fridge.

## Slide 19: Low Temperature Measurement
8
From Geophones, RMS displacement from 10 Hz to 200 Hz:
Z Axis = 86.7 nm
X Axis = 209.4 nm
Y Axis = 70.2 nm
Calibration at cryogenic temperature not confirmed
Notes: These are some initial measurements of our sensors on the floor of our lab. The 2271A and 393B05 agree on peak locations quite well, but disagree at lower frequencies. Like I said before, the 2271A is going to be much more susceptible to noise, and the lab is a very noisy environment so this may be raising the noise floor. I have found that triboelectric noise, capacitancce from the sensor and cabling, and the 1/f electronic noise of the charge converter are big contributors for the 2271A. So, at lower frequencies the 2271A is hitting a noise floor and starts to match better with the 393B05 at higher frequencies. Most of the largest peaks are at 60 Hz and the harmonic sof 60 Hz, so this is likely just electric noise, which makes sense because there is nothing really shielding the equipment. If the 2271A sensor is put in a fridge which is meant to mitigate vibrations and shield from electromagnetic noise, then I think the noise floor would decrease even more. The pulse tube mainly induces vibrations at 1 Hz and then harmonics of 1 Hz, so its unfortunate that the low frequency noise has the most noise, but this could change once the sensor is put in a fridge.

## Slide 20: Noise Correlation
Exploring further vibration mitigation
Finding optimal compressor frequency combination for vibration and cooling power
Other tests can include tuning Minus-K, decoupling rotary valves, positional measurement
9

## Slide 21: SQUID Cross Correlation
Vibration
Vibration sensor
x(t)
B(x)
Φ(t)
Power spectral density
Pulse-tube
peaks
10
Notes: The idea is to use the vibration sensors as a reference for the SQUIDs. Pulse-tube vibration shakes everything on the mixing chamber plate, including the SQUID. If the SQUID sits in any residual field that isn't perfectly uniform, moving it by a small distance changes the flux through its loop, roughly the loop area times the field gradient times the displacement. So mechanical motion shows up in the SQUID as flux noise. That means the SQUID's power spectral density should show the same pulse-tube harmonics that the accelerometers and geophones see. Electromagnetic noise, like RF or line pickup, couples into the SQUID but not into the vibration sensors. Cross-correlating the two signals, in practice the coherence between them, separates the two: frequencies where the coherence is close to one are vibration, and peaks that only the SQUID sees, with low coherence, are electromagnetic.

## Slide 22
Vibration
EM noise
Vibration sensor
x(t)
B(x)
Φ(t)
Power spectral density
Pulse-tube
peaks
EM peak
DC SQUID
10
SQUID Cross Correlation
Notes: The idea is to use the vibration sensors as a reference for the SQUIDs. Pulse-tube vibration shakes everything on the mixing chamber plate, including the SQUID. If the SQUID sits in any residual field that isn't perfectly uniform, moving it by a small distance changes the flux through its loop, roughly the loop area times the field gradient times the displacement. So mechanical motion shows up in the SQUID as flux noise. That means the SQUID's power spectral density should show the same pulse-tube harmonics that the accelerometers and geophones see. Electromagnetic noise, like RF or line pickup, couples into the SQUID but not into the vibration sensors. Cross-correlating the two signals, in practice the coherence between them, separates the two: frequencies where the coherence is close to one are vibration, and peaks that only the SQUID sees, with low coherence, are electromagnetic.

## Slide 23
Vibration
EM noise
Vibration sensor
Cross- correlate
x(t)
B(x)
Φ(t)
Power spectral density
Calculated Coherence
Pulse-tube
peaks
EM peak
DC SQUID
10
SQUID Cross Correlation
Notes: The idea is to use the vibration sensors as a reference for the SQUIDs. Pulse-tube vibration shakes everything on the mixing chamber plate, including the SQUID. If the SQUID sits in any residual field that isn't perfectly uniform, moving it by a small distance changes the flux through its loop, roughly the loop area times the field gradient times the displacement. So mechanical motion shows up in the SQUID as flux noise. That means the SQUID's power spectral density should show the same pulse-tube harmonics that the accelerometers and geophones see. Electromagnetic noise, like RF or line pickup, couples into the SQUID but not into the vibration sensors. Cross-correlating the two signals, in practice the coherence between them, separates the two: frequencies where the coherence is close to one are vibration, and peaks that only the SQUID sees, with low coherence, are electromagnetic.

## Slide 24: Conclusions
Bi-Weekly vibrational noise meetings:
Functional sensor setup, looking for ideas to improve calibration and sensor noise
Cross-correlation functions with SQUID magnetometer, transition edge sensors (TES), superconducting tunnel junctions (STJ), superconducting qubits to identify and mitigate noise
11
TES Picture
STJ Picture
Notes: My setup with 2271A and 393B05 shows promising resolution for seeing the pulse tube vibrations. Much of the noise from the 2271A can be attributed to triboelectric, 1/f charge converter noise.
In the future I will measure in the maybell quantum fridge plus for a quieter environment, I will create thermalization anchors so that I can use my low noise coaxial cabling instead of the fridges built in wiring tree.
Finally, I plan to cross correlate these measruements with a SQUID magnetometer so I can extrapolate my characterization out to all frequencies and fully understand the vibrational noise floor.

## Slide 25: Acknowledgements
jtfisher@mines.edu
https://qtsf.mines.edu/
Vibrational noise
meeting Interest:

## Slide 26: References
[1] K. J. Vetter et al., “Improving the performance of cryogenic calorimeters with nonlinear multivariate noise cancellation algorithms,” The European Physical Journal C, vol. 84, no. 3, Mar. 2024. doi:10.1140/epjc/s10052-024-12595-y
[2] Augier, C., Beaulieu, G., Belov, V. et al. Ricochet Progress and Status. J Low Temp Phys 212, 127–137 (2023). https://doi.org/10.1007/s10909-023-02971-5
[3] H. P. Binney et al., “Distinguishing Types of Correlated Errors in Superconducting Qubits,” Physical Review Applied, vol. 26, no. 2, Art. no. 024025, Aug. 2026, doi: 10.1103/j235-47t6.
[4] - curie paper / supercdms
[5] E. Olivieri, J. Billard, M. De Jesus, A. Juillard, and A. Leder,"Vibrations on pulse tube based Dry Dilution Refrigerators for low noise measurements" vol. 858, pp. 73-79, Jun. 2017, doi:https://doi.org/10.1016/j.nima.2017.03.045.
[6] PCB Piezotronics, Vibration Fundamentals, https://www.pcb.com/contentstore/mktgcontent/linkeddocuments/Vibration/TM-VIB_PosterBook_Lowres.pdf (accessed Mar. 2, 2026).

## Slide 27: Supplemental slides

## Slide 29: 2271A Noise
Strange vibration events, some sort of electrical discharges?

## Slide 30: Calculating Coherence between Lowtemp and Geophones
2271A and random geophone:
Two random geophones:

## Slide 31: Compressor Frequency Sweep
Measured Displacement RMS for each 10 minute compressor frequency change
The typical variance in displacement RMS is a +/- 2 nm, so these values should be fairly accurate.

## Slide 32: Geophone Analysis
Geophone Calibration Data ( V per m/s vs frequency)
Numpy.interp function used to create an interpolation for sensitivity (V / m/s) at a given frequency
Calculated PSD as shown previously, then converting to acceleration PSD
In the above equation, the sensitivity converts voltage to m/s, and the other term converts velocity PSD to acceleration PSD.

## Slide 33: Rolling RMS

## Slide 34: Time Series data
Geophones
2271A

## Slide 35: Room Temperature, PT Off