## Slide 1: Progress towards seamless large area X-ray and gamma-ray detectors
M. Campbell1, J. Alozy, R. Ballabriga, P. Christodoulou, X. Llopart, F. Piernas Dias, V. Sriskaran, and L. Tlustos
CERN,
EP Department
1211 Geneva 23
Switzerland
1 Honorary Professor at Glasgow University

## Slide 2: Hybrid Silicon Pixel Detectors
Standard CMOS can be used allowing on-pixel signal processing
Sensor material can be changed (Si, GaAs, CdTe..)
Notes: Mention signal generation Add X section?

## Slide 3: Gas detector readout - InGrid
Semiconductor detector is replaced with charge amplification grid
Permits lower energy events to be detected
NB: GEM foils may be used in place of the InGrid foils
If placed in a photo tube together with a MCP visible photons may be detected
(see: M. Fiorini et al 2018 J.Instrum 13 C12005)

## Slide 4: Timepix3 Photo

## Slide 5: Tiling larger areas
Wire Bonding
Target to build large area detectors by combining single chip modules.
Through-silicon vias (TSVs) are the key technology enabler.
Medipix4 and Timepix4 use on-chip interposer for bump to pixel redistribution layer (RDL).
The detector is fully sensitive: even above the peripheral circuit and I/O pads.
ASIC
PCB
Sensor
FC
FC
BGA

## Slide 6
Medipix2 (1999 -> ) | Medipix3 (2005 -> ) | Medipix4 (2016 -> )
Albert-Ludwig Universität Freiburg, Germany | Albert-Ludwig Universität Freiburg, Germany | CEA, Paris, France
CEA, Paris, France | AMOLF, Amsterdam, The Netherlands | CERN, Geneva, Switzerland
CERN, Geneva, Switzerland | Brazilian Light Source, Campinas, Brazil | DESY-Hamburg, Germany
Czech Academy of Sciences, Prague, Czechia | CEA, Paris, France | Diamond Light Source, England, UK
ESRF, Grenoble, France | CERN, Geneva, Switzerland | IEAP, Czech Technical University, Prague, Czeciah
IEAP, Czech Technical University, Prague, Czech Republic | DESY-Hamburg, Germany | IFAE, Barcelona, Spain
IFAE, Barcelona, Spain | Diamond Light Source, England, UK | JINR, Dubna, Russian Federation
Mid Sweden University, Sundsvall, Sweden | ESRF, Grenoble, France | NIKHEF, Amsterdam, The Netherlands
MRC-LMB Cambridge, England, UK | IEAP, Czech Technical University, Prague, Czech Republic | University of California, Berkeley, USA
NIKHEF, Amsterdam, The Netherlands | KIT/ANKA, Forschungszentrum Karlsruhe, Germany | University of Canterbury, Christchurch, New Zealand
University of California, Berkeley, USA | Mid Sweden University, Sundsvall, Sweden | University of Geneva, Switzerland
Universität Erlangen-Nurnberg, Erlangen, German | NIKHEF, Amsterdam, The Netherlands | University of Glasgow, Scotland, UK
University of Glasgow, Scotland, UK | Univesridad de los Andes, Bogota, Columbia | University of Houston, USA
University of Houston, USA | University of Bonn, Germany | University of Maastricht, The Netherlands
University and INFN Section of Cagliari, Italy | University of California, Berkeley, USA | University of Oxford, England, UK
University and INFN Section of Pisa, Italy | University of Canterbury, Christchurch, New Zealand | INFN, Italy
University and INFN Section of Napoli, Italy | Universität Erlangen-Nurnberg, Erlangen, German | Chinese Spallation Neutron Source, Dongguan City,China
 | University of Glasgow, Scotland, UK | Brazilian Light Source, Campinas, Brazil
 | University of Houston, USA | Philippine Nuclear Research Institute, Manila, Philippines
 | University of Leiden, The Netherlands | 
 | Technical University of Munich, Germany | 
 | VTT Information Technology, Espoo, Finland | 

## Slide 7: Acknowledgements – Commercial Partners
COLLABORATION NAME | Medipix2 |  |  | Medipix3 |  | Medipix4 | 
ASICS | Medipix2 | Timepix | Timepix2 | Medipix3 | Timepix3 | Medipix4 | Timepix4
ADVACAM s.r.o., Czech Republic | X | X | X | X | X |  | X
Amsterdam Scientific Instruments, The Netherlands | X | X | X | X | X |  | X
Kromek, UK | X | X | X |  | X |  | 
Malvern-Panalytical, The Netherlands | X | X | X | X |  |  | 
MARS Bio Imaging, New Zealand |  |  |  | X |  |  | 
PITEC, Brazil |  |  |  | X |  |  | X
Quantum Detectors, UK |  |  |  | X | X |  | X
Sydor Technologies, USA |  |  |  |  |  |  | X
X-ray Imaging Europe, Germany | X | X | X |  |  |  | 
X-spectrum, Germany |  |  |  | X |  |  | X
X = commercial licensee X = R and D licensee

## Slide 8: The Medipix and Timepix ASICs - Timeline
Year | 1999 | 2003 | 2006 | 2013 | 2014 | 2017 | 2018 | 2020 | 2023 | 2025?
ASIC | Alice1LHCb | Medipix2 | Timepix | Medipix3 | Timepix3 |  | Timepix2 | Timepix4 | Medipix4 | 
LHCb |  |  |  |  |  | VELOpix |  |  |  | VELOpix2
Medipix chips aim at energy sensitive photon counting and typically use frame-based readout
Timepix chips are more oriented towards single particle detection
This talk will focus on Timepix4 and our efforts towards large area tiling
1st large high resolution photon counting chip
Add possibility to measure time or energy
1st chip with multiple energy bins
Data driven readout (timestamp 1.6ns)
Tileable on 4 sides (timestamp 200ps)
Tileable on 4 sides (8 energy bins)

## Slide 9: Timepix3 Specs
CMOS node | 130nm
Pixel Array | 256 x 256
Pixel pitch | 55mm
Charge collection | e-, h+
Pixel functionality | TOT (Energy) and TOA (Arrival time)
Preamp Gain | ~47mV/ke-
ENC | ~60e-
FE Linearity | Up to 12ke-
TOT linearity (resolution) | Up to 200ke- (<5%)
TOA resolution* | Up to 1.6ns
Time-walk | <20ns
Minimum detectable charge | ~500e-  2 KeV (Si Sensor)
Max Analog power (1.5V) | 500 mA/chip
Digital Power (1.5V) | ~400mA data driven
Maximum hit rate | 80Mhits/sec (in data driven)
Readout | Data driven (44-bits/hit @ 5Gbps)
* Thanks to V. Gromov, et al. Nikhef, C. Brezina et al., Bonn

## Slide 10: Timepix3 Specs
CMOS node | 130nm
Pixel Array | 256 x 256
Pixel pitch | 55mm
Charge collection | e-, h+
Pixel functionality | TOT (Energy) and TOA (Arrival time)
Preamp Gain | ~47mV/ke-
ENC | ~60e-
FE Linearity | Up to 12ke-
TOT linearity (resolution) | Up to 200ke- (<5%)
TOA resolution* | Up to 1.6ns
Time-walk | <20ns
Minimum detectable charge | ~500e-  2 KeV (Si Sensor)
Max Analog power (1.5V) | 500 mA/chip
Digital Power (1.5V) | ~400mA data driven
Maximum hit rate | 80Mhits/sec (in data driven)
Readout | Data driven (44-bits/hit @ 5Gbps)
* Thanks to V. Gromov, et al. Nikhef, C. Brezina et al., Bonn

## Slide 11: Timepix3 Specs
CMOS node | 130nm
Pixel Array | 256 x 256
Pixel pitch | 55mm
Charge collection | e-, h+
Pixel functionality | TOT (Energy) and TOA (Arrival time)
Preamp Gain | ~47mV/ke-
ENC | ~60e-
FE Linearity | Up to 12ke-
TOT linearity (resolution) | Up to 200ke- (<5%)
TOA resolution* | Up to 1.6ns
Time-walk | <20ns
Minimum detectable charge | ~500e-  2 KeV (Si Sensor)
Max Analog power (1.5V) | 500 mA/chip
Digital Power (1.5V) | ~400mA data driven
Maximum hit rate | 80Mhits/sec (in data driven)
Readout | Data driven (44-bits/hit @ 5Gbps)
* Thanks to V. Gromov, et al. Nikhef, C. Brezina et al., Bonn

## Slide 12: Timepix3 Specs
CMOS node | 130nm
Pixel Array | 256 x 256
Pixel pitch | 55mm
Charge collection | e-, h+
Pixel functionality | TOT (Energy) and TOA (Arrival time)
Preamp Gain | ~47mV/ke-
ENC | ~60e-
FE Linearity | Up to 12ke-
TOT linearity (resolution) | Up to 200ke- (<5%)
TOA resolution* | Up to 1.6ns
Time-walk | <20ns
Minimum detectable charge | ~500e-  2 KeV (Si Sensor)
Max Analog power (1.5V) | 500 mA/chip
Digital Power (1.5V) | ~400mA data driven
Maximum hit rate | 80Mhits/sec (in data driven)
Readout | Data driven (44-bits/hit @ 5Gbps)
* Thanks to V. Gromov, et al. Nikhef, C. Brezina et al., Bonn

## Slide 13: Timepix3 Specs
CMOS node | 130nm
Pixel Array | 256 x 256
Pixel pitch | 55mm
Charge collection | e-, h+
Pixel functionality | TOT (Energy) and TOA (Arrival time)
Preamp Gain | ~47mV/ke-
ENC | ~60e-
FE Linearity | Up to 12ke-
TOT linearity (resolution) | Up to 200ke- (<5%)
TOA resolution* | Up to 1.6ns
Time-walk | <20ns
Minimum detectable charge | ~500e-  2 KeV (Si Sensor)
Max Analog power (1.5V) | 500 mA/chip
Digital Power (1.5V) | ~400mA data driven
Maximum hit rate | 80Mhits/sec (in data driven)
Readout | Data driven (44-bits/hit @ 5Gbps)
* Thanks to V. Gromov, et al. Nikhef, C. Brezina et al., Bonn

## Slide 14: Timepix3 Specs
CMOS node | 130nm
Pixel Array | 256 x 256
Pixel pitch | 55mm
Charge collection | e-, h+
Pixel functionality | TOT (Energy) and TOA (Arrival time)
Preamp Gain | ~47mV/ke-
ENC | ~60e-
FE Linearity | Up to 12ke-
TOT linearity (resolution) | Up to 200ke- (<5%)
TOA resolution* | Up to 1.6ns
Time-walk | <20ns
Minimum detectable charge | ~500e-  2 KeV (Si Sensor)
Max Analog power (1.5V) | 500 mA/chip
Digital Power (1.5V) | ~400mA data driven
Maximum hit rate | 80Mhits/sec (in data driven)
Readout | Data driven (44-bits/hit @ 5Gbps)
* Thanks to V. Gromov, et al. Nikhef, C. Brezina et al., Bonn

## Slide 15: Timepix3 miniaturised readout
Advacam s.r.o., Prague

## Slide 16: Timepix3 Demo

## Slide 17: Tracking in a single Si layer

## Slide 18: Test with 120GeV/c Pion Track
60 deg
p+ in n Si sensor
500mm thick
Vbias = 130V
Colour (and diameter) indicate charge
Measured z resolution ~50mm
Slide courtesy of B. Bergmann, S. Pospisil, IEAP, CTU, Prague

## Slide 19: Single Layer Compton Camera with MiniPIX TPX3
Compton camera principle
Typical two detectors
primary gamma is scattered in first detector (position and energy recorded), scattered gamma continues to second detector (absorbed, position and energy recorded)
from energies - > scattering angle calculated
from position and energies -> possible position of the source on the surface of a cone
Multiple cones intersection - > source position
Single Timepix3 layer camera
  Instead of 2 detectors, only single TPX3
  Using time of charge collection to determine relative depth
Courtesy of D. Turecek, Advacam s.r.o

## Slide 20
Reconstruction of position of three 131I gamma sources (364 keV)
131Iodine gamma source
3 different Iodine solution in small bottles positioned in a room at different positions
Distance from detector 3.5 m (activity 10’s of MBq)
Mapped on photograph of the room
Sources located correctly within minutes
Image took hours to collect
Single Layer Compton Camera with MiniPIX TPX3
Energy Spectrum of 131I
Courtesy of D. Turecek, Advacam s.r.o

## Slide 21: Gamma camera application: Thyroid diagnostics
Thyroid cancer diagnostics and treatment monitoring:
The second most frequent cancer for women (after breast cancer)
Current imaging methods offer resolution of about 12 mm in 2D
Our technology allows
  5 times better resolution and 3D (2.5 mm)
  4 times lower dose
Courtesy of D. Turecek, Advacam s.r.o

## Slide 22: Timepix3  Timepix4
 |  |  | Timepix3 (2013) | Timepix4 (2018/19)
Technology |  |  | 130nm – 8 metal | 65nm – 10 metal
Pixel Size |  |  | 55 x 55 µm | 55 x 55 µm
Pixel arrangement |  |  | 3-side buttable 256 x 256 | 4-side buttable 512 x 448
Sensitive area |  |  | 1.98 cm2 | 6.94 cm2
Readout Modes | Data driven (Tracking) | Mode | TOT and TOA | 
 |  | Event Packet | 48-bit | 64-bit
 |  | Max rate | <80 Mhits/s | <365 MHz/cm2/s
 |  | Max pix rate | 1.3kHz/pixel | 10.6kHz/pixel
 | Frame based (Imaging) | Mode | PC (10-bit) and iTOT (14-bit) | CRW: PC (8 or 16-bit)
 |  | Frame | Zero-suppressed (with pixel addr) | Full Frame (without pixel addr) CRW (8-bit / 16-bit) Up to 44 KHz frame @8b
 |  | Max count rate | 82 Ghits/cm2/s | ~800 Ghits/cm2/s
TOT energy resolution |  |  | < 2KeV | < 1Kev
Time resolution (bin size) |  |  | 1.56ns | ~200ps
Readout bandwidth |  |  | ≤5.12Gb (8 x SLVS@640 Mbps) | ≤163 Gbps (16 x 10.24 Gbps)
Target global minimum threshold |  |  | <500 e- | <500 e-

## Slide 23: Timepix4 Pixel Schematic
Global threshold
Front-end
Leakage
Current
compensation
Preamp
Input
pad
1 pixel
VCO
@640MHz
Super pixel
Control
voltage
8 pixels
640MHz
TpA
TpB
TestBit
MaskBit
Front-end
Counters
&
Latches
Time stamp
14-bits
Synchronizer
&
Clock gating
OP Mode
5-bit Local
Threshold
~50mV/ke-
2.5fF
3fF
Sync
Readout
Data out
to EOC
8-bits
clock
(40MHz)
8-bits
4
8-bits
ADB
32 pixels
8
Clock
local
 | bits
ADDR | 9
ToA | 16
ufToA_r | 4
ufToA_f | 4
fToA_r | 5
fToA_f | 5
ToT | 12
Pileup | 1
SPG

## Slide 24: Pixel Operation in TOA & TOT [DD]
Preamp Out
Disc Out
Clk (40MHz)
VCO Clk (640MHz)
TOT Clk (40MHz)
TOA (16-bit)
Global TOA (16-bit)
UFTOA_Start (4-bits)
UFTOA_Stop (4-bits)
65534
X
TOA (16 bits)=65534
Tpeak < 25ns
TOT (10 bits) =4
65535
0
2
3
4
65534
65533
1
FTOA_R (4 bits)=7
FTOA_F (4 bits)=11
0
X
8
X

## Slide 25: Full digital double column DLL [448 dDLL: 224 Top Matrix and 224 Bottom Matrix]
ADB
coarse
delay
section
fine
delay
section
14
15
in
out
output buffer
ctrl_coarse
ctrl_fine
Chart: Delay code
BCspectre | 0.27757 | 0.33289 | 0.3165 | 0.37181 | 0.73063 | 0.78595 | 0.77124 | 0.82655
TCspectre | 0.40701 | 0.47999 | 0.46629 | 0.53927 | 1.102 | 1.175 | 1.1647 | 1.2376
WCspectre | 0.68415 | 0.79703 | 0.78717 | 0.90004 | 1.9021 | 2.0149 | 2.0124 | 2.1253
iWoRID 2018
X. Llopart et al 2019 JINST 14 C01024
Timepix4 ~23 mW/cm2 @40MHz clock with a 100 psrms
Timepix3 ~100mW/cm2 @40MHz clock with ~1.2ns skew
Dynamic digital power consumption is distributed across the clock period
Measured in Timepix4v2

## Slide 26: Timepix4 floorplan
BOTTOM Matrix
256 x 448
55 µm2
TOP Matrix
256 x 448
55 µm2
WB Extender
WB Extender
TOP Periphery + TSVs
CENTER Periphery + TSVs
BOTTOM Periphery + TSVs
512 x 448 of 55 x 55 µm pixels
  2 Matrices (TOP and BOTTOM)
3 ‘peripheries’ with TSV (Through-Silicon-Vias):
  TOP, BOTTOM (TSV, WB): Data Readout (16x 10.24 Gbps Serializers)
  CENTER (TSV): Analog Blocks (DACs, ADC, Band-Gap…)
On-chip bump to pixel redistribution layer (RDL):
  Pixel matrix pixels are shorter (51.4 µm) than sensor pixels (55 µm)
  Equalized Cin for all pixels  ~46 fF increase for a 460 µm periphery
Edge peripheries include 1mm Wire Bond Extender
Dicing options:
  With WB (Wire-Bonds Extenders): 29.96 mm x 24.7 mm
    >93.7% active area (28.16mm x 24.64mm)
  Without WB (TSV Only) : 28.22 mm x 24.7 mm
    >99.5% active area (28.16mm x 24.64mm)
    Through Silicon Vias (TSV) requires post processing at wafer level to create TSV and on the ASIC back sides RDL + BGA pads
Periphery
256 x 256
55 µm2
(65 Kpixels)
Timepix3
24.64 mm
28.16 mm
Timepix4

## Slide 27: Timepix4 Bottom left detail
BOTTOM Periphery
BOTTOM Matrix
BOTTOM
WB Extender
Dicing lane #1
Dicing lane #2
TSVs landing pads in M1
460 µm

## Slide 28: Analog (static) power supply distribution
 | Total I (chip) |  | 2 WB | 3 TSV
Nominal Analog Power [10 µA/pixel] | ~2300 mA | Vdrop [VDDA-GNDA] | 19.6 mV | 6.9 mV
 |  | Imax pad | 60 mA | 57 mA
Low Analog Power [1 µA/pixel] | ~230 mA | Vdrop [VDDA-GNDA] | 1.96mV | 0.69mV
 |  | Imax pad | 6 mA | 5.7 mA
2 wire bonds
3 TSV

## Slide 29: Timepix4 submissions
Timepix4v2
4 FEOL + 4 BEOL masks changed
19 wafers received
Q2 2021
TDC and High speed links working as expected
Further improvement in RDL shielding in peripheries
Chip at its final version
Timepix4v0
Full mask engineering run
6 wafers received
Q4 2019
Excess noise coupling from peripheries to FE
640 MHz clock in edge peripheries
VCO not oscillating at nominal frequency
Chip is operational
Timepix4v1
4 BEOL masks changed
Small test VCO chip
6 wafers received
Q3 2020
VCO not oscillating at nominal frequency
Improved RDL shielding in peripheries
640MHz in peripheries recovered
Timepix4v3
2 BEOL masks changed
25 wafers received
Q3 2022
Larger IO Pads
Chip at its final version

## Slide 30: Uniformity of response Timepix4 (all versions)
(a)
Threshold adjustment bits

## Slide 31: Gain slopes for different FE Gain [TOA-TOT, few pixels]
Chart: [mV]
| 1 | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17
LG | 19.3135 | 61.632 | 103.041 | 145.7073 | 187.4373 | 231.2538 | 271.4857 | 301.1783 | 319.5823
LG-1sig | 17.7141 | 60.5397 | 101.2845 | 143.749 | 185.9718 | 229.2379 | 268.452 | 293.06 | 310.5059
LG+1sig | 20.9129 | 62.7243 | 104.7975 | 147.6655 | 188.9027 | 233.2696 | 274.5195 | 309.2965 | 328.6586
HG | 33.6248 | 105.181 | 177.299 | 248.24 | 300.242 | 322.7655 | 343.0955 | 362.0077 | 379.636 | 379.636
HG-1sig | 32.2772 | 103.7246 | 175.5425 | 245.7853 | 290.6817 | 311.7572 | 332.0749 | 350.5474 | 368.1112 | 367.8745
HG+1sig | 34.9723 | 106.6374 | 179.0555 | 250.6947 | 309.8023 | 333.7738 | 354.1161 | 373.4681 | 391.1608 | 391.3975
Low gain mode
~20.5 mV/ke-
High gain mode
~35.8 mV/ke-
Chart: [mV]
| 1 | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17
LG | 18.511 | 61.097 | 103.255 | 144.771 | 186.822 | 227.696 | 266.323 | 301.098 | 330.63
LG-1sig | 16.2719 | 59.121 | 101.1721 | 142.6423 | 184.6848 | 225.6987 | 264.1718 | 298.3667 | 325.9789
LG+1sig | 20.7501 | 63.073 | 105.3379 | 146.8997 | 188.9592 | 229.6933 | 268.4742 | 303.8293 | 335.2811
HG | 29.318 | 99.403 | 170.023 | 235.935 | 292.966 | 333.305 | 360.055 | 380.171 | 396.435
HG-1sig | 27.2819 | 97.6363 | 168.3984 | 234.2521 | 288.9534 | 325.9612 | 350.5637 | 369.3195 | 385.1993
HG+1sig | 31.3541 | 101.1697 | 171.6476 | 237.6179 | 296.9786 | 340.6488 | 369.5463 | 391.0225 | 407.6707
LOG | 14.659 | 55.319 | 97.156 | 137.816 | 171.842 | 199.02 | 219.457 | 235.935 | 249.31
LOG-1sig | 12.6617 | 52.7893 | 94.3443 | 134.6801 | 166.7826 | 191.2885 | 210.4987 | 225.9246 | 238.4757
LOG+1sig | 16.6563 | 57.8487 | 99.9677 | 140.9519 | 176.9014 | 206.7515 | 228.4153 | 245.9454 | 260.1443
Low/LOG gain mode
~20.5 mV/ke-
High gain mode
~34.5 mV/ke-
Hole collection
Electron collection
kh+
ke-

## Slide 32: Noise uniformity - Timepix4v0, v1 and v2
Chart: [e-rms]
Chart: Row #
When a Si sensor is attached the noise slightly increase ~3e- / pixel
v2

## Slide 33: Timepix4v1 and v2: Number of Noisy pixels (>1 count) Data-Driven mode [DD]
Chart: Thr[e-]
| 2000 | 1950 | 1900 | 1850 | 1800 | 1750 | 1700 | 1650 | 1600 | 1550 | 1500 | 1450 | 1400 | 1350 | 1300 | 1250 | 1200 | 1150 | 1100 | 1050 | 1000 | 950 | 900 | 850 | 800 | 750 | 700 | 650 | 600 | 550
DD-TPIX4v1 (TOA ON) 0.1s | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 10 | 11 | 11 | 14 | 17 | 22 | 26 | 35 | 43 | 54 | 66 | 88 | 102 | 129 | 185 | 237 | 348 | 478 | 820 | 1250 | 2996 | 27343
DD-TPIX4v1 (TOA OFF) 0.1s | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 9 | 10 | 11 | 12 | 16 | 17 | 22 | 25 | 28 | 34 | 39 | 44 | 49 | 56 | 64 | 82 | 98 | 134 | 174 | 290 | 536 | 1939
DD-TPIX4v2 (TOA ON) 0.1s | 1 | 1 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 6 | 6 | 9 | 9 | 14 | 18 | 25 | 28 | 34
DD-TPIX4v2 (TOA OFF) 0.1s | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 3 | 5 | 9 | 11 | 13 | 15 | 23 | 31 | 39
Less than 40 noisy pixels @ 550 e-

## Slide 34: Timepix4v2 2D VCO frequency distribution
Measured VCO oscillation frequency in Timepix4
Calibration is required in order to get to the designed time resolution (~60psrms)
How could this be improved/simplified?
Chart: VCO frequency [MHz]
Mean=639.5 MHz
σ=6.6 MHzrms

## Slide 35: TOA Resolution [TOA-TOT, 1 pixel, 10000 samples, HG e-]
Chart: samples
Chart: samples
Chart: samples
Digital TestPulse
TPulse=2 ke-
TPulse=15 ke-
Chart: [ke-]
| 0.99 | 1.49 | 1.99 | 2.4900000000000002 | 2.99 | 3.48 | 4 | 4.5 | 5 | 5.5 | 5.99 | 6.49 | 6.99 | 7.49 | 7.99 | 8.49 | 8.99 | 9.48 | 10 | 10.5 | 11 | 11.5 | 11.99 | 12.49 | 12.99 | 13.49 | 13.99 | 14.49 | 14.99 | 15.49
Analog TestPulse | 930.577 | 645.847 | 455.7614 | 355.5959 | 292.8421 | 264.5544 | 209.902 | 194.3859 | 185.6355 | 169.126 | 168.6735 | 149.4169 | 118.3641 | 108.0262 | 117.9701 | 123.8986 | 121.526 | 120.4151 | 112.6594 | 103.1192 | 94.6877 | 98.2889 | 103.4312 | 103.8366 | 107.9257 | 109.477 | 102.0421 | 92.2383 | 82.8884 | 83.3341
Digital TestPulse | 70.3897 | 30.4062 | 75.6846 | 40.2742 | 39.9939 | 68.0964 | 43.7006 | 26.1088 | 41.9034 | 36.4075 | 22.5515 | 68.9907 | 75.3258 | 59.4983 | 74.4332 | 68.3992 | 27.0307 | 35.7647 | 64.1901 | 58.2965 | 60.366 | 74.8893 | 26.1088 | 27.9184 | 66.2877 | 66.8364 | 67.9944 | 54.3493 | 34.434 | 75.507
TOA resolution
Timepix3 TOA resolution limit

## Slide 36: Timepix4 assembly (300mm Si sensor)

## Slide 37: Data driven mode
10s exp. 90Sr
Threshold ~ 800e-
6.1 M packets @ 5 Gbps

## Slide 38: Photon counting image Timepix4

## Slide 39: Energy calibration using test pulses
Per-pixel energy vs ToT calibration using test pulses
137Cs and 241 Am
All pixel clusters
137Cs and 241 Am
All pixel clusters
R. Bolzonella et al., TIPP Cape Town, 4-8 Sept
Under submission to JINST
After test pulse
calibration

## Slide 40: Timing test setup with laser
σ = 1.4 pixel = 77 μm
most of the charge is released in 4-5 pixels
Wide range of released charge covered
1020nm laser pulsed with 6ps jitter generator
Generator connected to electrical test pixels
Rear side metallization with holes
Timepix4V2 bonded to a 100μm thick n-on-p Si detector:
Biased at -150 V
R. Bolzonella et al., TIPP Cape Town, 4-8 Sept
Under submission to JINST

## Slide 41: Single pixel pulse resolution
For the pixel [305,144], where the laser is focused, the standard deviation saturates at 129±1 ps rms
Subtracting the contribution of the reference TDC (60 ps), a resolution of 111±1 ps rms is obtained
Timing resolution dependency on cluster charge best result: σToADiffAvg=79 ± 1 ps rms
Timing resolution subtracting reference TDC contribution: σToAAvg=49 ± 1 ps rms
R. Bolzonella et al., TIPP Cape Town, 4-8 Sept
Under submission to JINST

## Slide 42: June 2023 – TSV processes TPIX4v0 delivered

## Slide 43: Rear side Re-Distribution Layer (RDL)

## Slide 44: Visual Inspection – front and rear side
Redistribution Layer
1020 BGA pads
TSV
Residues of dicing tape
Front Side
Back Side
Chip bow ~ 100mm (over 3cm)
Magnified bowing (the chip is thin: ~120mm)
Pixel side
/14

## Slide 45: Inspection with Electron Microscope of one TSV
Courtesy of IZM
Landing pad of the TSV
Bottom of the TSV
Pixel pads
/14

## Slide 46: Mounting of TSV processed chips on Nikhef carrier board
ACF: Anisotropic Conducting Film
Conductive particles
Pressure needed
Temporary bonding
3M
We could communicate with the chip, test the DAC…
After releasing the clamp and putting it again, chip not responding, maybe due to reusing the tape.
Work in progress
ACP: Anisotropic Conducting Paste
Araldite with conducting particles
Done at UniGe (Mateus Vicente Barreto Pinto m.vicente@cern.ch)
Flip chip bond machine
Results Timepix4V0:
Good communication with the chip
Promising low cost approach

## Slide 47: TSV-processed successfully mounted
Chip behaves identically to a wire bonded version

## Slide 48: Output from TSV processed Timepix4v0

## Slide 49: Medipix3  Medipix4
 | Medipix3RX (2013) | Medipix4 (2022)
Technology | CMOS 130nm | CMOS 130nm
Pixel Size | 55/110 µm | 75/150 µm
Matrix Size | 256 x 256 / 128 x 128 | 320 x 320 / 160 x 160
Tile-ability | 3-side | 4-side buttable
Thresholds | 1/4 Continuous RW | 2/8 Continuous RW
Readout scheme | Sequential RW & Continuous RW | 
Count Rate (10% deadtime loss) | 4.3 x 106 ph/mm2/s (CSM 110) | 19 x 106 ph/mm2/s (CSM 140)
Dynamic Range | 25 Ke- | 32 Ke-
CSM energy resolution (FWHM) | ~ 4.4 KeV (CSM 110) | ~ 2.5KeV (CSM 150)
Readout bandwidth | <1.6 Gbps (8x LVDS) | 1,2 Serializers @5.12 or 2.56 Gbps
Power | <1W @1.5V | <1W/cm2 @1.2V
Status: Medipix4 chip first version under test – bugs being addressed

## Slide 50: Conclusions
Hybrid pixel detectors offer unique solutions for X-ray and gamma-ray imaging and have found widespread applications
The Timepix4 (and Medipix4) chips are designed explicitly for tiling in 2 dimensions using Through Silicon Vias for IO
Measurements with Timepix4 in data driven mode are consistent with the design values
Single pixel time resolution is ~111ps rms
TSVs have been successfully implemented on Timepix4v0 and the chip behaves as normal
Future work includes demonstrating a large area implementation with minimal dead area

## Slide 51
Thank you for your attention!
Medipix3RX images: S. Procz et al.