## Slide 1
A Spectroscopic Imaging Camera with a 1 MHz Frame Rate Continuous Readout
Ben Cline et al.
Detector Development Group, UKRI STFC
Email: ben.cline@stfc.ac.uk
07/09/2023

## Slide 2
ASIC Design
Stephen Bell
Thomas Gardiner
Lawrence Jones
Mark Prydderch
Detector Development
Ivan Church
Ben Cline
Matt Hart
Sion Richards
Paul Seller
Dave Sole
Matt Wilson
Rhian Wheater
Matt Veale
Interconnect
Paul Adkin
Paul Booker
Toby Brookes
Navid Ghorbanian
John Lipp
Andreas Schneider
PCB Design Office
Darren Ballard
Dan Becket
Chris Day
Electronic Systems Design
Adam Barcock
Rob Halsall
John Holden
Chris Lyford
Sooraj Pradeep
Matt Roberts
Detector Systems Software
Dominic Banks
Victor Bozhinov
Adam Davies
Josh Harris
Ashley Neaves
Tim Nicholls
Joseph Nobes
Funded by the: Centre for Instrumentation (CFI) run by Marcus French
Detector Systems
William Helsby

## Slide 3
1 HEXITEC vs HEXITECMHz
The next generation of HEXITEC systems
2 HEXITECMHz Overview
An introduction to our new HEXITECMHz ASIC including its architecture and specification
3 Test Results
Characterisation of spectroscopic performance using a lab-based X-Ray source and Diamond Light Source synchrotron X-Rays
4 Next Steps
A look to the next 12 months
Agenda

## Slide 4
3 Test Results
Characterisation of spectroscopic performance using a lab-based X-Ray source and Diamond Light Source synchrotron X-Rays
1 HEXITEC vs HEXITECMHz
The next generation of HEXITEC systems
2 HEXITECMHz Overview
An introduction to our new HEXITECMHz ASIC including its architecture and specification
4 Next Steps
A look to the next 12 months
Agenda

## Slide 5
Multiple frames
Per-pixel/ global histograms
Parameter | HEXITEC
Pixel Pitch (µm) | 250
Array Size | 80 × 80
Max Frame Rate (kHz) | ~10
Max Spectroscopic Flux (ph s-1 mm-2) | ~104
Digitisation | Off-chip
Detector Type | Peak Track + Hold
Gain stages (keV in CZT) | 200
 | 600
FWHM@60keV (keV) | HF-CZT: 0.79 keV [1] p-type Si: 0.59keV [2]
HEXITEC specifications
Fully spectroscopic X-Ray Imaging with CdTe/ CdZnTe (CZT) detectors
80 × 80 pixel array on a 250 μm pitch
Analogue readout via a rolling shutter to 4 ADCs in the DAQ
Maximum ~10 kHz data output (~100 MB s-1)
AMS 0.35 µm - since 2010
[1] DOI: 10.3390/s20102747 [2] DOI: 10.1088/1748-0221/17/05/P05030

## Slide 6
Examples from
HEXITEC has been used across many applications including:
Compton X-Ray Imaging [3]
X-Ray Fluorescence Imaging [4]
Hyperspectral X-Ray Tomography [5]
[3] DOI: 10.1016/j.mtener.2022.101224 [4] DOI: 10.1557/mrs.2020.270 [5] DOI: 10.1038/s41598-022-23592-0
Figure 7. (a) Pt and Er XRF intensity maps as a function of temperature during the solidification of an Al-23Pt-20Er alloy at 0.1 Ks–1.
Figure 1. Schematic of the correlative X-Ray Compton scattering imaging (XCS-I) and X-Ray computed tomography (XCT) technique experimental set-up
Figure 2. Voxel spectra analysis for double-stained hindlimb specimen. (A) Single image slice in the sagittal plane across three monochromatic energy channels, following iterative reconstruction. A set of three regions-of-interest (ROIs) are highlighted for voxel spectra analysis.

## Slide 7
BUT …
HEXITEC is frame-rate and therefore flux limited
  Charge-sharing corrections require <10 % frame occupancy
  This limits spectroscopic imaging @10 kHz to a max X-Ray flux of ~104 ph s-1 mm-2
→ There is a need for a continuous spectroscopic X-Ray detector that operates at faster frame rates
Faster (MHz) imaging requires:
  In-pixel digitisation
  High-speed serialisers
  Dedicated FPGA processing
These requirements apply to a large number of STFC’s upcoming detector development projects
  E.g. DynamiX, C100

## Slide 8
Parameter | HEXITEC | HEXITECMHz
Pixel Pitch (µm) | 250 | 250
Array Size | 80×80 | 80×80
Max Frame Rate (kHz) | ~10 | 1000
Max Spectroscopic Flux (ph s-1 mm-2) | ~104 | >106
Digitisation | Off-chip | On-chip
Detector Type | Track + Hold | Integrating
Gain Stages (keV in CZT) | 200 | 100
 | 600 | 200
 |  | 300
FWHM@100keV (keV in CZT) | < 1 | < 1
Power Consumption (W) | 1.5 | 15
Comparison of HEXITEC and HEXITECMHz specifications
A new fully spectroscopic X-Ray imaging detector:
  80×80 pixels on a 250 μm pitch
  New front-end design – now an integrating detector
  On-chip 12-bit digitisation (no external ADCs)
    Data outputted via 20 × 4.1 Gbps serialisers
  1 MHz continuous frame rate
    Spectroscopic X-Ray fluxes of >106 ph s-1 mm-2
Photo of ASIC

## Slide 9
(Left) ASIC block diagram, (Right) HMHz ASIC

## Slide 10
3 Test Results
Characterisation of spectroscopic performance using a lab-based X-Ray source and Diamond Light Source synchrotron X-Rays
1 HEXITEC vs HEXITECMHz
The next generation of HEXITEC systems
2 HEXITECMHz Overview
An introduction to our new HEXITECMHz ASIC including its architecture and specification
4 Next Steps
A look to the next 12 months
Agenda

## Slide 11
- Integrating Front End
  Properties:
  CSA + CDS amplifier
  3 static dynamic ranges
  t10-90 = 50 ns
  treset ≈ 90 ns
  ENC < 70 e-
Front end schematic

## Slide 12
- In-Pixel TDC
Super-pixel schematics
TDC signal timing
12-bit digitisation
  The Sample & Hold signal is compared with a ramp with a programmable slew
The ASIC comprises 800 super-pixels of 2 × 4 pixels
  These contain one TDC block and shared readout logic

## Slide 13
- Data Output
80×80 array is divided into divisions of 4 columns
  Each division has packet-assembler and dedicated serialiser
  Packets constructed using Xilinx’s Aurora 64B/66B protocol
  Serialisers operate at 4.1 Gb s-1 (total data rate ≈ 10 GBs-1)
What to do with all this data?
Two receiving data planes:
  First-stage FPGA – Alpha Data card
    Recovers and reorders the received optical data
  Second-stage FPGA – Xilinx Alveo U50 board
    Data corrections (darks, energy calibration)
    Data reduction (charge-sharing discrimination, histogramming)
Then received by ODIN Data, a scalable data-processing and acquisition network
100GbE
100GbE
Processed
Data
4.1Gb Aurora × 20 lanes
Alveo board
Alpha Data card

## Slide 14
- Data Output
80×80 array is divided into divisions of 4 columns
  Each division has packet-assembler and dedicated serialiser
  Packets constructed using Xilinx’s Aurora 64B/66B protocol
  Serialisers operate at 4.1 Gb s-1 (total data rate ≈ 10 GBs-1)
What to do with all this data?
Two receiving data planes:
  First-stage FPGA – Alpha Data card
    Recovers and reorders the received optical data
  Second-stage FPGA – Xilinx Alveo U50 board
    Data corrections (darks, energy calibration)
    Data reduction (charge-sharing discrimination, histogramming)
Then received by ODIN Data, a scalable data-processing and acquisition network
100GbE
100GbE
Processed
Data
4.1Gb Aurora × 20 lanes
Alveo board
Alpha Data card
DOI: 10.1088/1748-0221/17/10/C10012

## Slide 15
3 Test Results
Characterisation of spectroscopic performance using a lab-based X-Ray source and Diamond Light Source synchrotron X-Rays
1 HEXITEC vs HEXITECMHz
The next generation of HEXITEC systems
2 HEXITECMHz Overview
An introduction to our new HEXITECMHz ASIC including its architecture and specification
4 Next Steps
A look to the next 12 months
Agenda

## Slide 16
ASIC is fully functioning
Using test enclosure
FPGA firmware in development – at present, 1 channel (4 × 80 pixels) over fast data
Current Status
HEXITECMHz test enclosure
50 cm
30 cm
15 μs
1 μs
Fast data output

## Slide 17
Testing – Initial X-Ray Tests
Ta Kα and Kβ XRF
Cd & Te XRF + Escape
Ta L lines
X-Ray transmission measurements of a DDR Ram Card using a HEXITECMHz HF-CZT sensor
X-Ray Fluorescence (XRF) measurements of Ta foil using a HEXITECMHz HF-CZT sensor in medium-gain mode

## Slide 18
B16 Beamline: August and December 2022
  Monochromatic X-Rays: 10 – 20 keV
  Photon fluxes: 105 – 108 ph s-1 mm-2
  1 MHz data stream on one fast-data channel
  Tested HF-CZT (2 mm), p-type Si (300 µm), GaAs (500 µm) devices
Beamline scientists: Vishal Dhamgaye, Oliver Fox, Kawal Sawhney
Experimentation at
B16 setup photos

## Slide 19
B16 Beamline: August and December 2022
  Monochromatic X-Rays: 10 – 20 keV
  Photon fluxes: 105 – 108 ph s-1 mm-2
  1 MHz data stream on one fast-data channel
  Tested HF-CZT (2 mm), p-type Si (300 µm), GaAs (500 µm) devices
Beamline scientists: Vishal Dhamgaye, Oliver Fox, Kawal Sawhney
Experimentation at
B16 setup photos
DOI: 10.1088/1748-0221/18/07/P07048

## Slide 20
–Characterisation
Occupancy across fast-data channel
HF-CZT; 20 keV; High Gain; 7 mm Al
Calibrated CSD histogram of Pixel (37,1) @ 21 keV – 0.2 keV bin width. Gaussian FWHM fit shown

## Slide 21
– Characterisation
Occupancy across fast-data channel
Single photon FWHM in beam area
Calibrated CSD histograms in beam area

## Slide 22
– Characterisation
p-type Si; 15 keV; High Gain; 3 mm Al
Occupancy across fast-data channel
Single photon FWHM in beam area
Calibrated CSD histograms in beam area

## Slide 23
– Characterisation Results
Peak Energy [keV] | HG FWHM [keV] | MG FWHM [keV]
10 | 0.66 ± 0.06 | 0.77 ± 0.06
15 | 0.68 ± 0.06 | 0.81 ± 0.06
Peak Energy [keV] | HG FWHM [keV] | MG FWHM [keV] | LG FWHM [keV]
20 | 0.85 ± 0.10 | 0.92 ± 0.11 | 1.13 ± 0.12
HF-CZT
p-type Si
p-type Si FWHM in beam area
Electron-hole pair generation energies
CZT – 4.6 eV (e-h pair)-1
Si – 3.6 eV (e-h pair)-1

## Slide 24
– Integrating Measurements
P-type Si; 15 keV; High Gain; 3 mm Al – 0 mm Al
Calibrated histograms of pixel (48,2)
Occupancy across fast-data channel
0.4 × 106 ph s-1 mm-2

## Slide 25
– Integrating Measurements
P-type Si; 15 keV; High Gain; 3 mm Al – 0 mm Al
Calibrated histograms of pixel (48,2)
Occupancy across fast-data channel
1.0 × 106 ph s-1 mm-2

## Slide 26
– Integrating Measurements
P-type Si; 15 keV; High Gain; 3 mm Al – 0 mm Al
Calibrated histograms of pixel (48,2)
Occupancy across fast-data channel
3.0 × 106 ph s-1 mm-2

## Slide 27
– Integrating Measurements
P-type Si; 15 keV; High Gain; 3 mm Al – 0 mm Al
Calibrated histograms of pixel (48,2)
Occupancy across fast-data channel
9.0 × 106 ph s-1 mm-2

## Slide 28
– Integrating Measurements
P-type Si; 15 keV; High Gain; 3 mm Al – 0 mm Al
Calibrated histograms of pixel (48,2)
Occupancy across fast-data channel
20 × 106 ph s-1 mm-2

## Slide 29
HF-CZT: Changing attenuators over 20 seconds (30 ms/video frame)
– Dynamic Measurements
Dynamic measurement schematic

## Slide 30
HF-CZT: Changing attenuators over 20 seconds (30 ms/video frame)
– Dynamic Measurements
Dynamic measurement schematic

## Slide 31
HF-CZT: Changing attenuators over 20 seconds (30 ms/video frame)
– Dynamic Measurements
Real time Xilinx U50 FPGA processing will enable live corrections of this effect
Dark Offset varies with incident flux
Dark Offset varies with time at a given flux
Dynamic measurement schematic
More details at HiZPAD Workshop
20th – 21st September, DESY
https://indico.esrf.fr/event/18/

## Slide 32
3 Test Results
Characterisation of spectroscopic performance using a lab-based X-Ray source and Diamond Light Source synchrotron X-Rays
1 HEXITEC vs HEXITECMHz
The next generation of HEXITEC systems
2 HEXITECMHz Overview
An introduction to our new HEXITECMHz ASIC including its architecture and specification
4 Next Steps
A look to the next 12 months
Agenda

## Slide 33
Next Steps
Achieving a 20-channel fast data output
  Reading out the full 80 × 80 array using the fast data
Implementation of in-FPGA histogramming
Delivery of HEXITECMHz detector system
  Smaller, form factor production-grade system
Further lab-based ASIC and sensor characterisation
Matt Hart
Delivering Science
  5D-CT for Materials Science – University of Manchester
  Mammography – Institute of Cancer Research, UK
  Security imaging – Nottingham-Trent University
Possible designs for HEXITECMHz detector system

## Slide 34
Summary
If you have any further questions, please contact me at: ben.cline@stfc.ac.uk
Parameter | HEXITECMHz
Max Frame Rate (MHz) | 1
Max Spectroscopic Flux (ph s-1 mm-2) | >106
Digitisation | On-chip
Detector Type | Integrating
Measured FWHM (High Gain) | 0.68 keV @ 15 keV in p-type Si 0.85 keV @ 20 keV in HF-CZT
HEXITECMHz is a fully-spectroscopic X-Ray detector capable of operating continuously at 1 MHz
  Testing underway to characterise the ASIC and optimise register settings
  The next year will include delivery of a full 80 × 80 pixel readout and a new smaller form-factor system
  CdZnTe offsets to be presented at HiZPAD

## Slide 35
Backup Slides

## Slide 36
Testing – Test Setup
HEXITECMHz test enclosure (exterior)
HEXITECMHz test enclosure (interior)
50 cm
30 cm

## Slide 37
Available Frame Rates
Two chip registers can be altered:
  Frame Time: 200→255 clocks frame-1
  Integration Length: 1 →255 frames
Available Frame Rates
Default values

## Slide 38
Test Pulse Sequence
Test pulse sequence
Fast Data output using test pulse sequence
15 μs
1 μs

## Slide 39
Potential application - Results from Veale et al., 2019;
  Data collected with STFC’s PIXIE ASIC (DOI: 10.1016/j.nima.2019.01.045)
Negative offset circuit
Front End – Negative Offset Circuit

## Slide 40
Data Output Plane
1st Level Receiver FGPA
ASIC
Image © STFC John Dawson
20x 4.1G Optical Aurora (10m+)
2nd Level Processing FPGA
2x 100G Optical Network Protocol
ODIN Data
Zynq SoC
ODIN Control
Control
Data
Data output path schematic

## Slide 41
Synchronising the ASIC
Synchronisation circuit
ASIC synchronised by controlling transition of SYNC input
  Logic 0 → logic 1
  Synchronisation occurs on 3rd subsequent rising clock edge
Used for external synchronisation and multi-ASIC systems

## Slide 42
Electronic Noise Contribution
HF-CZT FWHM @ 20 keV ≈ 0.9 keV
HF-CZT w factor = 4.67 eV
→ ENC =

## Slide 43
– Charge Sharing
The distribution of energies in two-pixel charge sharing events using a 10 keV beam energy at an occupancy of ~2% - 0.2 keV bin width
P-type Si; 10 keV; High Gain; ~2 % occupancy; 0.2 keV bin width

## Slide 44
– Temporal Stability
Evolution of high-gain calibrated global CSD spectra during exposure @ 10 keV for ~8 h – 0.2 keV bin width. Inset plot shows uncalibrated CSD spectra for pixel (14,45) – 1 ADU bin width
Change in performance of p-type HEXITECMHz device during exposure @ 10 keV for ~8 h. Plots show fluctuation in peak position (a), change in FWHM (b), and fractional change in peak counts (c) of 10 keV photo peak. Peak counts comprise events within the FWHM of peak, and error bars given by standard deviation across beam region

## Slide 45
TSVs on HEXITEC
https://doi.org/10.1016/j.nima.2021.166083
Wafer maps showing results of probe testing. Green – functioning; Red – non-functioning
(a) I/O pads on reverse of the 200 mm HEXITEC4S wafer post-TSV-last processing. (b) RDL pads dimensions on ASIC’s reverse. (c) SEM micrograph of exposed TSVs in cleaved test chip. (d) SEM micrograph showing contact between TSV and ASIC’s bottom metal layer

## Slide 46
MIC-12-434: On the feasibility of using HEXITECMHZ and fully-spectral x-ray imaging to detect breast tumours: an in-silico study
New Project: 5DCT – Dynamic Colour X-ray Computed Tomography
Next steps – Current Applications

## Slide 47
Testing – SPI vs Fast Data Comparison
Fast data matches SPI output 
Comparison of fast data and SPI measurements using test pulse

## Slide 48
Offset effect – Incrementing beam size within a pixel

## Slide 49
Offset effect – Impeding CSA reset
0 Reset – 0R | 1 1R | 2 2R | 3 3R | 4 4R | 5 5R | 6 6R | 7 7R | 8 8R | 9 9R
10 Reset – 0R | 11 1R | 12 2R | 13 3R | 14 4R | 15 5R | 16 6R | 17 7R | 18 8R | 19 9R
20 Reset – 0R | 21 1R | 22 2R | 23 3R | 24 4R | 25 5R | 26 6R | 27 7R | 28 8R | 29 9R
30 Reset – 0R | 31 1R | 32 2R | 33 3R | 34 4R | 35 5R | 36 6R | 37 7R | 38 8R | 39 9R
Group + histogram
all 1R frames

## Slide 50
Finding the centroids of each of the noise peak distributions
Plotting noise peak centroid against frames following reset
Now showing all pixels in channel
Illuminated Area
Offset effect – Impeding CSA reset

## Slide 51
Pixel outside of beam region
6.85 ADU frame-1 gradient of fit
Using a gradient of ~ 0.036 keV ADU
  0.036 × 6.85 = 0.2466 keV frame-1
CZT w-factor = 4.6 eV per e-h pair
  0.2466/(4.6/1000) = 53 e-
  53 e- = 8.48 × 10-18 C frame-1
1 frame = 1 µs
  8.48 x 10-12 A ≈ 8.5 pA
This is a sensible value for the leakage current
Pixel within beam region
23.87 ADU frame-1 gradient of fit
Using a gradient of ~ 0.036 keV ADU
  0.036 × 23.87 = 0.8593 keV frame-1
CZT w-factor = 4.6 eV per e-h pair
  0.8593/(4.6/1000) = 195 e-
  195 e- = 3.12 × 10-17 C frame-1
1 frame = 1 µs
  3.12 x 10-11 A ≈ 31.2 pA
This is ~24 pA higher than outside of the beam
Offset effect – Impeding CSA reset