## Slide 1: Recent results from the MIMOSIS-1 CMOS MAPS
M. Deveaux on behalf of the
IPHC-IKF-GSI (CBM-MVD) Collaboration
(05P19RFFC1)
Supported by

## Slide 2: The CBM Micro Vertex Detector (MVD)
CBM
CBM-MVD
13th POS conference, Oxford, Sept 2023
M. Deveaux
Darmstadt, Germany

## Slide 3: The (initial) mission: Reconstruct short - lived particles
Compute mother particle properties by:
Adding energy + momentum of daughters (invariant mass).
Challenge:
Select good daughter particles by separating primary and secondary vertex.
Need 5µm spatial resolution,
Need 300 - 500 µm Si equivalent material per station (0.3% X0).
… like ALICE ITS-2…
CBM 8 AGeV: ~10% of all particles shown
We are here
High rate capability needed:
5µs time resolution.
20 MHz/cm² (peak 80 MHz/cm²).
~1014 neq/cm² radiation tolerance.
Tolerance to heavy ion hits (direct beam impacts).
(Find full list of requirements in the backup)
Conclusion (2003, still true today):
Need specifically designed CMOS Monolithic Active Pixel Sensors.
13th POS conference, Oxford, Sept 2023
~10 x ALPIDE

## Slide 4: MIMOSIS R&D plan
MIMOSIS-0 (2018)
Demonstrate pixel concept.
Demonstrate zero suppression.
Demonstrate readout concept.
MIMOSIS-1 (2020)
Full dimension sensor
Add buffer structure.
SEE hardening 1/2
MIMOSIS-2 (Q2/2023)
On-chip pixel grouping.
Final pixels.
SEE hardening 2/2
All submissions:
Additional CE18 test structures to study specific design questions.
MIMOSIS-1, 60µm thick
Discussed today
MIMOSIS-1
MIMOSIS-3
Final sensor for mass production
In synergy with:
13th POS conference, Oxford, Sept 2023

## Slide 5: MIMOSIS-1
Size: 504 x 1024 pixels (27x30µm²)
FEE integrated.
Pixels types: 2x DC, 2x AC (>20V depletion voltage)
MIMOSIS-1, 60µm thick
p-stop
n-gap
standard
Standard
Rad. hard, candidate 1
Rad. hard, candidate 2
First proposed by W. Snoeys et al.
H. Pernegger et al., 2017 JINST 12 P06008
A
B
C
D
AC-pixel
>20V HV top bias
Capacitor blocks HV
Back bias possible.
P-Well bias (ALPIDE) possible
Most likely Fully Depleted
13th POS conference, Oxford, Sept 2023

## Slide 6
MIMOSIS-1 beam test setup
MIMOSIS-1 beam telescope:
6 MIMOSIS-1 sensors
4 reference sensors (standard epi-layer)
2 device under test
Reference Track Uncertainty:
  2.5 µm + 1.5 µm Mult. Scattering (DESY only)
Irradiation:
1 MeV reactor neutrons (TRIGA, Ljubjana).
Few 10 keV X-rays (KIT)
Storage at room temperature.
Beam test:
5 GeV e- Beam @ DESY
120 GeV Pion Beam @SPS-CERN
~ 1 GeV d – beam @ COSY
Stabilized room temperature
preliminary
13th POS conference, Oxford, Sept 2023

## Slide 7
Sensor performance: New and after 5 MRad + 1014neq/cm²
Best performing pixel: AC P-stop
Efficiency >99% (end of life-time).
Spatial resolution: ~6 µm
Dark rate (not shown, see backup):
  Marginal before irradiation.
  <10-6 after irradiation.
Short dimension
of 27x30µm² pixel.
Conclusion on sensor performance:
All pixels work excellent before irradiation.
Standard pixels show best spatial resolution. (see backup)
P-stop AC pixel most radiation hard, matches requirements of CBM.
13th POS conference, Oxford, Sept 2023
CBM

## Slide 8
Immunity to heavy ions (Single Event Effects)
Beam
Beam hole: 11 mm
Beam halo
Origin:
Limited focus - few ions miss hole all the time.
Beam impact.
1 cm
M. Deveaux for NA61/SHINE, SQM 2017
SPS Pb beam halo @ NA61/SHINE
as seen with MIMOSA-26
MIP (proton)
~2 fired pixels
~2x2 mm
Heavy ion:
~200 fired pixels
Minority charge carriers excited by ions may:
Switch digital electronics => Bit flip
Open unwanted conduction paths => Latch-up
(like short cut, extinguish by power cycle)
Macroscopic damage by individual ion: Single Event Effect.
Thermal destruction if ignored
Malfunction if ignored
13th POS conference, Oxford, Sept 2023

## Slide 9: Observation on Single Event Effects
 | GSI - 3.2 A MeV Ca 20 MeV cm² / mg
Latch - ups | 0
ROI (all except DAC) | 3.2 cm²
Ions on ROI [1/cm²] | > 3 x 1010 / cm²
per 5.3cm² sensor | 1.5 x 10-10 cm²
 | GSI - 3.2 A MeV Ca 20 MeV cm² / mg
Latch - ups | 0
ROI (all except DAC) | 3.2 cm²
Ions on ROI [1/cm²] | > 3 x 1010 / cm²
 | 1.5 x 10-10 cm²
Observation on DAC:
Bit flips in DAC generate current fluctuation.
  Bit correction bug to be fixed with MIMOSIS-2.
Latch-up (Ion generated short circuit, recover by power cycle):
MIMOSIS-1 seems to tolerate up to 300 MHz/cm² Ions at 20 MeV cm² / mg .
Latch-up cross-section orders of magnitude better than required (all structurs except DAC) .
Protection of steering registers to be fixed.
Data registers not protected – Expect occasional data corruption.
No latch-ups observed.
Pulses: Few 106/cm²
No beam
MIMOSIS-1
ROI = DAC
preliminary
Current [a.u.]
13th POS conference, Oxford, Sept 2023

## Slide 10: Response to different dE/dx
Known for elder sensors:
Individual pixel is not energy sensitiv but:
Pixel multiplicity scales with dE/dx [1,2].
Does this hold for MIMOSIS (dependence on depletion)?
[1] C. A. Reidel et al., "Response of the Mimosa-28 pixel sensor to a wide range of ion species and energies", NIM-A Vol. 1017 (2021) 165807
Experimental approach:
Use deuterium beam of COSY to create dE/dx higher than m.i.p.
Histogram multiplicity of hits associated to a identified track.
13th POS conference, Oxford, Sept 2023
MIP (proton)
~2 fired pixels
~2x2 mm
Heavy ion:
~200 fired pixels
[2] M. Deveaux for NA61/SHINE, SQM 2017
First observation:
Fully depleted pixels (p-stop) show marginal sensitivity to dE/dx.
Standard DC-pixel (lowest depletion) shows response. Significant?
m.i.p
1.4 GeV d
0.8 GeV d
dE/dx [m.i.p]
DC - pixel
AC - pixel
Fully depleted
Standard pixel

## Slide 11: Response to different dE/dx
CERN SPS pion
COSY 0.8 GeV d
1
5
10
Pixels/cluster
10%
20%
30%
40%
50%
Cluster distribution (DC standard, normalized)
Preliminary conclusion:
Response to increasing dE/dx observed (usefulness for PID to be studied in detail).
High rad. tolerance configuration with high depletion not compatible with good dE/dx response.
Next step: Reproduce findings in device simulation.

## Slide 12
Summary and conclusion
MIMOSIS-1 forms the first full size prototype of the MIMOSIS sensor for the CBM-MVD.
    5 µs / 5 µm time/spatial resolution.
    80 MHz/cm² peak rate.
MIMOSIS-1 irradiated with up to 3x1014 neq/cm² were tested in laboratory and in beam.
    >> 99% detection efficiency after 1014 neq/cm², 5 MRad and a combination of both
    <10-6 fake hit rate (after 1014 neq/cm², <0.05% of all pixels masked)
    ~5 µm spatial resolution before and after irradiation => Matches requirement.
    Tolerates HI with up to LET=20 MeV cm² / mg without observed latch-up.
    Response of partially depleted MIMOSIS-1 to different dE/dx may add information to CBM-PID
      (to be followed up).
MIMOSIS-1 matches the requirements of the CBM-MVD…
… and forms a mile-stone toward a sensor for higgs-factories.
13th POS conference, Oxford, Sept 2023

## Slide 13: CBM-MVD collaboration members
Julio Andary1, Benedict Arnoldi-Meadows1, Ole Artz1, Jérôme Baudot2, Grégory Bertolone2, Auguste Besson2, Norbert Bialas1, Roma Bugiel2, Gilles Claus2, Claude Colledani2, Hasan Darwish1,2,3, Michael Deveaux1,3,6, Andrei Dorokhov2, Guy Dozière2, Ziad El Bitar2, Ingo Fröhlich1,3, Mathieu Goffe2, Benedict Gutsche1, Fabian Hebermehl1, Abdelkader Himmi2, Christine Hu-Guo2, Kimmo Jaaskelainen2, Oliver Michael Keller6, Michal Koziel1, Franz Matejcek1, Jan Michel1, Frédéric Morel2, Christian Müntz1, Hung Pham2, Christian Joachim Schmidt3, Stefan Schreiber1, Matthieu Specht2, Dennis Spicker1, Joachim Stroth1,3,4, Isabelle Valin2, Roland Weirich1, Marc Winter5and Yue Zhao2 1Goethe-Universität Frankfurt, Germany 2Université de Strasbourg, CNRS, IPHC UMR 7178, Strasbourg, France 3GSI Helmholtzzentrum für Schwerionenforschung GmbH, Germany 4Helmholtz Forschungsakademie Hessen für FAIR, Germany 5IJCLab, UMR9012 – CNRS / Université Paris-Saclay / Université de Paris, France 6Facility for Antiproton and Ion Research in Europe GmbH,Germany
Supported by:
(05P21RFFC2)
13th POS conference, Oxford, Sept 2023

## Slide 14
Backup

## Slide 15: Beam tests (MIMOSIS-1)
Date | Location | Beam | Goal
13. – 14. Mar 2021 | GSI / mCBM | 1 AGeV Pb | Single-Event-Effects (SEE)
23. – 24. May 2021 | GSI / mCBM | 1 AGeV Xe | SEE
07. – 13. Jun 2021 | DESY | 5 GeV e- | Performance
19. – 26. Sep 2021 | DESY | 5 GeV e- | Performance (X-ray irradiated)
05. – 12. Oct 2021 | CERN | ~100 GeV | Performance (neutron irradiated)
14. – 20. Feb 2022 | DESY | 5 GeV e- | Performance (mixed irradiated) ++
21. – 28. Mar 2022 | COSY | 0.3 – 3 GeV p | Performance, dE/dx?
23. – 29. May 2022 | GSI/UNILAC | ~4 MeV Ca | SEE, slow fragments
01. – 07. Sep 2022 | CERN | ~100 GeV | Response to inclined tracks,…
Date | Location | Beam | Goal
13. – 14. Mar 2021 | GSI / mCBM | 1 AGeV Pb | Single-Event-Effects (SEE)
23. – 24. May 2021 | GSI / mCBM | 1 AGeV Xe | SEE
07. – 13. Jun 2021 | DESY | 5 GeV e- | Performance
19. – 26. Sep 2021 | DESY | 5 GeV e- | Performance (X-ray irradiated)
05. – 12. Oct 2021 | CERN |  | Performance (neutron irradiated)
14. – 20. Feb 2022 | DESY | 5 GeV e- | Performance (mixed irradiated) ++
21. – 28. Mar 2022 | COSY | 0.3 – 3 GeV p | Performance, dE/dx?
23. – 29. May 2022 | GSI/UNILAC | ~4 MeV Ca | SEE, slow fragments
01. – 07. Sep 2022 | CERN |  | Response to inclined tracks,…
Date | Location | Radiation
Jul – Aug 2021 | Ljubjana (TRIGA) | ~1 MeV reactor neutrons
Sep 2021 | Karlsruhe (KIT) | ~10 keV X-rays
Aug 2022 | Karlsruhe (KIT) | ~10 keV X-rays
Special thanks to IPHC for massive support in beam time preparation.
Meanwhile: 14 IPHC people (9-10 FTE) involved in MIMOSIS.
Irradiation campaigns:
Shielding for PCB-ICs
(X-rays @ KIT)

## Slide 16
CBM-MVD Sensor requirements
How to arrive there (starting from ALPIDE) based on 180 nm technology?
 | Requirement
Time resolution | ~5 µs
Spatial resolution | ~5 µm
Sensor thickness | ~50 µm
Power dissipation | 200 mW/cm²
Radiation doses (non-ionizing) | > 7x1013 neq/cm²
Radiation doses (ionizing) | > 5 Mrad
Radiation gradient on chip | 100%
HI-tolerance | 10 Hz/mm²
Rate (average/peak) | 150/700 kHz/mm²
 | Requirement
Time resolution | ~5 µs
Spatial resolution | ~5 µm
Sensor thickness | ~50 µm
Power dissipation | 
Radiation doses (non-ionizing) | > 7x1013 neq/cm²
Radiation doses (ionizing) | > 5 Mrad
Radiation gradient on chip | 100%
HI-tolerance | 10 Hz/mm²
Rate (average/peak) | 150/700 kHz/mm²
3rd High-D consortium meeting, 9-10 Feb. 2023
Mostly established by ALPIDE
(Sensor of ALICE ITS2 upgrade)
Incompatible with ALPIDE
> 20x internal bandwidth needed
per year
no safety factor
10 x ALPIDE

## Slide 17
MIMOSIS-1: Tolerance to non-ionizing radiation
Vbb=-3V
HV=10V
1 hit/frame
<0.05% of all pixels masked
Observation:
Det eff. >> 99% for 1014 neq/cm².
Resonable performance after 3x 1014 neq/cm².
AC-pixel shows better efficiency (as expected).
Dark rate < 10-6 for all pixels tested.
  => Holds for DC-pixels (not shown).
  => Matches requirements.
Preliminary
Preliminary
Vbb=-3V
Preliminary
13th POS conference, Oxford, Sept 2023

## Slide 18
Mimosis-1: In beam performance
Vbb=-3V
HV=10V
After 1014neq/cm² (end-of-lifetime):
Standard epi-layer reaches limits.
Good performance for p-stop, n-gap.
Best performance: p-stop AC pixel.
  n-gap AC worse than n-gap DC => follow up.
Preliminary
Vbb=-3V
Preliminary
Before irradiation:
All pixels show excellent performance.
13th POS conference, Oxford, Sept 2023

## Slide 19
MIMOSIS-1 spatial resolution
?
?
Preliminary
Preliminary
Observation:
Best resolution for standard epi.
  Larger charge sharing (measured, not shown).
Resolution mildly worse after irradiation.
  Bulk damage reduces cluster size.
Plots hold for small dimensions of pixel
  ~10 % worse resolution for long dimension (not shown).
Results for all pixels match CBM requirements at default 120 – 150 e threshold.
Note: Preliminary results
Still preliminary alignment.
  Fixing this may or may not eliminate outlayers
  (work in progress).
3rd High-D consortium meeting, 9-10 Feb. 2023
Approx. alignment
uncertainty.

## Slide 20
MIMOSIS-1: Tolerance to ionizing radiation
Observation:
Good detection efficiency also after ionizing doses.
Radiation seems to improve performance sometimes
  Radiation may modify pixel tuning (7 parameters)
  Probably room for improvement.
Dark rate stays below 10-7
No impact on spatial resolution within uncertainties
3rd High-D consortium meeting, 9-10 Feb. 2023
standard
p-stop
standard
p-stop

## Slide 21
1.4 GeV d
0.8 GeV d
m.i.p
2x m.i.p
4x m.i.p