## Slide 1: Diamond Detectors
Alexander Oh
University of Manchester
Advanced UK Instrumentation Training
Notes: Time: 45min
Title: “Diamond Detectors - Status and Perspectives”
Abstract:
Diamond is a promising material with remarkable mechanical and electronic properties to provide the basis for radiation resistant particle detectors for high energy particle physics (HEP). In the presentation the current state of diamond detectors in particle physics will be discussed, focussing on recent developments of 3D diamond detectors and test-beam results with prototype pixel devices.

## Slide 2: Outline
Diamond basics and detector principle
Diamond strip and pixel detectors
Radiation Hardness
3D Diamond detectors
Current and future diamond detector installations
Thanks for the material from the RD42 and ADAMAS collaborations! Very soon new working group organisation: DRD3 WG6
Part 1
Part 2

## Slide 3: PART 1
Introduction to Diamond detectors
  properties
  principle of operation
Strip and Pixel detectors
Radiation tolerance
High rate capability

## Slide 4: Challenges Ahead
Luminosity upgrades of the LHC will increase the luminosity by factor ~3.
Luminosity ~ Radiation damage.
Need new technologies in the innermost layers to survive the radiation levels.

## Slide 5: diamond
Allotrope of Carbon
Hardest natural material
Tetrahedral structure
  Sp3 bonds
Notes: Allotrope of carbon
The hardest natural material
Sp3 bonds in tetrahedral
Greatest number of atoms per unit volume
Zinc blende structure, shared with a lot of other compounds e.g. silicon. Why even though diamond is considered an insulator, can be described the same way as a semiconductor.

## Slide 6: material properties
 | Diamond | Silicon
Band Gap [eV] | 5.5 | 1.1
Lower leakage current
Notes: Silicon is diamond’s natural competitor as a solid state detector medium. Shown in this table are some material properties to consider when choosing a detector. First on this list is the Band Gap, this is the energy required to promote an electron from the valance band to the conduction band. As this is larger in diamond than Silicon, it means that diamond detectors have less leakage current or dark current, that is the current that is seen when the detector is not being irradiated.
On the other hand, this does mean that you get roughly half the signal, shown explicitly by the second quantity, describing average number of electron hole pairs produced per micron of thickness in the material by a minimum ionising particle or MIP.
Although a larger signal is generally more desirable, the larger bandgap also means that the detector will be more thermally stable, and is optically transparent, so can be used in a lit room where as silicon detectors are sensitive to light.
Diamond has strong interatomic bonds, so it requires more energy to displace a carbon atom from the its lattice. This makes it less susceptible to radiation damage, which is an ever growing area of importance in detector development, particularly in high energy physics.
Diamond has the highest thermal conductivity of any material, therefore it can quickly dissipate heat that would otherwise cause damage to the system, allowing diamond detectors to work at room temperature without degrading. Silicon detectors, on the other hand, often require extensive cooling systems and be operated below zero degrees centigrade to ensure that their performance is maintained.
The atomic number of carbon is 6, which is much closer to the effective mass of human tissue around 7.5, this is of significant importance for devices used in the medical setting.
Another interesting property is the mobilities of the electrons and holes in diamond. These govern the drift velocity of the induced charges inside the detector, therefore the timing of the signal. As detectors require faster and faster timing resolution, it should be noted that diamond has mobilities which outweigh those of silicon by some margin.
The properties of Diamond detectors are most attractive in scenarios where the device is used regularly in an area that cannot be cooled and is likely to be exposed to light, such as around people.

## Slide 7: material properties
 | Diamond | Silicon
Band Gap [eV] | 5.5 | 1.1
Average Ionisation Density for MIP [eh/µm] | 36 | 81
Lower signal
Lower leakage current
Notes: Silicon is diamond’s natural competitor as a solid state detector medium. Shown in this table are some material properties to consider when choosing a detector. First on this list is the Band Gap, this is the energy required to promote an electron from the valance band to the conduction band. As this is larger in diamond than Silicon, it means that diamond detectors have less leakage current or dark current, that is the current that is seen when the detector is not being irradiated.
On the other hand, this does mean that you get roughly half the signal, shown explicitly by the second quantity, describing average number of electron hole pairs produced per micron of thickness in the material by a minimum ionising particle or MIP.
Although a larger signal is generally more desirable, the larger bandgap also means that the detector will be more thermally stable, and is optically transparent, so can be used in a lit room where as silicon detectors are sensitive to light.
Diamond has strong interatomic bonds, so it requires more energy to displace a carbon atom from the its lattice. This makes it less susceptible to radiation damage, which is an ever growing area of importance in detector development, particularly in high energy physics.
Diamond has the highest thermal conductivity of any material, therefore it can quickly dissipate heat that would otherwise cause damage to the system, allowing diamond detectors to work at room temperature without degrading. Silicon detectors, on the other hand, often require extensive cooling systems and be operated below zero degrees centigrade to ensure that their performance is maintained.
The atomic number of carbon is 6, which is much closer to the effective mass of human tissue around 7.5, this is of significant importance for devices used in the medical setting.
Another interesting property is the mobilities of the electrons and holes in diamond. These govern the drift velocity of the induced charges inside the detector, therefore the timing of the signal. As detectors require faster and faster timing resolution, it should be noted that diamond has mobilities which outweigh those of silicon by some margin.
The properties of Diamond detectors are most attractive in scenarios where the device is used regularly in an area that cannot be cooled and is likely to be exposed to light, such as around people.

## Slide 8: material properties
 | Diamond | Silicon
Band Gap [eV] | 5.5 | 1.1
Average Ionisation Density for MIP [eh/µm] | 36 | 81
Displacement Energy [eV] | 43 | 25
Lower signal
Radiation Hardness
Lower leakage current
Notes: Silicon is diamond’s natural competitor as a solid state detector medium. Shown in this table are some material properties to consider when choosing a detector. First on this list is the Band Gap, this is the energy required to promote an electron from the valance band to the conduction band. As this is larger in diamond than Silicon, it means that diamond detectors have less leakage current or dark current, that is the current that is seen when the detector is not being irradiated.
On the other hand, this does mean that you get roughly half the signal, shown explicitly by the second quantity, describing average number of electron hole pairs produced per micron of thickness in the material by a minimum ionising particle or MIP.
Although a larger signal is generally more desirable, the larger bandgap also means that the detector will be more thermally stable, and is optically transparent, so can be used in a lit room where as silicon detectors are sensitive to light.
Diamond has strong interatomic bonds, so it requires more energy to displace a carbon atom from the its lattice. This makes it less susceptible to radiation damage, which is an ever growing area of importance in detector development, particularly in high energy physics.
Diamond has the highest thermal conductivity of any material, therefore it can quickly dissipate heat that would otherwise cause damage to the system, allowing diamond detectors to work at room temperature without degrading. Silicon detectors, on the other hand, often require extensive cooling systems and be operated below zero degrees centigrade to ensure that their performance is maintained.
The atomic number of carbon is 6, which is much closer to the effective mass of human tissue around 7.5, this is of significant importance for devices used in the medical setting.
Another interesting property is the mobilities of the electrons and holes in diamond. These govern the drift velocity of the induced charges inside the detector, therefore the timing of the signal. As detectors require faster and faster timing resolution, it should be noted that diamond has mobilities which outweigh those of silicon by some margin.
The properties of Diamond detectors are most attractive in scenarios where the device is used regularly in an area that cannot be cooled and is likely to be exposed to light, such as around people.

## Slide 9: material properties
 | Diamond | Silicon
Band Gap [eV] | 5.5 | 1.1
Average Ionisation Density for MIP [eh/µm] | 36 | 81
Displacement Energy [eV] | 43 | 25
Thermal Conductivity [W/cm.K] | 10-20 | 1.5
Lower signal
Radiation Hardness
Room temperature operation
Lower leakage current
Notes: Silicon is diamond’s natural competitor as a solid state detector medium. Shown in this table are some material properties to consider when choosing a detector. First on this list is the Band Gap, this is the energy required to promote an electron from the valance band to the conduction band. As this is larger in diamond than Silicon, it means that diamond detectors have less leakage current or dark current, that is the current that is seen when the detector is not being irradiated.
On the other hand, this does mean that you get roughly half the signal, shown explicitly by the second quantity, describing average number of electron hole pairs produced per micron of thickness in the material by a minimum ionising particle or MIP.
Although a larger signal is generally more desirable, the larger bandgap also means that the detector will be more thermally stable, and is optically transparent, so can be used in a lit room where as silicon detectors are sensitive to light.
Diamond has strong interatomic bonds, so it requires more energy to displace a carbon atom from the its lattice. This makes it less susceptible to radiation damage, which is an ever growing area of importance in detector development, particularly in high energy physics.
Diamond has the highest thermal conductivity of any material, therefore it can quickly dissipate heat that would otherwise cause damage to the system, allowing diamond detectors to work at room temperature without degrading. Silicon detectors, on the other hand, often require extensive cooling systems and be operated below zero degrees centigrade to ensure that their performance is maintained.
The atomic number of carbon is 6, which is much closer to the effective mass of human tissue around 7.5, this is of significant importance for devices used in the medical setting.
Another interesting property is the mobilities of the electrons and holes in diamond. These govern the drift velocity of the induced charges inside the detector, therefore the timing of the signal. As detectors require faster and faster timing resolution, it should be noted that diamond has mobilities which outweigh those of silicon by some margin.
The properties of Diamond detectors are most attractive in scenarios where the device is used regularly in an area that cannot be cooled and is likely to be exposed to light, such as around people.

## Slide 10: material properties
 | Diamond | Silicon
Band Gap [eV] | 5.5 | 1.1
Average Ionisation Density for MIP [eh/µm] | 36 | 81
Displacement Energy [eV] | 43 | 25
Thermal Conductivity [W/cm.K] | 10-20 | 1.5
Atomic Number | 6 | 14
Lower signal
Radiation Hardness
Room temperature operation
Tissue equivalence
Lower leakage current
Notes: Silicon is diamond’s natural competitor as a solid state detector medium. Shown in this table are some material properties to consider when choosing a detector. First on this list is the Band Gap, this is the energy required to promote an electron from the valance band to the conduction band. As this is larger in diamond than Silicon, it means that diamond detectors have less leakage current or dark current, that is the current that is seen when the detector is not being irradiated.
On the other hand, this does mean that you get roughly half the signal, shown explicitly by the second quantity, describing average number of electron hole pairs produced per micron of thickness in the material by a minimum ionising particle or MIP.
Although a larger signal is generally more desirable, the larger bandgap also means that the detector will be more thermally stable, and is optically transparent, so can be used in a lit room where as silicon detectors are sensitive to light.
Diamond has strong interatomic bonds, so it requires more energy to displace a carbon atom from the its lattice. This makes it less susceptible to radiation damage, which is an ever growing area of importance in detector development, particularly in high energy physics.
Diamond has the highest thermal conductivity of any material, therefore it can quickly dissipate heat that would otherwise cause damage to the system, allowing diamond detectors to work at room temperature without degrading. Silicon detectors, on the other hand, often require extensive cooling systems and be operated below zero degrees centigrade to ensure that their performance is maintained.
The atomic number of carbon is 6, which is much closer to the effective mass of human tissue around 7.5, this is of significant importance for devices used in the medical setting.
Another interesting property is the mobilities of the electrons and holes in diamond. These govern the drift velocity of the induced charges inside the detector, therefore the timing of the signal. As detectors require faster and faster timing resolution, it should be noted that diamond has mobilities which outweigh those of silicon by some margin.
The properties of Diamond detectors are most attractive in scenarios where the device is used regularly in an area that cannot be cooled and is likely to be exposed to light, such as around people.

## Slide 11: material properties
 | Diamond | Silicon
Band Gap [eV] | 5.5 | 1.1
Average Ionisation Density for MIP [eh/µm] | 36 | 81
Displacement Energy [eV] | 43 | 25
Thermal Conductivity [W/cm.K] | 10-20 | 1.5
Atomic Number | 6 | 14
Electron Mobility [cm2/V.s] | 1900-3800 | 1350
Hole Mobility [cm2/V.s] | 2300-4500 | 480
Lower signal
Radiation Hardness
Room temperature operation
Tissue equivalence
Lower leakage current
Fast signal
Notes: Silicon is diamond’s natural competitor as a solid state detector medium. Shown in this table are some material properties to consider when choosing a detector. First on this list is the Band Gap, this is the energy required to promote an electron from the valance band to the conduction band. As this is larger in diamond than Silicon, it means that diamond detectors have less leakage current or dark current, that is the current that is seen when the detector is not being irradiated.
On the other hand, this does mean that you get roughly half the signal, shown explicitly by the second quantity, describing average number of electron hole pairs produced per micron of thickness in the material by a minimum ionising particle or MIP.
Although a larger signal is generally more desirable, the larger bandgap also means that the detector will be more thermally stable, and is optically transparent, so can be used in a lit room where as silicon detectors are sensitive to light.
Diamond has strong interatomic bonds, so it requires more energy to displace a carbon atom from the its lattice. This makes it less susceptible to radiation damage, which is an ever growing area of importance in detector development, particularly in high energy physics.
Diamond has the highest thermal conductivity of any material, therefore it can quickly dissipate heat that would otherwise cause damage to the system, allowing diamond detectors to work at room temperature without degrading. Silicon detectors, on the other hand, often require extensive cooling systems and be operated below zero degrees centigrade to ensure that their performance is maintained.
The atomic number of carbon is 6, which is much closer to the effective mass of human tissue around 7.5, this is of significant importance for devices used in the medical setting.
Another interesting property is the mobilities of the electrons and holes in diamond. These govern the drift velocity of the induced charges inside the detector, therefore the timing of the signal. As detectors require faster and faster timing resolution, it should be noted that diamond has mobilities which outweigh those of silicon by some margin.
The properties of Diamond detectors are most attractive in scenarios where the device is used regularly in an area that cannot be cooled and is likely to be exposed to light, such as around people.

## Slide 12: Natural and synthetic diamond
Natural diamonds have a high defect concentration
  Grow in different structure to synthetic diamonds
  Compete with jewellery market
  There are radiation sensors using natural diamond!
CVD
HPHT
Mined
Notes: Cvd plate
Natural octahedral shape
Hptht octohedral

## Slide 13: Diamond
1941 – Diamond as particle detector (Stetter)
1953- CVD process, synthesis of diamond (Eversole)
~1980 – polycrystalline CVD diamond.
1994 – first diamond strip detector
1996 – first diamond pixel detector
2011 – first 3D diamond detector
G. Stetter. Verhandl. deut. physik. Ges., 22:13, 1941.

## Slide 14: Synthesis of Diamond
Chemical Vapour Deposition (CVD) of diamond in the graphit phase space.

## Slide 15: Synthesis of Diamond
Hydrogen terminated substrate surface
Methan and Hydrogen gas are heated with microwaves to form a plasma
Radicals form

## Slide 16: Synthesis of Diamond
Hydrogen atoms are replaced with Carbon

## Slide 17: Synthesis of Diamond
SP2 bonds (graphite) are weaker then SP3 bonds (diamond)
Hydrogen radicals will etch away graphite, but leave diamond
A diamond film is grown

## Slide 18: Development of CVD Diamond for detector applications
Today two main manufacturers of detector grade diamond
  ElementSix Ltd
    large polycrystalline wafers
    single crystal diamonds
  II-VI Semiconductors
    large polycrystalline wafers
    relatively recent entry
Alternative sources
  Diamond on Iridium (DoI) (Audiatec, Germany)
  Hetero-epitaxially grown -> large area
  Highly oriented crystallites.
5’’ waver
single crystal
DoI, 155ct
Notes: 5 inch wafer (dots are 1cm apart)

## Slide 19
Principle of detector operation
R
e
h
Substrate-Side
Growth-Side
d
d
collected charge
“collection distance”
e = Q / Q0
collection efficiency
d

## Slide 20
MIP signal is measured, expressed in charge collection distance defined as d [mm]=Qm[e] / 36 [e/mm]
More accurately the “Schubweg” (l).
  Relation between MIP signal efficiency e, ”collection distance” d, and “Schubweg” l:
e=2l[1-l/d(1-exp(-d/l))]
MIP signal efficiency
l/d

## Slide 21
Principle of detector operation
R
e
h
Substrate-Side
Growth-Side
d
dQ
collection distance
Schubweg

## Slide 22: Development of CVD Diamond for detector applications
Impressive progress over the last 20 years.
Current state of the art for polycrystalline CVD diamond d ~ 250 mm (~9000 e/MIP) commercially available.
R
e
h
Substrate-Side
Growth-Side
d
d

## Slide 23: Development of CVD Diamond for detector applications
Impressive progress over the last 20 years.
Single crystal diamond ~ 100% efficient
Diamond on iridium ~ 97% efficient
J. Appl. Phys. 97, 073704 (2005)
GSI annual report 2013
DoI
single crystal
241Am 5.4 MeV a

## Slide 24: Strip Detectors
First position sensitive diamond detectors where strip detectors.
Many prototypes tested starting around 1994

## Slide 25
The charge signal is picked up by the strip(s) next to the particle track.
The charge is shared by multiple strips if the charge collection is incomplete.
The position of the particle track can be reconstructed by calculating the charge weighted impact point (Center of Gravity)
Principle

## Slide 26
A Diamond Testbeam Telescope
Diamond detectors

## Slide 27
strip
strip
~10ke mean signal

## Slide 28
Uniformity in Charge Collection of CVD Diamonds
Measured with MIPS
Polycrystalline CVD diamond exhibits non-uniform signal response due to crystallite structure.
Similar patterns observed as with photon beam measurement

## Slide 29: Irradiated Strip Detectors
Proton Irradiation
15% loss of S/N after 2e15 p cm-2
35% improvement in resolution

## Slide 30: Pixel Detectors
Several prototypes of Diamond pixel detectors have been developped and tested since around 1996.
Read-out chips use ROC (CMS), FE-I4 (ATLAS)
More recently tested 3D pixel detectors (see later).
Some historic examples in the following.

## Slide 31
Diamond Pixel Detectors

## Slide 35
Efficiency = 80%
Resolution = digital
Results from Atlas Diamond Pixel Detectors
s=115mm
s=14mm

## Slide 36
Large track residuals
Non-uniformity of response qualitatively reproduces by modeling
Results from Atlas Diamond Pixel Detectors
Tommaso Lari (INFN)
Alexander Oh (CERN)
Norbert Wermes (University Bonn)

## Slide 37
Task: Find a simple model to explain the peculiarities of CVD diamond
    priming
    polarisation
    non-uniformities
First Step: Make an assumption about the spatial distribution of carrier lifetime
    lifetime correlates to the crystallite structure
    lifetime depends on the distance to the nearest grainboundary
    lifetime can be parametrized by
Second Step: Develop a simple growth model to approximate the growth morphology found in CVD diamond
    in 3D crystallites start to grow with equal probability with a given seed density
    the space points claimed by concurring seeds are conquered randomly
    the probability to sucessfully claim is weighted with the crystallites size
S. Waidmann, Inst. f. Festkorper und Werkstofforschung, Dresden
Modelling

## Slide 38
The Growth Generator can be adjusted for different lateral growth speeds:
k=0.5
k=0.8
k=0.9
k=2 vvert vlateral-1
Modelling

## Slide 39: Modelling
The carrier drift process is simulated for single carriers (electrons and holes) accordingly to the lifetime distribution.
Recombination and trapping is distinguished.
The electric field due to trapped charge is calculated.
The drift path for each carrier is determined by the local electric field and thermal diffusion with 1mm resolution.
Recombination and trapping is distinguished, priming effect can be modelled….
trapped electrons
trapped holes
polarisation field

## Slide 40: Material Studies
Non Uniformities qualitatively reproduced by modeling

## Slide 41: Modelling
Comparison of track residuals in the model with and without grain structure
Model with grain structure can qualitatively reproduce residuals observed with diamond pixel detectors.
Tommaso Lari (INFN)
Alexander Oh (CERN)
Norbert Wermes (University Bonn)

## Slide 42: Radiation Tolerance

## Slide 43: Tests of Radiation Tolerance
Irradiate with proton, pions and neutrons.
  Energies within the expected radiation profile at HL-LHC.
  HL-LHC fluence requirement about 2e1016 neq.
Paul Miagawa, 31st RD50 Workshop on Radiation hard
semiconductor devices for very high luminosity colliders
FLUKA simulation
 | Proton | Pion | Neutron
Energy | 25MeV – 24GeV | 300 MeV | 1-10 MeV
Fluence | 1.27e16 p cm-2 | 6e14 p cm-2 | 1.3e16 n cm-2
Notes: Paul Miagawa, 31st RD50 Workshop on Radiation hard
semiconductor devices for very high luminosity colliders
Say: Goal is 10e17

## Slide 44
Assume simple effective model for radiation damage:
Radiation damage constant
is fitted with simple model:
damage constant
particle flux

## Slide 45: Radiation Tolerance: Characterization
  Typical Landau Spectra after irradiation of pCVD.
  For pCVD see reduction of FWHM / MP with irradiation.
    Expected from polycrystalline nature of material!
    Single crystal material almost flat.
800 MeV protons

## Slide 46: Radiation Tolerance: Characterization
  Resistivity
    No dose dependence.
    Due to large bandgap no significant temperature dependence at RT or below.
cf. Silicon
S Seidel 2014 JINST 9 C01013

## Slide 47: Radiation Tolerance: Characterization
Damage factor k is determined for each sample.
pCVD diamonds are offset by l0 to account for initial finite carrier lifetime.
Final damage factor averaged over all samples.
24 GeV p
RD42 Preliminary
single crystal
poly1
poly2
poly3

## Slide 48: Radiation Hardness
Describe radiation damage using Norget-Robinson-Torrens theorem to predict displacements per atom (DPA).
  (M. Guthoff et al., arXiv:1308.5419)
  Diamond displacement energy: 43.3 eV
  Reasonable agreement for E>100MeV.
Notes: We were (are?) working hard on the paper that describes these
results. At those working meetings Harris has impressed upon us
all the things that contribute to the systematic uncertainties on
the relative 'k' factors. Not having completed an analysis of the
systematics we are reluctant to quote only the statistical
uncertainties from the fits. Marko has shown those at internal
RD42 meetings, but we have not shown them outside. I think all
you can do at this point is point out the monotonicity (or lack
there-of) of the 4+1 points plotted in the figure to give "some
sense of the size of the uncertainties".

## Slide 49: Radiation Tolerance
24 GeV protons
  kl = 0.67 ± 0.04 ×10-18 cm2mm-1
  polycrystalline diamond sample offset by F~5×1015 to account for existing traps.
  Poly and single crystal diamond show consistent damage constants.
L. Baeni ETHZ Thesis
https://www.research-collection.ethz.ch/handle/20.500.11850/222412
70 MeV p
800 MeV p
24 GeV p
Notes: Look at Nicola’s write up.
Plots has all the data.

## Slide 50: Radiation Tolerance
Summary of RD42 irradiation results:
*normalized to 24GeV protons
*
54
Notes: 800 MeV proton update: https://indico.cern.ch/event/505107/contributions/1182410/attachments/1272791/1887710/RD42Meeting-May2016-800MeVResults.pdf

## Slide 51: Radiation Tolerance: Comparison to Si
k factors typically 2-3 times higher for Silicon.
A comparison to Si needs to take into account:
  leakage current
  capacitance
Possible figure of merit Signal to noise ratio:
24 GeV p
d=200mm
24GeV p
JINST 7 (2012) P09009

## Slide 52: High rate capability

## Slide 53: High Rate tests
Tests the pulse height as function of particle rate.
Test single and poly crystalline diamond.
Irradiated and un-irradiated.
Investigations triggered by indication of rate dependence of
of single crystal diamond pixel detector installed in CMS in 2012.
Notes: High rate tests need update

## Slide 54: High Rate tests
single and poly sample irradiated with 5×1013 reactor n.
Tested with 250MeV pions.
Slight rate dependence observed in irradiated single crystal sample.
No rate dependence observed for irradiated polycrystalline sample.
single
poly
RD42
RD42
RD42

## Slide 55: END OF PART 1
In part 2 next week we look at:
  3D Diamond detectors
  Application of diamond detectors in HEP

## Slide 56: Diamond Detectors
Alexander Oh
University of Manchester
Advanced UK Instrumentation Training
Notes: Time: 45min
Title: “Diamond Detectors - Status and Perspectives”
Abstract:
Diamond is a promising material with remarkable mechanical and electronic properties to provide the basis for radiation resistant particle detectors for high energy particle physics (HEP). In the presentation the current state of diamond detectors in particle physics will be discussed, focussing on recent developments of 3D diamond detectors and test-beam results with prototype pixel devices.

## Slide 57: PART 2
3D Diamond detectors
Application of diamond detectors in HEP

## Slide 58: 3D diamond detectors

## Slide 59
Planar
3D
Electrode spacing determines drift distance to induce 1e charge.
3D has shorter electrode spacing compared to planar.
Charge carriers need less drift distance (and time) in 3D then in planar to induce equal signal.
Influence of traps and resulting limited lifetime suppressed in 3D.
3D Diamond Detectors
Notes: 3D detector section changes
oxford method, SLM theory and practice
two-photon absorption theory and results
simulation results
tribic and ibic comparing hexa and cubic cells
first test-beam results with cern testbeam (check with harris)

## Slide 60: 3D Diamond Research
  Lasered graphitic structures in pCVD (’10)
  Single crystal with column structure (‘11)
  Pico second laser for improved graphitic electrodes (‘12)
  Several Prototypes tested at Diamond Light Source CERN test beam RBI proton beam
  Clinical Photon beam
Manchester,
Oxford, CEA Saclay, CERN, ETH Zurich,
Ohio State University
  Femto second laser processing and TCAD simulation (’13,’14)
  Phase modulated laser processing (’15,’16)
  Polarised & Phase modulated laser processing (’17,’18)

## Slide 61: 3D Diamond Research - A relatively young field
Laser induced phase change in diamond.
  E.g. T.V. Kononenko et al, Diamond & Related Materials 18 (2009) 196–199 “Femtosecond laser microstructuring in the bulk of diamond “
3D “Pad” detector
  A. Oh, B. Caylar, M. Pomorski, T. Wengler, Diamond and Related Materials, 38 , (2013), ”A novel detector with graphitic electrodes in CVD diamond”
  S. Lagomarsino et al, Appl. Phys. Lett. 103, 233507 (2013), “Three-dimensional diamond detectors: Charge collection efficiency of graphitic electrodes”
3D “strip array” detector with position resolution.
  E.g. F. Bachmaier et al, NIM A, 786, (2015) 97-104, “A 3D diamond detector for particle tracking”
Radiation damage studies.
  Eg. S. Lagomarsino et al, Applied Physics Letters 106, 193509 (2015) “Radiation hardness of three-dimensional polycrystalline diamond detectors”
Improvements in graphitization process.
  Eg. B. Sun et al., Applied Physics Letters 105, 231105 (2014), “High conductivity micro-wires in diamond following arbitrary paths”
3D pixel detectors
  RD42, CERN-LHCC-2018-015 ((2018), Development of Diamond Tracking Detectors for High Luminosity Experiments at the LHC, HL-LHC and Beyond
  L. Anderlini et al, Front. Phys., 04 November 2020, Fabrication and Characterisation of 3D Diamond Pixel Detectors With Timing Capabilities

## Slide 62: Fabrication
Conductive columns are created by changing diamond into a graphitic material with a short laser pulse:

## Slide 64: Laser Set-up in Manchester
Laser specs:
Wavelength: 800 nm
Repetition rate: 1 kHz
Pulse duration: 100 fs
Max power: 1 W
Just moved to MECD!
Notes: 3D diamond detectors are fabricated at eh university of Manchester using this set-up
The key features are
SLM to correct for abbherationd
The 4F distance system which reduces the abbherations that might occur in the section of beam path between the SLM image and the objective lens
The stage can move with sub-micron precision in the 3 directional axes which is critical for fabricating the fine structures in the diamond

## Slide 65
University of Manchester, Laser Processing Research Center.
Wavelength = 800 nm
Repetition rate = 1 kHz
Pulse duration = 100 fs
Spot size = 10μm
Pulse Energy ~ 1 μJ
Spatial light modulator

## Slide 66: SLM – Phase Spatial Light Modulation
Comparison SLM vs standard process.
66
Simulated
depth = 40μm
Measured
depth = 40μm
depth = 80μm
depth = 130μm
66
 | Std. | SLM
Resistivity | 1 Wcm | 0.1 Wcm
Diameter | ~3mm | ~1mm
Diamond to graphite ratio | ~4 | ~0.2

## Slide 67: SLM – Phase Spatial Light Modulation
Comparison SLM vs standard process.
67
67
 | Std. | SLM
Resistivity | 1 Wcm | 0.1 Wcm
Diameter | ~3mm | ~1mm
Diamond to graphite ratio | ~4 | ~0.2
STD
SLM

## Slide 68: X-polariser image
4 mm
4 mm
Optical grade scCVD diamond.
4 mm
4 mm
Post processing.

## Slide 69
SEM surface image
Seed surface
Exit surface
With SLM
10um/s
400nJ
2um
5um
Without SLM
10um/s
400nJ
Without SLM
10um/s
400nJ
5um
2um
5um

## Slide 70: Making the thinnest column
More energy = thicker column
Non-linear breakdown of diamond
  More focused beam spot at depth makes thinner column
  Immersion Oil helps to reduce refraction loss from air-diamond interface
    SLM still key!
Lopez Paz, I., Allegre, O., Li, Z., Oh, A., Porter, A. and Whitehead, D. (2019), Study of Electrode Fabrication in Diamond with a Femto-Second Laser. Phys. Status Solidi A, 216: 1900236.
oil
Where we are now!
Notes: Need to say why thin columns are better over thicker columns
Show that the set up reaches down here, thin columns resisitivity limits

## Slide 71: Moving sideways
Also have the possibility to move in arbitrary direction
Wavefront correction needs to be tailored in real-time
  For vertical columns have mainly spherical corrections
  For horizontal processing, the correction is ~elliptical
  Gets even trickier at depth >200µm
Sun, B., Salter, P. S., & Booth, M. J. (2014). High conductivity micro-wires in diamond following arbitrary paths. Applied Physics Letters, 105(23), 231105.

## Slide 72
Internal structure
Patrick S. Salter et al., APPLIED PHYSICS LETTERS 111, 081103 (2017)
Prepare sample with horizontal graphitic wires.
STEM image of wire cross section.
Optical and spectral data points to micro-cracks and nano-clusters of sp2 bonded carbon.
Micro wires are not macroscopic structures!

## Slide 73
Parameter space scan
 | 5um/s | 10um/s | 20um/s | 30um/s
100nJ | x | x |  | 
200nJ | x | x | x | 
300nJ |  | x | x | x
400nJ |  | x | x | x
500nJ |  |  | x | x
600nJ |  |  |  | x
Laser translation speed
Laser beam energy
Repeat with and without SLM correction.
Patrick Salter, Oxford Iain Haughton, AO, Manchester

## Slide 74
Metallisation
X polarisers
Metallisation mask
(Seed surface)
Metallisation:
Chromium-Gold
Seed surface structured.
Exit surface pad.
Microscope
100um
100um

## Slide 75
IV curves
Ohmic and barrier potential curves observed.
Ub [V]
Current [A]
Current [A]
Ub [V]
Continuous.
Bulk effect?
Micro gaps?

## Slide 76
Barrier potential
2Uϕ
Ub [V]
Current [A]

## Slide 77
Resistance measurement
ΔU-b
Ub [V]
Current [A]
ΔI-
R = ΔUb/ΔI
ΔI+
ΔU+b

## Slide 78
With SLM
Resistance [MOhm]
Metallisation mask

## Slide 79
With SLM
Resistance [MOhm]

## Slide 80
With SLM
Uϕ [V]

## Slide 81
Raman – with SLM
Resistance increase as power law  multi-photon process.
Clear discrepancy at 30um/s.
Resistance

## Slide 82
Raman – with SLM
Barrier energy
Uϕ [V]
Reduction in barrier with increased energy.
Discrepancy at 30um/s.

## Slide 83
Raman – with SLM
Multiple passes reduces resistance and increases uniformity of the columns.
Multiple passes

## Slide 84
Raman – with SLM
Multiple passes also reduces Uϕ.
Multiple passes
Uϕ [V]

## Slide 85
Raman – with SLM
With and without SLM
Uϕ [V]
SLM clearly beneficial to reduce
resistance and barrier potential

## Slide 86
SLM parallel processing?
y/ pixel
x/ pixel
Phase shift/ π
normalised intensity

## Slide 87: 3D Diamond detector tests with relativistic charged particles
Types
  100x100um cell size ganged to form strips
  100x100um cell size, bonded to pixel read-out
  50x50um cell size, bonded to pixel read-out
All detectors made from polycrystalline diamond.
Beam tests
  CERN beam line H6 : protons ~ 120 GeV/c
  PSI : pions ~ 250 MeV/c
Thanks for material from the RD42 collaboration!

## Slide 88: 3D Diamond prototype
Proto-type
  Strip detector with back side contact
  3D metal only pattern
  3D metal + graphitic columns
  Cubic cell base size 150mm
  99 cells
Measure response with 120 GeV protons.
Paper published NIMA "A 3D diamond detector for particle tracking", NIM A, 786 (2015)
Strip
3D mask no columns
3D mask with columns
Read-out
Bias
Cell

## Slide 89
Analysis steps
Ub(3D)=40V
Ub(strip)=500V
Identify continuous region of intact cells for analysis.
Exclude contribution of negative signals.
Average charge Strip: 16.8ke 3D: 15.9ke
MP: Strip: 14.7ke 3D: 15ke
3D and Strip show comparable response.
Conclusion -> 3D works!

## Slide 90: Test of first 3D pCVD diamond detectors
Ub(3D)=75V
Ub(strip)=500V
Selected 16 adjacent cells
Notes: Describe red/Blue
----- Meeting Notes (16/02/16 14:40) -----
change range of plot such that nothing is white

## Slide 91: Test of first 3D pCVD diamond detectors
Red line estimate the Mean for Full Charge Collection (100%)
71% of
Full Charge Collection, corresponding to ~13 ke.
Highest charge collection ever measured for pCVD diamonds
Measured Mean 904
Full Charge Mean: 1269
Notes: scCVD 1117 @ 440 um
1117/440*500 = 1269 ADC

## Slide 92
Large area 3D, pCVD, 100x100
In May/Sept 2016 tested the first full 3D device fabricated in pcCVD with three dramatic improvements:
  1. An order of magnitude more cells (1188 vs 99).
  2. Smaller cell size (100um vs 150um).
  3. Higher column production efficiency (>99% vs ~90%).
HV side
Readout side

## Slide 93
Large area 3D, pCVD, 100x100
Largest charge collection to date in pcCVD diamond!
  >85 % of charge collected in continuous region, about twice as much as planar.
x coordinate (um)
y coordinate (um)
avrg. signal / ADC
charge / ADC
number of entries
RD42 Preliminary

## Slide 94
First assembly with ROC chip produced.
  Bump bonded in Princeton.
  Cr-Au on bias side.
  Ti-W under-bump metal.
  Indium bumps on sensor.
Bump pads
3D sensor
3D sensor bonded to ROC
Pixel 3D, pCVD, 100x100
  Bump pads
  Columns
  Bias grid

## Slide 95
Production of first pixel device using CMS readout electronics.
Active region 3x3 mm with cell size ~100x100 um.
Pixel 3D, pCVD, 100x100

## Slide 96
  3D diamond device and Silicon reference planar device.
  Pixel threshold 1500e.
  Check hit efficiency over time.
  Device works!
3D Diamond
Silicon planar
e(hit) = 99.3%
e(hit) = 98.5%
RD42 Preliminary
RD42 Preliminary
Pixel 3D, pCVD, 100x100

## Slide 97: Next generation 3D Diamond
  Produced 3500 Cell pixel protoype, 50x50um cell size.
  Sample production:
    Oxford (2x cubic cells)
    Manchester set-up in progress (expected production date end of month.)
    Bump bonding
      For ROC (CMS) Princeton.
      For FE-I4 (ATLAS) IFAE.
  Data taking in August 2017 at PSI.

## Slide 98: First 50x50 mm cell 3D Diamond (2017)
  Readout with CMS pixel readout.
  Bump bonding issue in upper right edge (Indium bump deposition machine not working properly)
  6 columns (3x2) ganged together.
  Preliminary hit efficiency 99.2%
  Preliminary: Collect >90% of charge!
  Rate dependence tested with 10 kHz/cm-2 and 10 MHz/cm-2 -> no dependence observed.
RD42 Preliminary
RD42 Preliminary

## Slide 99: 3D diamond radiation tolerance
Tested a 3D device irradiated to 3.5e15 p/cm2 and compare to a planar diamond device at same fluence.
Signal reduction: Planar 3D
Assuming scaling is similar 3D should operate at 10e17 p/cm2

## Slide 100: Applications in HEP

## Slide 101: Applications in HEP
Vertex detectors with CVD Diamond are not considered yet as an option for LHC.
For Beam monitoring CVD Diamond is used at CMS and ATLAS at the LHC.
  BaBar and Belle test already CVD Diamond in their beam monitoring system.

## Slide 102: Diamond in current HEP experiments
Beam monitors to protect experiments against beam losses at the LHC, CERN.
  For Silicon Vertex systems careful monitoring is crucial.
  Beam monitors have to be radiation hard.
  Abort beam when monitors signal dangerous beam conditions.
  False signals must be avoided.
During run-1 diamond beam monitors operated in ATLAS, CMS, and LHCb.
Previously diamond beam monitors were installed in BaBar(SLAC), CDF & D0 (Tevatron).
ATLAS DBM
Beam
monitors
Particle
trackers

## Slide 103: ATLAS beam conditions monitor
Use 2x polycrystalline CVD diamonds per station (10 x 10 mm).
4 stations on each side of the ATLAS pixel detector
  z = ±183.8 cm (~12.5ns) and r ~ 5 cm

## Slide 104: ATLAS beam conditions monitor
Single particle counting with s=0.7ns.
  Distinguish between collision events and out-of-time background.
Good stability
  Used for luminosity determination.

## Slide 105: Run 2: ATLAS Diamond Beam Monitor
8 mini-trackers of 3 planes each using pixel-detectors.
polycrystalline diamond sensors, 18mm x 21mm, d>250mm.
bump-bonded to FE-I4 pixel read-out chip.
  336 x 80 pixels
  pixel size : 50mm x 250 mm
Purpose:
  Bunch-by-bunch luminosity monitor (aim < 1 % per BC per LB)
  Bunch-by-bunch beam spot monitor

## Slide 106: Run 2: ATLAS Diamond Beam Monitor
Installed in ATLAS during LS1, but switched off due to unexpected death of Si and Diamond modules.
DBM recommissioned in 2017/18 with 50% working modules.
Notes: Radial distance of the closest approach of the projected particle tracks to the interaction point as recorded by a single DBM telescope. The Gaussian fit yields a D0 with a mean of -8.4~mm and an RMS of 7.2~mm. No alignment has yet been included. The two side peaks of the tracks in the distribution for the tracks with unpaired bunches correspond to secondary particles created by collisions of protons with the pipe and the surrounding material. Data acquired during one run in July 2015.

## Slide 107: Examples of diamond detectors in related areas
Synchrotron labs
  beam position monitor
Radiation Therapy
  small field dosimetry
Heavy Ion (GSI, FAIR)
  beam diagnostic
  particle tracking and TOF
  hadron spectroscopy
[1] M. Pomroski, CEA-LIST, MRS Fall meeting, Boston 28/11/2012
[2] F. Marsolat et al. / Diamond & Related Materials 33 (2013) 63–70
3 µm thick membrane
in 40 µm thick scCVD
[2]
[1]
scCVD dosimeter, 0,4 mm3 active vol. [2]

## Slide 108: Detectors for Heavy Ion
In-beam START-VETO detectors for HADES and CBM
  High beam intensities (107 #/s).
  Protons (MIPS) up to very heavy ions (Au, U)
  Excellent time resolution (30-100 ps)
  Radiation hard up to 1012 /cm Au ions
Start: SC diamond,
70 m thickness, 4.7mm x 4.7mm
Veto: PC diamond,
100 m thickness, 20 mm x 20 mm
4,7 mm
2
3
4
1
5
6
7
8
4.7mm
Start - Halo
Veto detector,
reference detector
HI, T0 sensor
J.Pietraszko, 21st CBM Collab. Meeting

## Slide 109: Upgrade for LHC Phase-2: BCM’ modules
abort
luminosity

## Slide 110: First BCM’ 3D test structure with new “cage” design
Fabricated March 2020 (~week of UK Lockdown) first 3D diamond device with horizontal ganging
Test structure has horizontal wires at depths of 125 and 375 µm in a 500 µm thick substrate (ρ measurement)
Metallised 3D detector with alternating ganging (~model) read out in current mode in RD42 Zagreb Testbeam 06-2021, second version tested April 2022
Notes: The first test structure for this design was fabricated in early march of 2020. Days before the first UK lockdown. On the right of this image are horizontal graphitic wires fabricated at depths of 125, 250 and 375 microns below the surface of the diamond. As you can see the darker the line looks the closer towards the top surface the wire is. These are used to measure the resistivity of the wires. On the left side of the image are 3D diamond test structes with internal ganging. The structures have a configuration with alternating ganging of the field and sense columns, similar to what is shown in this TCAD model. The successful implementations of these test structures were metallised and read out in current mode at the proton microbeam facility as part of a Rd42 test beam that took place last month.
Preliminary results that cant yet be shown from this, work in progress.
Theses are the structures that are available to make now
This is an image of the first proof of concept device that was made just before the UK lockdowns. Sadly I can’t yet how you results from this, but there has been a second (improved iteration) that will be a large part of my thesis

## Slide 111: Dark matter
Dark matter detection via elastic spin-indepent interaction with Carbon nuclei.
Diamond operated as cryogenic calorimeter.
Reach of first prototypes is comparable to CRESST silicon sensors.
Promising if scaled up (mass, time, background shielding)

## Slide 112: Physics motivation
Sensitive to dark matter and ALPs.
Predictions for 1 kg Y *
* 1kg high purity synthetic diamond costs O(1 M)

## Slide 113: Current Diamond TES landscape
Relatively new field.
Diamond proposed as a DM detector.

## Slide 114: Why diamond?
Operation as Transition Edge Sensor (TES)
  W Al metalisation
Diamond has high Debye temperature:
  high thermal conductivity
  excellent phonons transport.
Light nucleus (A=12)
  probe lower DM masses

## Slide 115: Summary
Diamond systems are used as beam and luminosity monitors in current HEP experiments and foreseen for future experiments.
Radiation hardness and rate dependence has been studied.
3D diamond has been demonstrated to work.
The understanding of diamond as a detector material is advancing.

## Slide 116: EXTRA MATERIAL

## Slide 117: 3D Detector Characterization
Proton Micro-beam: 4.5 MeV p

## Slide 118
0 50 100
Depth (um)
Energy loss (a.u.)
4.5 MeV protons
Read-out
channel 1
Proton micro-beam measurements
Energy deposition -
Bragg peak
Single particle beam.
Rate ~ 1kHz.
Beam position resolution < 2μm.
Read-out
channel 2
3D diamond
detector

## Slide 119
IBIC - setup
Instantaneous
current
pulses
Integrating
amplifiers
Shaping
amplifiers
Multi Channel Analyser and Data Aquisition
Two read-out channels and amplification chains.
Q
~ 2μs
~ 0.1μs

## Slide 120
0 50 100
Depth (um)
Energy loss (a.u.)
4.5 MeV protons
Read-out
channel 1
Proton micro-beam measurements
Energy deposition -
Bragg peak
Single particle beam.
Rate ~ 1kHz.
Beam position resolution < 2μm.
Read-out
channel 2
3D diamond
detector

## Slide 121
(TR)IBIC (Time Resolved) Ion Beam Induced Current
Instantaneous
current
pulses
Integrating
amplifiers
Shaping
amplifiers
Multi Channel Analyzer and Data Acquisition
Two read-out channels and amplification chains.
Q
~ 2μs
~ 0.1μs

## Slide 122
Signal efficiency at +20V
Channel 1
Channel 2
Ch. 1
Ch. 2
Ch. 2
Electronics chain calibrated with silicon detector.
CCE/ %
CCE/ %

## Slide 123
Charge sharing
Square
Hexagonal
Ch. 1
Ch. 2
Ch. 2
r = Q2/(Q1+Q2)
Induced charge on Ch. 2
Charge Sharing
Induced charge on Ch. 1
Charge sharing
Charge sharing
Charge sharing
Charge sharing

## Slide 124
Charge sharing comparison
Ch. 1
Ch. 2
Ch. 2
r width/ μm

## Slide 125
Combined Ch. 1 and Ch. 2 at +20V
Square
Hexagonal
Combine
CCE/ %
CCE/ %

## Slide 126
Signal efficiency comparison
Combine

## Slide 127: Summary IBIC
Measured the Signal Efficiency versus position.
Both geometries show > 98% CCE above 20V.
There is a measured drop in signal due to columns of diameter ~ 6μm.
A small region of charge sharing, ~ 20μm, is observed in both geometries.

## Slide 128
Time Resolved IBIC
(TRIBIC)
i = Ewqν
Instantaneous current
Weighting field
Carrier charge
Carrier velocity

## Slide 129
Motivation
+40V
100μm depth
100µm
+20V
Simulation

## Slide 130
Motivation
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 131
Motivation
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 132
Motivation
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 133
Motivation
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 134
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 135
Motivation
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 136
Motivation
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 137
Motivation
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 138
Motivation
+40V
100μm depth
(25μm,25μm)
100µm
+20V
4.5MeV Proton
Simulation

## Slide 139
TRIBIC - setup
Current
amplifiers
Two read-out channels and amplification chains.
Instantaneous
current
pulses

## Slide 140
TRIBIC signals
Plot to be made
Instantaneous current pulse
Integral
Full Width Half Maximum
FWHM.
Signal/ V
Integral/ a.u.

## Slide 141
Position dependence – near electrode
Square
Hexagonal
Simulated electric field at 20V
x (0μm, 15μm)
x (0μm, 15μm)
Measured current pulse at 20V
Signal/ V
4x104
2x102
E/ V/m
4x104
2x102
E/ V/m
Integral/ a.u.

## Slide 142
Position dependence – far from electrode
Square
Hexagonal
Simulated electric field at 20V
x (-15μm, 50μm)
Measured current pulse at 20V
4x104
2x102
4x104
2x102
x (-15μm, 50μm)
Signal/ V
Integral/ a.u.
E/ V/m
E/ V/m

## Slide 143
Position dependence - cell
Square
Hexagonal
Simulated electric field at 20V
Measured FWHM of pulse at 20V
4x104
2x102
4x104
2x102
E/ V/m
E/ V/m
FWHM/ ns
FWHM/ ns

## Slide 144: TRIBIC
Ub=-20V

## Slide 145: TRIBIC: Results
Cubic
Hexagonal
Comparison with TCAD Simulation model:
basic features qualitatively reproduced.
Reasonable agreement, but simplified model.

## Slide 146: Summary TRIBIC
First study of hexagonal and cubic 3D diamond cells with TRIBIC.
  Hexagonal and cubic cells show expected characteristics.
  Qualitative agreement with simulations.
Results shown are preliminary, analysis ongoing.
Next steps:
  Quantitatively verify electrical field distribution.
  Disentangle hole and electron contributions.
  Radiation hardness studies.

## Slide 147: BACKUP

## Slide 148: 3D Diamond detector for medical dosimetry

## Slide 149: Dosimetry application
Planning of dose distribution delivered dose distribution challenging with narrow field beams.
Need high spatial resolution of tissue equivalent dose deposited.
Target numbers:
  Dose uncertainty <1%
  Spatial resolution ~0.1mm

## Slide 150: Dosimetry application
Diamond key properties for dosimetry
  Tissue equivalence
  Radiation hardness
  Room temperature operation
  bio-compatibility
3D Advantage
  Flexible active volume
  Radiation hardness
  Potential for 3D position information
IEEE TRANSACTIONS ON NUCLEAR SCIENCE, VOL. 61, NO. 4, AUGUST 2014
100 mm
100 mm

## Slide 151: Dosimetry application
The Christie Hospital, Manchester
  medical linear accelerator (Elekta Synergy Sband)
  6MV and 10MV accelaration.
  10x10cm radiation field.
  Dose rate dependence.
  Photon beam profile.
The Christie
Manchester
Florence
Perugia

## Slide 152: Dosimetry application
Test set-up:
Asymmetric leakage current.
Ileak < 1nA for -100V to +60V
The Christie
Manchester
Florence
Perugia

## Slide 153: Dosimetry application
First 3D diamond results:
Pre-irradiated with 5 Gy.
Clear response to presence to 6MV photons.
Return to baseline <1s, no significant baseline shift.
Plateau stability needs further studies.
On-off response to 6MV photon beam with 4Gy/min. Variation of bias voltage.
The Christie
Manchester
Florence
Perugia

## Slide 154: Dosimetry application
First 3D diamond results:
Good linearity of ~1% over dose rate range of 2-4 Gy/min.
Dose rate linearity for 6MV photons at -80V
The Christie
Manchester
Florence
Perugia

## Slide 155: Dosimetry application
First 3D diamond results:
Good linearity of ~1% over dose rate range of 2-4 Gy/min.
Good linearity of ~2% over dose range of 0.5 to 7 Gy.
Dose linearity for 6MV photons at -80V
The Christie
Manchester
Florence
Perugia

## Slide 156: Dosimetry application
First 3D diamond results:
Good linearity of ~1% over dose rate range of 2-4 Gy/min.
Good linearity of ~2% over dose range of 0.5 to 7 Gy.
Beam width well reproduced to 1% when compared to GafChromic film measurement.
10cm beam profile measured with 3D diamond at -80V, 4Gy/min and film.
The Christie
Manchester
Florence
Perugia

## Slide 157: Next generation:
Next generation tests with variable array sizes.
The Christie
Manchester
Florence
Perugia