## Slide 1
TRACK
Sen Zhao,On behalf of tracker team and Suyu Xiao

## Slide 2
What For
System level
Signal from detectors
Readout-board
DAQ
Readout board:
BMOS v1
Trigger for BMOS
4 channels readout board
Tools:
LTSPICE: theory simulation
Function level simulation
HFSS: S parameters extraction
Transmission Line Theory and Impedance Matching
KICAD/JIALICHUANG: Layout design
DAQ:
BMOS FPGA
Tools:
VIVADO

## Slide 3
Detector: signal of detector
If decay time of signal << constant time of readout circuit
For extremely short time ,during time from t to t+ ,detector collect charge of Q, the current pulse of signal can be described by:
In simulation, a pulse function can be used to approach the signal from unit function:
With consideration of statistical fluctuations

## Slide 4
Readout-board : noise
The noise level of a system cannot​ be judged from the average but from the RMS , for the average of the noise is always 0.
From a typical detector system:
ENV,ENC, ENN and ENE can also reveal the noise level in a circuit including the amplifier.

## Slide 5
Readout-board : noise
Shot noise:
  Origin from the leakage current of detector
  For shot noise, a higher leakage current or​ a wider amplifier bandwidth results in​ higher noise.
    Shot noise from detector
    Shot noise from amplifier
Thermal noise:
  From the thermal movement of carriers
  For thermal noise, higher temperature results in higher noise.(also depends on the material due to different R)
Low-frequency noise
  For Low-frequency noise, higher the smallest limit frequency results in lower noise.

## Slide 6
Readout-board : Filter and shaper circuit
For:
Reduce the noise in system
Reshape the signal from the detector
CR-RC CIRCUIT
Pole-Zero Cancellation Circuit​

## Slide 7
DAQ: FPGA
PLD(Programmable Logic Device)
PROM: Programmable Read-Only Memory
PLA : Programmable Logic Array
PAL : Programmable Array Logic
GAL : Generic Array Logic
CPLD : Complex PLD
FPGA : Field Programmable Gate Array
De Morgan’s Theorems
&
Karnaugh Map
"Any logic function can be simplified into a Sum of Products (SOP) logic equation.”

## Slide 8
DAQ: FPGA Spartan II
CLB module
CLB primarily includes 2 LUT-FF pairs with carry logic and 2 global tri-state buffers (BUFT). A single LUT-FF consists of 2 4-input LUTs, 2 control and carry logic units, as well as 2 storage elements.

## Slide 9
DAQ: FPGA Spartan II
The LUT (Look-Up Table) is the primary element for logic implementation, and each LUT can implement any 4-input function.

## Slide 10
DAQ: FPGA Design
Demand analysis
Sub-module design
Pre-simulation
System simulation
Logic Synthesis
Layout
Simulation
Test

## Slide 11
Begin a project BMOS
The energy of the linear accelerator has been increased to 300 MeV.
The beam power has been increased to 500 kW.
The neutron flux has reached 1×108 n/cm2/s
A total of 11 new neutron spectrometers and experimental end stations have been added and constructed.
  A high-energy proton experimental station (HPES) capable of providing a 1.6 GeV test beam.
  A surface muon source (MELODY)
Superconducting Linear Accelerator -300 MeV
Rapid-Cycling Synchrotron -500 kW
High-Energy Proton
Experimental Hall
11 Neutron Spectrometers and Experimental End Stations
Phase II Spallation Project
China Spallation Neutron Source is located in Dongguan City, Guangdong Province, China. It mainly provides beams such as protons and neutrons.

## Slide 12
Begin a project
Phase II Spallation Project
High-Energy Proton Experimental Station (HPES)
According to testing requirements, both low-intensity and high-intensity proton test beams are provided.
Parameter | Design Targets
Proton Energy Range | 0.8-1.6 GeV
Beam Pulse Frequency | 24 Hz
High-Intensity Beam Flux | 106-109 p/s
Low-Intensity Beam Flux | 103-104 p/s
Proton ENERGY

## Slide 13
Begin a project
Phase II Spallation Project
Due to the beam characteristics of CSNS, the proton beam exhibits a two-level bunch structure in time ：
  Macro Pulse Frequency: 25 Hz (24 Hz effective)
  Macro Pulse Length: 1 ms, Interval: 39 ms
  Micro-pulse in macro-pulse
  Micro-pulse period: 410 ns，Micro-pulse Length : 80 ns
…
…
410ns
Macro Pulse Period: 40 ms
1 ms
Time
Macro Pulse Interval: 39 ms
80ns
410ns
Schematic Diagram of Micro Pulse Time Structure

## Slide 14
Begin a project
Phase II Spallation Project
Design target
Real-time monitoring of beam intensity, position,
and uniformity
Design life：5000 hours/year
Detection Strategy: Indirect Measurement Method
  Place the detector at the adjustable edge of the beam center to avoid the high-dose irradiation zone, thereby ensuring detector lifespan and beam stability.
  1. Effective Beam Working Area:​​ A central 2cm × 2cm square.
  2. ​Detector Layout:​​ Located in the edge non-direct irradiation zone.
BMOS ---- Beam Monitor Online System

## Slide 15
Begin a project
Phase II Spallation Project
Experimental Verification
Test Platform：Based on IHEP SiC-PIN detector + UCSC board, for testing responses to different particles.
Beta particles:​​ 90Sr radiation source (approaching Minimum Ionizing Particles, MIPs, for validating low-energy signal response).
Test Results
signal peak ~ 20mV
amplifier gain ~ 18
Board Rf ~ about 400 Ohm
SiC Peak of Intrinsic Current Signal~2 uA

## Slide 16
Begin a project
Phase II Spallation Project
Signal Response per Micro-pulse:​
Pulse information for 1–16 protons is generated using random numbers.
Each proton excites a peak current of 2 μA (intrinsic signal for 1.6 GeV proton incidence).
Macro-pulse Intrinsic Current Signal
1ms intrinsic current signal
…
…
410ns
Macro Pulse Period: 40 ms
1 ms
Time
Macro Pulse Interval: 39 ms
80ns
410ns
Micro-pulse Intrinsic Current Signal
(ignoring the 80ns micro structure)

## Slide 17
Begin a project
Phase II Spallation Project
Key configure of ADA4817
Input resistance ~ 500 GOhm
Open-Loop Gain ~ 65 dB
The BMOS PCB design employs a charge integration circuit and a first-order CRRC shaper circuit.
Core Circuit Parameter Design：
Integration capacitance: 10 pF
Discharge time: 15 ms (considering parasitic capacitance of 2~5 pF)

## Slide 18
Begin a project
Phase II Spallation Project
Signal of output
  More than 600 mV
  4 ms zeroing time
Considering the ADC and FPGA , the total system can be reset within 40 ms

## Slide 19
Begin a project
Phase II Spallation Project
High-voltage traces should be kept away from signal traces.
Power supply traces should be thicker.
No ground plane is provided for high-bandwidth chips.
Front View
Rear View
HV pad
Signal bonding pad
Signal grond bonding pad

## Slide 20
Begin a project
Phase II Spallation Project
CAL_IN test
Inject 42pC of Charge
Cal_in:10nF
Cal_in:4mV

## Slide 21
Begin a project

## Slide 22
Begin a project
ADC 16bit
FIFO 16-64
FPGA TOP
FIFO 64-16
Ethernet
Analyzer
DDR3
16bit
64bit
64bit
16bit
Ethernet Header
IP header
UDP header
Config
ADC Data
FCS
14byte
8byte
20byte
1024byte
5byte
4byte
Send data (UDP Data)
Receive data(Ethernet Frame)
ADC Data
1024byte
512 Sample Count
Command Code
FPGA MAC
Channal Sel
1byte
6byte
4byte
4byte
Header
Sample Num
4byte
Eg :00 00 01 00 02 00 0A 35 01 FE C0 00 00 00 01 00 00 40 00
Sample Num
4byte
00 00 40 00
16384 Sample count
32 Frame Count
Send
Command
Send command
Receive data
Receive data
(UDP Data)
2*32bit

## Slide 23
Begin a project
Phase II Spallation Project

## Slide 24
Begin a project
Phase II Spallation Project
Tune = 0： peak：34.04mV
Tune = 30： peak：23.60mV

## Slide 25
Begin a project
BMOS system
Reach :
Stable long time monitor
Linear relation with e/h pairs

## Slide 26
Begin a project
Core
ADC
HV
LV
使能
监控
模拟信号
Adc数据进入core进行处理：电荷，展宽，峰值
GPIO+升压
低压控制
高压控制
温度，电流，
工作状态

## Slide 27
Begin a project
Track upon the trace.
Welcome to join us!