## Slide 1: Proton Computed Tomography with Particle Physics Tracking Detectors Phil Allport Fast timing silicon pixel detectors for new applications: proton therapy 22/11/19 University of Manchester
Development of Silicon Detector Arrays in Particle Physics
  Strip Detectors
  Hybrid Pixel Detectors
  Charge Transfer Detectors
  Monolithic Active Pixel Sensors
  Silicon Fast Timing Detectors
Applications in Hadron Radiotherapy
  Dosimetry
  Motivation for Proton CT
  PRaVDA Results
  Selected International Activities
Conclusions
Dose range
X-ray radiotherapy
Proton radiotherapy

## Slide 2
Silicon Sensor Technology Developments .
For charged particle tracking in particle physics, a large number of different solid-state detector technologies have been developed over the last four decades.
Key requirements have included efficiency (signal/noise), resolution (electrode size and pitch), speed/rate (microelectronics), large area (wafer size, yield and cost), low multiple scattering (minimal total material including cooling and services) and radiation hardness.
Gaseous and scintillating fibre technology also widely used but not discussed here due to rate and radiation hardness limitations.
Applications of silicon strip and pixel-based particle tracking detectors - Nature Reviews Physics - https://doi.org/10.1038/s42254-019-0081-z - Allport2019ER

## Slide 3
All these figures showed doubling times of < 2 years up to now. Some scalings stopping, but other improvements conceivable. Can still hope for major detector improvements and enhanced data aquisiiton plus computing capabilities. However, storage and CPU costs may not scale as fast as might be desired for future projects.
Data Storage
Transistors/mm2
Bandwidth
ADC pJ/conversion
Commercial Microelectronics Evolution

## Slide 4
104m2 = 1 hectare
Historical Development of Silicon Sensor Arrays .

## Slide 5
10
MARK II (SLAC)
128 Channel
Microplex
Chip (NMOS)
DELPHI (LEP) MX6 (CMOS)
Application Specific Integrated Circuit
(ASIC) development has been essential
Strip Readout ASICs
High Density Wire-bonds
Hybrid Pixel
Detectors
Monolithic Pixel Detectors
Historical Development of Silicon Sensor Arrays .

## Slide 6
10
The highest channel count arrays are based on pixelated detectors.
For hybrid pixel sensors connection to the electronics requires flip-chip technologies.
True space-points
Hybrid Pixel Readout ASICs
Two orthogonal projections
Required in high track density environments to avoid ambiguities
Hybrid Pixel
Detectors
Historical Development of Silicon Sensor Arrays .

## Slide 7
1.2×1015neq/cm2
1.9×1016neq/cm2
n+ strips
in p- bulk
Federico Faccio: PMOS turn-on Vg https://indico.cern.ch/event/468486/
1.3×107Gy
For Pixels: also need rad-hard data links
with up to ~40 hits/cm2 each 25ns
Strips
Pixels
niel
Radiation Effects: Sensors and Electronics
5×105Gy
Fluence (1014neq/cm2)
Hybrid silicon detectors (pixels/strips) signal output drops with irradiation to very high doses
RD50
ATLAS
CMS
LHCb
Bulk damage (measured in units of 1MeV equivalent neutrons/cm2 (assuming scaling with non-ionising energy loss) drives the deterioration of sensors.
For microelectronics worry about total ionising dose.
(65nm CMOS - RD53) can start to see significant deterioration above 500Mrad (5MGy) ◊Many different effects◊
Example of radiation expected at High Luminosity
LHC (HL-LHC)

## Slide 8
CCD, Drift and DEPFET
A number of technologies have been adopted and could be suitable for future facilities where the charge is moved through the silicon by clocking potentials (CCD), lateral field (Drift) or stored in the silicon beneath a transistor allowing non-destructive read-out (DEPFET)
Arrival time gives lateral drift distance
Charge stored beneath gate can be read out as giving increased Id for given Vg
BELLE II Vertex Detector at SuperKEKB
SLD Vertex
CCD
Drift
DEPFET

## Slide 9
Monolithic: SoI and 3D Integration
  Wafer bonding, through silicon via (TSV) and back-thinning technologies allow circuits to be stacked in 3D.
Could imagine detector concept with stacked opto-electronics digital electronics, analogue electronics and sensor substrate
Monolithic pixel sensors combine the functionality of the microelectronics with the sensor substrate as part of a single silicon object.
CCDs and DEPFETs already incorporate active circuit elements in the sensors which are therefore no longer simple passive devices.
However, these devices are potentially slow and prone to radiation damage.
Silicon on Insulator technology allows integration of full read-out circuitry above a buried oxide over the sensing substrate.

## Slide 10
Commercial CMOS Image Sensors offer possible dramatic decrease in costs (Monolithic Active Pixel Sensors)
MAPS can deliver very low power consumption at low R/O speeds, possibly <100mW/cm2 i.e. simple water cooling
  Ultra low material budget (cf ALICE ITS upgrade: <0.5% for inner layers, <1% for outer layers)
  But these devices limited in speed and radiation hardness
  Current and near future MAPS for heavy ion experiments
    integration time up to 4μs (noise, electron diffusion)
    radiation resistance up to few 1013 neqcm-2
Major developments in HV/HR-CMOS → deep depletion region with charge collection by drift not diffusion → huge improvements in collection speed and radiation hardness
Can usually either have small collecting node (and therefore faster and low noise) but shallow charge collection or deplete from the deep n-well with larger signal produced in up to 100μm of silicon but higher capacitance (= noise+slower)
MAPS: HV/HR-CMOS Detectors
Modified process with junction in substrate
If DMAPS could approach hybrid pixel radiation hardness they should offer much less material and cost

## Slide 11
RD50, ATLAS, CMS: Low Gain Avalanche Detectors (LGAD) (Radiation issues still under study for HL-LHC applications)
Corner of 15×15 array
2/3 layers of LGAD sensors
Ultra Fast Timing Detectors
11
Bump-bonded hybrid 6.4m2 of 1.3mm×1.3mm
pads proposed in ATLAS

## Slide 12: Hadron Radiotherapy
Proton and Ion therapy
https://www.ptcog.ch/index.php/
Lists 94 hadron radiotherapy facilities “in operation”, 45 “in construction” and 21 “in planning stage”
Tumour
Proton vs photon
Medulloblastoma
Paediatric CNS Tumour
www.pravda.uk.com

## Slide 13: Hadron Radiotherapy
Proton Computed Tomography (pCT) hoped to be the solution
Tumour
Impact of under/over estimating stopping power
https://www.ptcog.ch/index.php/
Lists 94 facilities internationally “in operation”, 45 “in construction” and 21 “in planning stage”
www.pravda.uk.com

## Slide 14
Need to know momentum vector and entry point of each proton going into the patient and also momentum vector and exit point coming out.
Need tracks plus incoming and outgoing energies
Proton Radiotherapy Verification and Dosimetry Applications
https://physicsworld.com/a/proton-imaging-comes-to-the-christie/
14
Uses HL-LHC strip tracking detectors
Except at the highest energies multiple scattering dominates, so thin is better than fine pitch
www.pravda.uk.com

## Slide 15: Computed Tomography
Most common imaging modality for treatment planning is Computed Tomography (CT)
Measure the linear attenuation of x-rays (μ) along a path through patient (line integral) to form a radiograph
Rotate the source and detectors around the patient and take another radiograph from different angles
Use a reconstruction algorithm to combine all of the line integrals as a function of position in the patient
Information in CT image stored as the Houndsfield unit which is the linear attenuation relative to water and air
Lung = -850 to -910
Bone = 700 to 3000
Water = 0
Currently x-ray CT is performed, which is then converted into stopping powers
“The values recommended in this study based on typical treatment sites and a small group of patients roughly agree with the commonly referenced value (3.5%) used for margin design.”
Main range uncertainty caused by imaging the patient with x-rays, but treating with protons
M Yang et al PMB. 57
4095–4115 (2012)
Uncertainties in RSP Estimation (1σ)

## Slide 16: Proton Computed Tomography
x-u-v: A disadvantage of other systems with crossed strips is that they cannot cope with two or more protons per frame without ambiguities.
The usual way round this with strips is to have a third layer
Depending on strip pitch, still problems above several hits / frame
Then need truly pixelated sensor, but needs to be fast
(3N Channels)
(N2 Channels)
www.pravda.uk.com

## Slide 17: Proton Computed Tomography
Proton energy after passing through the patient found from the distance penetrated in the range telescope
Need to know momentum vector and entry point of each proton going into the patient and also momentum vector and exit point coming out.
Need tracks plus incoming and outgoing energies
https://gow.epsrc.ukri.org/NGBOViewGrant.aspx?GrantRef=EP/R023220/1
www.pravda.uk.com

## Slide 18: Proton Computed Tomography
21 silicon planes interspersed with absorber to measure proton stopping distance
In final realisation, range telescope and tracker used common (n-in-p) strip technology
4 sets of 3 silicon strip planes reading out at the duty cycle of the cyclotron
www.pravda.uk.com

## Slide 19
Installed at iThemba for first tests May 2016 (125 MeV degraded beam).
Compensator in place.
180 rotations at 1o steps.
Over 1M protons / rotation.
Second run November 2016.
Proton Radiotherapy Verification and Dosimetry Applications
www.pravda.uk.com

## Slide 20
Proton scattering power CT
Installed at iThemba for first tests May 2016 (125 MeV degraded beam).
Compensator in place.
180 rotations at 1o steps.
Over 1M protons / rotation.
Second run November 2016.
280 M proton histories recorded.
Proton Radiotherapy Verification and Dosimetry Applications
www.pravda.uk.com

## Slide 21
Raw data rate ~ 360 Gbit/s
Installed at iThemba for first tests May 2016 (125 MeV degraded beam).
Compensator in place.
180 rotations at 1o steps.
Over 1M protons / rotation.
Second run November 2016.
280 M proton histories recorded.
Reconstructed using novel Back-Projection-Filtering algorithm specifically developed for proton CT.
Proton Radiotherapy Verification and Dosimetry Applications
www.pravda.uk.com

## Slide 22
High contrast inserts (LN10, Air, SB5)
Low contrast inserts (AP7, WT1, RB2)
Stopping Power
Installed at iThemba for first tests May 2016 (125 MeV degraded beam).
Compensator in place.
180 rotations at 1o steps.
Over 1M protons / rotation.
Second run November 2016.
280 M proton histories recorded.
Reconstructed using novel Back-Projection-Filtering algorithm specifically developed for proton CT.
Stopping power for all inserts (except lung) agree within 1.6% of expected values.
Proton Radiotherapy Verification and Dosimetry Applications
www.pravda.uk.com

## Slide 23
Proton CT
Cone beam CT
Lamb chop chosen as first test of proton CT on real tissue due to regions of bone, soft tissue and fat
Same parameters as for imaging phantom, but with 2o rotation steps
Installed at iThemba for first tests May 2016 (125 MeV degraded beam).
Compensator in place.
180 rotations at 1o steps.
Over 1M protons / rotation.
Second run November 2016.
280 M proton histories recorded.
Reconstructed using novel Back-Projection-Filtering algorithm specifically developed for proton CT.
Stopping power for all inserts (except lung) agree within 1.6% of expected values.
Still need optimal way to combine scattering and stopping power.
Proton Radiotherapy Verification and Dosimetry Applications
www.pravda.uk.com

## Slide 24
The pCT Collaboration
R. P. Johnson, Tia Plautz, Hartmut F.-W. Sadrozinski, A. Zatserklyaniy: SCIPP, U.C. Santa Cruz, Santa Cruz, CA , USA
V. Bashkirov, V. Giacometti, F. Hurley, P. Piersimoni, R. Schulte: Division of Radiation Research, Loma Linda University, CA, USA
P. Karbasi, K. Schubert, B.Schultze: Baylor University, Waco, TX, USA
Loma Linde with SCIPP (UCSC) see for example Robert P Johnson et al 2017 “Review of medical radiography and tomography with proton beams” https://doi.org/10.1088/1361-6633/aa8b1d
H.F.-W. Sadrozinski, et al, Development of a Head Scanner for Proton CT, Nucl. Instr. Meth. A 699 (2013) 205.
protons
5-Stage Energy Detector
Trackers
Rotation Stage
Four 9cm*9cm silicon strip sensors with thin-edge technology in x-y configuration
1 million individual protons reconstructed per second

## Slide 25
Proton
X-ray
3D rendering of the pCT-reconstructed RSP map of a pediatric anthropomorphic head phantom
The relative stopping power
(RSP) of the variety of materials in the phantom is accurately measured
Three cardinal planes of 3D RSP images obtained with Phase II scans of the anthropomorphic head phantom.
The pCT Collaboration

## Slide 26: Some Other Silicon Based Trackers
Item | Specifications
Outer dimensions | 89.5×89.5 mm2
Active area | 87.6×87.6 mm2
Substrate thickness | 410±10μm
Strip pitch | 228μm
Readout pitch | 456μm
Number of strips | 384
Full depletion voltage | <120 V
  Another group using similar sensors to the pCT Collaboration is based at Niigata University
  A very promising approach being followed by Padua with CERN is the use of MAPS technologies for the tracking but with intelligent on-chip data sparsification to deliver faster, low power readout
  However, the approach does not yet address the radiation tolerance challenges. Although these are being addressed for particle physics applications, these are yet to be demonstrated in a wafer-scale sensor, which is ideally what is needed for pCT.

## Slide 27: CMOS Imaging Sensors and BGO
  MAPS (Mimosa) are also being looked at by groups in the Netherlands and Norway
  There are a number of other initiatives using different tracking/calorimeter technologies, scintillating fibres, multi-wire proportional chambers, … all of which have potential advantages and disadvantages when compared with silicon detectors
  Also Micro-Pattern Gas Detectors (CERN, Vienna, ...)
Sodai Tanaka et al 2018 Phys. Med. Biol. in press https://doi.org/10.1088/1361-6560/aaa515
•••• CCD Camera
Chicken leg

## Slide 28
Fermilab, Delhi and Northern Illinois system uses scintillating fibre tracker
Area 20x24 cm² (4 planes upstream) and 24x30 cm² (4 planes downstream) with 15 cm separation between planes.
96 plastic scintillating tiles with dimensions of 27cm width, 36 cm height, and 3.2mm thickness were used in the range stack.
At the proton beam at the Central DuPage Hospital in Warrenville, IL
FERMILAB-TM-2617-AD-CD-E
Scintillating Fibre Tracking

## Slide 29
Limitation in Adaption for Clinical Use
The main issue for many systems is the size of the system and the problem with coping with multiple protons per cyclotron/synchrotron spill both in the tracker and in the calorimeter or range telescope
Spot scanning makes the issue with ambiguities more difficult. Also hard to get very low currents in clinical setting. → Are there proton time differences within each spill?
Example calculation: 10pA would imply 6.25109 protons per second.
cf pCT 106 protons reconstructed per second
If 72* MHz rep rate, then 0.87 protons per spill. Spot size ~cm2 anticipated to be continuously scanned across the field of view with 180 different projections needed.
Depending on volume to be imaged, could need 109-1010 protons, so 107-108 per angle.
Might need ~103 spot locations so ~(104-105) proton spills per location ie on average ≳ µs per 1cm2 location ≳ ms per angle
Another thought for fast timing is energy measurement where 10ps could give 1MeV energy resolution with 1m ToF up to ~200MeV
Chart: 
| 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300
1m ToF (ps) | 23009.9799 | 16398.7253 | 13493.7303 | 11775.7846 | 10612.6183 | 9760.7005 | 9103.7189 | 8578.2185 | 8146.2709 | 7783.638 | 7474.0167 | 7205.9878 | 6971.2853 | 6763.7606 | 6578.7335 | 6412.5713 | 6262.4061 | 6125.94 | 6001.3076 | 5886.9769 | 5781.6771 | 5684.3436 | 5594.0776 | 5510.114 | 5431.7973 | 5358.5621 | 5289.918 | 5225.4373 | 5164.745 | 5107.5111
d ToF (ps) / d KE = 1MeV | 661.1255 | 290.4995 | 171.7946 | 116.3166 | 85.1918 | 65.6982 | 52.55 | 43.1948 | 36.2633 | 30.9621 | 26.8029 | 23.4702 | 20.7525 | 18.5027 | 16.6162 | 15.0165 | 13.6466 | 12.4632 | 11.4331 | 10.53 | 9.7333 | 9.0266 | 8.3964 | 7.8317 | 7.3235 | 6.8644 | 6.4481 | 6.0692 | 5.7234
*Varian ProBeam®
https://www.varian.com/oncology/solutions/proton-therapy

## Slide 30
Radiotherapy Detector Systems
Reconstruction of a 1.5 x 1.5 cm2 square 6 MV beam, centred at (0.5 cm, 0.5 cm).
Collaboration of Birmingham, Queen Elizabeth Hospital, STFC TD, VivaMOS Ltd and NPL towards developing bulk-damage radiation-hard sensors to enhance the UK primary standard for absorbed dose for proton radiotherapy by incorporating real-time spot profile measuring apparatus with NPL calorimetry.
Currently using strips (from PRaVDA) in 3 projections and CMOS pixels
LASSENA is radiation-hard against ionising dose but not protons. Need to take advantage of latest HL-LHC CMOS MAPS developments to have truly 2D imaging capability.
NPL Calorimeter at Clatterbridge

## Slide 31
Also:
STFC Global Challenge Network+ in Advanced Radiotherapy*: many projects, including work combining UCL Calorimetry with PRaVDA tracker, DMAPS with Bonn, developments towards proton Micro-beam studies and several dedicated studies using Birmingham MC40 cyclotron.
ENLIGHT: The European Network for Light ion Hadron Therapy
FLASH (2-40 Gy in few 100ms) and Micro-beam (sub mm segmentation) currently very exciting as both show much better sparing of healthy tissue for a given dose although mechanisms still not fully understood.
STFC TD, NPL, VivaMOS Ltd and Birmingham demonstration of imaging of micro-beam slit system at clinical X-Ray source with large area CMOS sensor (225 kV, 1.65 Gy/min, Helmhotz Research Centre in Munich).
Resolution at least as good as film (in fact better contrast) and gives real time image.
Radiotherapy Detector Systems
*led by Karen Kirkby

## Slide 32: Conclusions
  Many individual ingredients for “4D” detectors now exist in particle physics (high sensor segmentation, sensors with intrinsic fast response, 40MHz data rates, 50µm×50µm pixel electronics, radiation hardness, …. ) which could give interesting devices for different aspects of hadron radiotherapy applications
  Many groups internationally looking at proton CT (only some discussed here) but a spot scanning clinical system will probably give a time structure leading to multiple nearby protons per frame (with frame rates in the many tens of MHz) that need resolving
  A critical requirement is to both use protons efficiently to minimise dose during imaging while taking the image in a time consistent with a sensible patient treatment plan and gantry/spot scanning hardware capabilities
  OPTIma, in the UK, is building on PRaVDA successes to address many of these requirements but is likely for the tracking aspects to use the mature silicon strip technology to meet the very high data rates, area cost and radiation requirements
  A key observation is that multiple scattering dominates so, for the tracker, material must be minimal so monolithic thin devices would be idea if able to cope with the frame rate
  Pitches can be large (or pixels can be ganged together) to reduce channel count
  Ambiguities for nearby protons might be resolvable with timing information (tbc)
  Time of flight with 10ps resolution could give useful energy resolution capability