## Slide 1
Investigations of Background and Compton Suppression Shields for GRIFFIN
Nikita Bernier, UBC and TRIUMF
for the GRIFFIN collaboration
2014 Canadian Association of Physicists Congress
June 17th, 2014.

## Slide 2: GRIFFIN Spectrometer
GRIFFIN : Gamma-Ray Infrastructure For Fundamental Investigations of Nuclei
  16 large-volume clover-type High Purity Ge [HPGe] detectors dedicated to decay spectroscopy research with the low-energy radioactive ion beams in ISAC-I at TRIUMF
  Five sub-systems are combined to create a high-efficiency decay spectrometer for sensitive measurements.
2
N. Bernier
See A.B. Garnsworthy (M2-2)
In-vacuum moving tape collector system

## Slide 3: GRIFFIN Spectrometer
GRIFFIN : Gamma-Ray Infrastructure For Fundamental Investigations of Nuclei
  16 large-volume clover-type High Purity Ge [HPGe] detectors dedicated to decay spectroscopy research with the low-energy radioactive ion beams in ISAC-I at TRIUMF
  Five sub-systems are combined to create a high-efficiency decay spectrometer for sensitive measurements.
2
N. Bernier
See A.B. Garnsworthy (M2-2)
In-vacuum moving tape collector system

## Slide 4: Suppression Shields
The performance of a γ- ray spectrometer is determined primarily by
  γ- ray energy resolution
  total γ- ray photopeak detection efficiency
  photopeak-to-total ratio, and
  suppression of environmental background signals.
Solutions for improved performance :
  Significantly reduce Compton scattering and escape peaks [high rate] by placing HPGe detectors in close proximity and add energy loss in each crystal
  Shield against the radioactive background [low rate] present in the experimental hall by surrounding HPGe with a high density scintillator.
3
N. Bernier

## Slide 5: Addback and Escape Suppression
If energy deposited in two crystals, treat as a single incident γ-ray : add energies.
If energy measured in a scintillator, veto any HPGe event in coincidence.

G. Hackman
High probability that an incoming γ- ray Compton scatters, or an annihilation photon, will escape the detector
  Results in a continuous spectrum of lost energy and escape peaks.
4
N. Bernier

## Slide 6
Experimental Setup
GRIFFIN clover set up with TIGRESS shields
20 optically-isolated scintillators per shield for crystal specific Compton suppression
  Retractable BGO front shields, BGO side shields and CsI back shields.
5
N. Bernier
TIGRESS array in ISAC-II experimental hall

## Slide 7: Background Characterization
Measurements taken with GRIFFIN clover in 2 locations in ISAC-I hall
  East : closer to sources safe
  West : closer to yield station.
6
N. Bernier
Yield Station
Target Stations

## Slide 8: Different Locations
Proton beam off, 48h
East : >124 Hz/crystal, West : 102 Hz/crystal
>18% difference between both locations.
! Different energy thresholds
Blue=59 keV Red=30 keV
238U and 232Th decay chains
40K in concrete
Calibration sources
7
N. Bernier

## Slide 9: Environmental Background
40K and U/Th series decays and cosmic rays
Activity generated by high-energy neutrons produced when the 500 MeV proton beam impinges on the high-power ISAC production targets located 2 stories below the ISAC-I experimental hall.
8
N. Bernier
M. Brown

## Slide 10: Different Orientations
Proton beam off, 48h
Down : 102 Hz/crystal, Horizontal : 87 Hz/crystal
14% decrease with orientation.
214Bi
238U series
40K
Concrete
208Tl
232Th series
214Pb
238U series
228Ac
232Th series
351.9
911.2
1764.5
2614.5
1460.8
9
N. Bernier

## Slide 11: Shielded Clover
Proton beam off, 2h
Passive Shielding : 101 Hz/clover, No Shielding : 321 Hz/clover
69% decrease with passive shielding only.
! 4 HPGe crystals summed without addback.
214Bi
238U series
40K
Concrete
208Tl
232Th series
214Pb
238U series
228Ac
232Th series
351.9
911.2
1764.5
2614.5
1460.8
10
N. Bernier

## Slide 12: ISAC Proton Beam ON
737 Hz/clover, ~12 kHz in the full GRIFFIN array without shielding
Expecting 0.1 mHz for ~0.01 ions/s radioactive beams
Background exceeding signal of interest by factor 108.
11
N. Bernier

## Slide 13: Active Suppression
TIGRESS clover in closed array, 24h
Passive : 50 Hz/clover, local active : 23 Hz/clover
Reduction of GRIFFIN background by a factor 20.
! Background rates will be moderately higher due the closer proximity of GRIFFIN to the ISAC production targets.
214Bi
238U series
40K
Concrete
214Pb
238U series
228Ac
232Th series
351.9
911.2
1764.5
2614.5
1460.8
208Tl
232Th series
12
N. Bernier

## Slide 14: Peak-to-Total Ratios
TIGRESS clover in closed array, 4h
Improved ratios from 9.8% to 11% for a 60Co source.
13
N. Bernier
1173.2
1332.5

## Slide 15: Example : Decay of 54K to 54Ca
C.E. Svensson
Simulation for a one week experiment using GRIFFIN at a beam rate of 0.01 ions/s using TIGRESS background measurements
β-γ coincidence condition already suppresses by a factor 106
Impossible without the suppression shields !
Photopeaks from
40K, Th series and neutron capture on 27Al
1656
2043
14
N. Bernier

## Slide 16: Overview
GRIFFIN array (16 clovers) will count ~12 kHz of background triggers without suppression shields, limiting experiments to isotopes with beam intensities ≥ 0.1 ions/s.
Suppression shields will represent a factor of 20 reduction in environmental background enabling an entire class of sensitive experiments down to beam intensities ~0.01 ions/s, approximately 200 additional exotic isotopes.
15
N. Bernier

## Slide 17
Merci!
Thank you!
P.C. Bender, R. Churchman, S. Ciccone, I. Dillmann, A.B. Garnsworthy, S. Georges, G. Hackman, R. Krücken, D. Miller, W.J. Mills, M.M. Rajabali, C. Unsworth
TRIUMF
E.T. Rand, C.E. Svensson
University of Guelph