## Slide 1: Plastic Scintillators as In-Vivo Dosimeters: Angular Dependence of Response
Ethan Avila Ruiz, MSc1, Thalat Monajemi, PhD1,2
1Department of Physics and Atmospheric Science, Dalhousie University
2Department of Radiation Oncology, Dalhousie University

## Slide 2: Plastic scintillation dosimetry (PSD)
Benefits of PSDs
  Similar density to water
  Less corrections
  High temporal resolution
  Linear response with dose

## Slide 3: Scintillator with spectrometer setup
1 cm scintillator
40’ optical fiber
Exemplar plus spectrometer
Placeholder
Scintillation detector: components
Jacket
SC SF-3HF
SC SF-78
PMMA core of
fiber-optic cable

## Slide 4: Validate spectral invariability of system signal
Well-characterized signal components in arbitrary dose measurements

## Slide 5: Determination of scintillation signal
Scintillation:
Proportional to dose
Acquired by subtracting Cerenkov
Cerenkov:
Highest intensity at 45o
Obtained from stem measurement
Total signal:
-Spectrometer signal

## Slide 6: Spectral variability with angle
Inline
Crossline

## Slide 7: Spectral variability of stem signal
Inline
Crossline

## Slide 8: Relative dose affected by variability
Inline
Crossline

## Slide 9: Effect is reduced with depth
Inline
Crossline

## Slide 10: Summary
Surface dosimetry requires a cautious approach
Subtraction correction does not account for Cerenkov in scintillator
Spectral correction more robust under all measurement conditions