## Slide 1: Improving the Super-Kamiokande PMT modelling using the the photosensor test facility (PTF)
Vincent Gousy-Leblanc, Elena Cornick,Skylar Wingfelder, Matej Pavin, Xiaoyue Li, Ryosuke Akutsu, Thomas Lindner, Patrick de Perio, Akira Konaka
June 8th

## Slide 2: Outline
Context
Why ?
  Motivations for measurements
How ?
  Overview of the set up
What ?
  Overview of the measurements
Results
  Integration of measurements into simulation

## Slide 3: Super-Kamiokande experiment
x
Goals: Detect neutrino mixing parameters, proton decay, search for dark matter
~11,000 photo-multiplier tube (PMT)
Super-Kamiokande
Neutrino events creating cherenkov radiation
Neutrino events
Neutrino interaction

## Slide 4: Motivations for PTF measurements
±1.9%
Stop-μ (multi-GeV)
Stop-μ (sub-GeV)
π0
Decay-e
(MC-Data)/Data (%)
Systematic error for Super-K high energy analysis (~2%)
Can these be partially explained by PMT response mismodelling (eg. angular/magnetic field/polarization/wavelength effects)
  Qualitative example
Will become even more important for next generation neutrino experiment (no longer limited by statistical uncertainties)
  1% required for Hyper-k
    See M.Hartz talk
Control samples
Photocathode hit region
PTEP 2019,053F01
Notes: . Previously discussed on bottom up calibration, let’s see what it is in terms of numbersHigh and low energy are corrected on average, geometrical dependance of the PMT light is left. The average calibration (cosmic muon and linac electrons) goes from top to bottom, only downward going light which opens a potential bias. MENTION Energy scale

## Slide 5: The Photosensor Test facility (PTF) at TRIUMF
3 pairs of Helmholtz coils (one in each direction)
  Can control magnetic field
2 optical boxs (laser, sensors to measure tilt, rotation angle and magnetic field)
DAQ to perform 2D scans of PMT
Angular response and reflection measurements
Water system
Electronic rack
PTF frame
PTF room
PTF measurements
Rotation:φ
Tilt :θ
PTF Geometry

## Slide 6: Potential measurements of PTF
PTF will be able to measure and separate external variables
  Magnetic field
  Angular dependance
  Polarization dependance
  Wavelenght dependance
QE vs wavelenght
G1
Calibration of the polarizer
Magnetic field compensation
Angular scan
Bx [mG]
Calibration of the tilt
Y[m]
X[m]
X[m]
Y[m]
Bx [mG]
Measured [deg]
Requested [deg|
Cumulative intensity
Angle between 2 polarizer
Rotator mount
Wavelength [nm]
Quantum efficiency [%]

## Slide 7: Polarization measurements
Effect of polarization seems to be diagonal
  Expectations: edges have more reflections so should be more affected
Further investigation is required
90˚ polarization
Ratio
0˚ polarization
=
nb of photoelectron pulse
nb of pulse
DE

## Slide 8: Integration of measurements into simulations

## Slide 9: Simulation pipeline
Simplest implementation :
  Replace Uniform DE by position dependant DE
SKG4
Uniform DE implemented
True hits
Analysis (energy reconstruction, vertex position etc)
Digitize hits
X position
Y position
Uniform DE
Super-K Geant4 simulation

## Slide 10: Simulation pipeline
Simplest implementation :
  Replace Uniform DE by position dependant DE
SKG4
PMT measurements done in 2020 (in water, 0mG, no acrylic)
True hits
Analysis (energy reconstruction, vertex position etc)
Digitize hits
X position
Y position
PTF position dependant DE
Super-K Geant4 simulation

## Slide 11: In detail: PMT Modelling
}
Simulation
PTF data
Step 2: Input PTF measurements
Step 1: Optical simulation assuming uniform DE in SKG4
}

## Slide 12: Modelling the experiment
Step 3: Build empirical model
Simulation output, assuming uniform PMT response
Empirical model of DE (corrected for known effects) for input to sim
÷
=
DE measurements
True hits
PTF position dependant DE
Analysis (energy reconstruction, vertex position etc)
PTF implementation in (SKG4)
Digitize hits
Notes: WCSIm, SKG4, SKDETSIM

## Slide 13: Integration of data into simulation
Step 4: Test empirical model ( first cross check seems to be good)
Implemented DE
DE measurements
True hits
PTF position dependant DE
Analysis (energy reconstruction, vertex position etc)
PTF implementation in (SKG4)
Digitize hits
Notes: WCSIm, SKG4, SKDETSIM

## Slide 14: Integration of data into simulations (2)
Step 5: Apply empirical model on simple study : particle gun μ+ ,E=1 GeV, dir= (0,0,-1;-1,0,0) vertical and horizontal, 1500 events
Horizontal
Vertical
-True hits
Get PTF position dependant DE
-Total charge
Uniform QE
MC-MC comparison
Super-K tank
Notes: WCSIm, SKG4, SKDETSIM

## Slide 15: Total charge horizontal vs vertical case
Muon
  Nominal
    Mean ratio : 0.98
  PTF
    Mean ratio: 0.96
Horizontal:
Mean:8694
RMS:171.5
Horizontal:
Mean: 8762
RMS:162.3
Muon (PTF case)
Vertical : Mean:8576
RMS:155
Vertical :
Mean:8315
RMS:161
Mean ratio
Mean Vertical
Mean Horizontal
Difference
Ratio_nominal
2.34%
Ratio_PTF
Second ratio
Muon (Nominal case)

## Slide 16: Total charge horizontal vs vertical case (2)
Decay electron
  Nominal
    Mean ratio :1.03
  PTF
    Mean ratio: 1.02
Angular distribution
Mean:277.4
RMS:104.6
Mean:296.5
RMS:108.6
Mean:288.1
RMS:106.8
Electron (Nominal case)
Mean:282.4
RMS:106.6
Electron (PTF case)
Electron angular distribution
ɸ
Difference
1.09%

## Slide 17: Conclusion
The PTF facility was rebuild and improve to measure the PMT response
  A better understanding of PMT response could decrease/characterize the systematic uncertainties associated to the detector
  More measurements to come soon !
Integration of measurements into simulation was started
  Only MC-MC comparison and for simple case
  An effect of ~2% was seen but more measurements are required

## Slide 18: Thank you

## Slide 19: Back up

## Slide 20: Total charge horizontal vs vertical case
Muon
  Nominal
    Mean ratio : 0.98
  PTF :
    Mean ratio: 0.96
  Difference: 2.34%
Mean:8694
RMS:171.5
Mean: 8762
RMS:162.3
Muon (Nominal case)
Muon (PTF case)
Mean:8676
RMS:155
Mean:8315
RMS:161
±1.9%
Stop-μ (multi-GeV)
Stop-μ (sub-GeV)
π0
Decay-e
(MC-Data)/Data (%)
Control samples

## Slide 21: Motivations for PTF measurements
±1.9%
Stop-μ (multi-GeV)
Stop-μ (sub-GeV)
π0
Decay-e
(MC-Data)/Data (%)
Significant systematic error for Super-K high energy analysis (~2%)
Also for low energy solar neutrino (~0.5%)
Can these be partially explained by PMT response mismodelling (eg. angular/magnetic field/polarization/wavelength effects)
  Part of bottom-up calibration
  Qualitative example
Will become even more important for next generation neutrino experiment (no longer limited by statisticals uncertainties)
Control samples
Photocathode hit region
PTEP 2019,053F01
Phys.Rev.D94,052010
Notes: . Previously discussed on bottom up calibration, let’s see what it is in terms of numbersHigh and low energy are corrected on average, geometrical dependance of the PMT light is left. The average calibration (cosmic muon and linac electrons) goes from top to bottom, only downward going light which opens a potential bias. MENTION Energy scale

## Slide 22: Compensating the magnetic field
Try to make the magnetic field as uniform as possible
  Using G-IRON passive shielding
Degauss procedure is needed
  Metals structure surrounding PTF creates their own magnetic field
  Needs to be simultaneously for all directions
2-Spatial scan :
3-Differential plot
1-Voltage scan
Example of a degauss procedure
Voltage[V]
Step
X[m]
Y[m]
Differential [mG]
Coil 3 voltage [V]
Coil 3 voltage [V]
Coil 4 voltage [V]
Bx [mG]
Bx [mG]
Notes: -Fitting function we had the famous +- 100mG
-Planar function

## Slide 23: Compensation of the magnetic field
Compensated to 0 but also different offset
  Input from Super-K magnetic field measurements
  Seems to work within 20mG or so
    Z direction has a larger gradient
Bx [mG]
By [mG]
Bz [mG]
Example setting X=100mG, y=0mG, Z=0mG
Compensation table
X[m]
Y[m]
X[m]
Y[m]
X[m]
Y[m]

## Slide 24: Effect of PMT Angular Response on Reconstruction
Varying angular response can change the total charge
Reconstructed Momentum Bias in HK (N. Prouse, HK calibration)
+10% →-10%
-10%→+10%
Electrons
Variations at MC level
+10% →-10%
-10%→+10%
-10%→+10%, fixed Q
+10%→-10%, fixed Q
Muons
Different trends indicate not fully correlated between e/μ
“fixed Q”: scaled overall efficiency averaged over all events to study case of no change in total charge
  Assumes degeneracies with other detector (e.g. water) parameters are fully constrained
  Can still have ~0.5% bias in reconstructed momentum
Different trend
PMT angular response (efficiency)
-10% → +10% linear scaling
Nominal (from WCSim)
+10% → -10% linear scaling
-10%→+10%, fixed Q
+10%→-10%, fixed Q
-10% → +10% linear scaling
-10% → +10% linear scaling, fixed Q
Nominal (from WCSim)
+10% → -10% linear scaling, fixed Q
+10% → -10% linear scaling
Input to fitqun

## Slide 25: PTF in pictures
PTF frame
Gantry system
20 inch PMT
Optical box
Stepper motors
Water tank
Polarizer
Monitor PMT

## Slide 26: Pipeline is ready !
Some text file are empty, partially finished
  Takes ~50hours for 9 text files on local cluster
Move to something quicker for first comparaison with PTF measurements
Cosmic muon text file
SKG4 (SK4 geometry)
root2zbs
Cosmic muon analysis
MC-MC comparaison
fiTQun (SKDETSIM tuning)
zbs2root
Superscan
Muon reconstructed energy
Distance from nearest wall

## Slide 27: SKG4 local coordinates system
S and P designation are in terms of the plane of incidence of the photon (parallel vs perpendicular)
Done for incoming photons
Calculate incidence plane angle and randomize s or p polarization
At the center no s or p polarization

## Slide 28: Absorption coefficient
Initial conditions
  X polarization
See effect of p vs s polarization
See the maximum predicted
S polarization
P polarization

## Slide 29: Potential study
Angular scan (60˚) + magnetic field
Angular response move region of high efficiency
  Same position but deformed ?
Highest possible resolution (0.5cm)
  Clearly see the effect of the temperature variations
In terms of gantry position
  Off center scan
+100mG in Y direction
+100mG in X direction
+100mG in Y direction
+100mG in X direction
Gantry X position
𝞍
N
Z axis

## Slide 30: Compensation of the magnetic field
Compensated to 0 but also different offset
  Input from Super-K magnetic field measurements
  Seems to work within 20mG or so
    Z direction has a larger gradient
Bx [mG]
By [mG]
Bz [mG]
Example setting X=100mG, y=0mG, Z=0mG
Compensation table
X[m]
Y[m]
X[m]
Y[m]
X[m]
Y[m]

## Slide 31: No compensation effect
Uncompensated magnetic field
DE plot were generated
  Results looks good
  Analysis pipeline is replicating previous results
  All PMT are working as expected
Monitor PMT
Detection efficiency
Detection efficiency (corrected)

## Slide 32: The magnetic field in Kamioka
Earth field is compensated in Super-K
Older measurements (2013)
  Showed ± 80 mG in Z, ± 100 mG in Y and ± 80 mG in X
Newer measurements
  Showed ± 100 mG in 3 directions
Does it as an impact ?
YES!

## Slide 33: Data taking plan
Relocation happened in 2020
  Delays due to COVID and unexpected problems
Finally ready for data taking
  First water scan was done !
  Running angular scan !
First water measurements!
Count of photoelectron pulse
The number of pulse
DE

## Slide 34
Decay electron
  Nominal
  Mean ratio :1.03
PTF
  Mean ratio: 1.02
Difference: ~1.09%
Angular distribution

## Slide 35
Decay electron
                Nominal
Mean ratio :1.03
PTF
  Mean ratio: 1.02
Difference: ~1.09%
Angular distribution