## Slide 1: Measurements of a Magnetically Shielded Room for a Neutron EDM Experiment
Thomas Hepworth
University of Winnipeg/TRIUMF on behalf of the TUCAN collaboration

## Slide 2: Overview
In this talk I will:
  Introduce the search for a non-zero neutron electric dipole moment (nEDM) at TRIUMF using ultracold neutrons (UCN)
  Motivate the need for a magnetically shielded room (MSR), and the required specs for the nEDM measurement
  Go over magnetically shielded room status: Results of our testing and next steps
Notes: Ucn production isn’t the actual constraint on finishing MSR. Being done and implementing further subsystems is the reason for wanting to finish the MSR. UCN prod. Is not a bottle neck.

## Slide 3: The Experiment
An nEDM measurement searches for new sources of T and CP symmetry violation
This has implications on new physics scenarios
Current world best limit is 1.8 *10-26 ecm (90% CL)
TUCAN aims for 10-27 ecm (factor of 10 improvement)
Notes: A non-zero electric dipole moment of the neutron would be a violation of parity (P) and time-reversal (T) symmetry. This also implies CP violation, which is needed to explain the dominance of matter over anti-matter in the present day universe.
This can be explained by the following picture: if the neutron has a finite EDM, the charge distribution is reversed under P; it is unchanged under T, but the orientation of a particle is specified by its spin, which is unchanged under P, but reverses under T. Therefore if the EDM is not zero then P and T are not conserved.
Say BSM physics searches is the motivation for searching for forms of CP violation.
Plan to improve current lim by factor of ~~10

## Slide 4
520 MeV
Cyclotron
My talk lives here
Notes: Describe experimental setup. We smash protons from the main 520 MeV TRIUMF cyclotron, they then go through our cryogenics system to slow the neutrons down until they are ultra cold. The UCN are then transported into the electric dipole moment precession chamber which is contained inside of the magnetically shielded room (MSR).
Need to spend real time going through this thing. Neutrons, cold stuff, MSR. Could add more labels to aid my explaining.

## Slide 5: Experimental Technique
dn is measured indirectly through the neutron’s Larmor spin precession frequency
Subtract and we get:
   B must cancel
= 0?
Notes: Precession of magnetic moment is the usual Larmor precession frequency, but if the EDM is non-zero, then then we should see a slight deviation in the precession frequency proportional to this value when we alternate the electric field between parallel and anti-parallel configurations. Measuring this deviation in the precession frequency proportional to the electric field E is extremely challenging is B is changing as well. To see a very small deviation in we must know B very well. It must be very homogenous over the neutron cell, as well as stable over long time frames. Slow drifts in the field which we can track are okay, but more rapid fluctuations in the field would be detrimental to the measurement process.
Say why we want E big, but precision isn’t deal breaker. B must be precisely controlled however.
The presence of an electric field aligned parallel or anti-parallel causes a shift in the precession frequency of the neutron, proportional to the nEDM.
Need to explain technique well so it flows into the next slide. Say more about how we flip E field then why it is so important for B to be known well, maybe.
Figure with E and B orientation.
LEGEND
From the deleted plot notes: If they interact with electric field longer, the dipoles can better align with the electric field as that is their tendency.
If E is off by 10%, who really cares, not a huge detriment to the experiment. If B is fluctuating by 10% we wouldn’t even be able to tell if the precession frequency is changing when we reverse the polarity of the electric field
We want E to be as large as possible so that we can easily measure the d times E term in the Hamiltonian
Ditch?
Write out MSR/
Precession of magnetic moment is the usual Larmor precession frequency, but if the EDM is non-zero, then then we should see a slight deviation in the precession frequency proportional to this value when we alternate the electric field between parallel and anti-parallel configurations. Measuring this deviation in the precession frequency proportional to the electric field E is extremely challenging is B is changing as well. To see a very small deviation in 𝜔 we must know B very well. It must be very homogenous over the neutron cell, as well as stable over long time frames. Slow drifts in the field which we can track are okay, but more rapid fluctuations in the field would be detrimental to the measurement process.
The presence of an electric field aligned parallel or anti-parallel causes a shift in the precession frequency of the neutron, proportional to the nEDM.
Need to explain technique well so it flows into the next slide. Say more about how we flip E field then why it is so important for B to be known well, maybe.

## Slide 6: MSR Design
Ambient magnetic field coils (AMC). Being installed in May 2024.
Door
Motion winches
Holes
Layer | Thickness | External side length
1- Outer | 4 mm (2 x 2 mm) | 3.5 m
2 | 3 mm (2 x 1.5 mm) | 3 m
3 | 3 mm (2 x 1.5 mm) | 2.6 m
4 - Cu | 8 mm | 2.46 m
5 - Inner | 2 mm (2 x 1 mm) | 2.39 m
Notes: Door is a removable wall on rails. We need this door to mate well with
Say the holes are for magnetometers, neutron guides, and I guess for letting heat out.
the room for it to perform well.
Cyclotron field is unique!
Layer 4 Copper primarily shields high frequency.
Other layers do the Low frequency (essentially static magnetic field) shielding.

## Slide 7: MSR Field requirements
Residual field of < 1 nT within the 1 m3 central volume .
Internal gradient must be < 100 pT/m
1pT stability over 100s of sec
To achieve this stability, field inside MSR needs to be around 50,000 times smaller
Notes: The higher the SF the better
Say why we picked the SF number. Shielding 100nT perturbations? Also need to shield slow drifts in the field.
Talk about internal fluctuation fields.
100s comes from measurement length

## Slide 8: Measuring the Shielding Factor
Basic principle:
    We also include several correction factors in this calculation
We previously measured perturbation coil field to be (amplitude) AC field at center.
Example: we measure a field of 18 nT
Notes: List all the correction factors we are accounting for!
Go through example. Explain the setup, the wooden coil etc.
MSR as of Friday!

## Slide 9: Results
Internal field: 6 nT
SF is a factor of 5 smaller than needed
Simulated SF were all > 100,000 at all frequencies of interest
Cyclotron is ON
Cyclotron is OFF
Notes: EXPLAIN FIG FIRST. AXIS and all. Point to the 0.01 Hz point. NEED TO THROUGHLY
State results.
Residual field is the field leftover due to magnetization of material, removed by degaussing.
In order to hit our desired sensitivity limit for the nEDM, we must improve the field inside, ie improve the SF. SF is just a metric we use to think about the field we need on the inside.

## Slide 10: When in doubt add more layers!
We believe simulations were too optimistic when guessing the relative permeability of shielding materials
We are adding an additional inner most mu-metal layer
    Completion by August 2024
According to COMSOL simulations, this should improve the SF by around a factor of 10 at low frequencies.
  Why is this the best course of action?
Notes: Say Why inside vs outside.
We don’t have room on outside, would need to remove cladding, more challenging to design, and more subsystems would have to move.
In some ways, more mu-metal is like the duct tape of magnetically shielded rooms.
Inside. Places less constraints on other subsystems, according to COMSOL simulations it should improve the SF by around a factor of 10.
Note redesign of B0, coil that sets the

## Slide 11: Acknowledgements
Thanks to the NSERC USRA program, TUCAN, and the B. G. Hogg Scholarship in Physics for supporting me in this research and attending this conference
Conclusion:
TUCAN is working toward world best 10-27 ecm nEDM measurement at TRIUMF
We need to improve our magnetic environment i.e. the MSR
We are adding a 6th mu-metal layer to boost our SF to an acceptable level