## Slide 1: The Interaction Region BPMs for the HiLumi LHC Upgrade
Douglas BETT, University of Oxford
JAI Fest 2024, Royal Holloway, 19 December 2024

## Slide 2: Contents
Introduction
Experimental setup
Results
Conclusion

## Slide 3: Introduction
The High-Luminosity LHC (HL-LHC) project aims to increase the integrated luminosity of the LHC by a factor of 10 beyond the original, producing >15 million Higgs boson / year from 2030
Requires upgrading quadrupoles, adding new crab cavities, replacing collimators etc.
Vidal and Manzano

## Slide 4: BPM upgrade
New stripline BPMs in the immediate vicinity of interaction regions 1 (ATLAS) and 5 (CMS), where the two beams coexist within a single pipe
6 either side of each IP (24 in total)
Two different types with type-B BPMs rotated by 45° to accommodate shielding elements
BPM | Type | s [m] | bxt [ns]
Q1 | A | 21.853 | 3.92
Q2A | B | 33.073 | 3.92
Q2B | B | 43.858 | 6.82
Q3 | B | 54.643 | 9.72
CP | B | 65.743 | 10.52
D1 | B | 73.697 | 7.36
CAD model of the Type-B BPM

## Slide 5: Signal processing with two beams
Signals from all eight ports required to calculate the horizontal and vertical position of both beams
Distortion of main beam signals due to presence of counter beam must be compensated for
Simulation framework developed to inform design of a processing algorithm (“counter beam compensation”)
Beam profile
Signal induced at upstream ports
“coupled”
Signal induced at downstream ports
“isolated”

## Slide 6: Example signals
Beam “9” only
Beam “0” only (3.92 ns later)
Blue: coupled; red: isolated
Blue: isolated; red: coupled

## Slide 7: RFSoC
Radio Frequency System-on-Chip: candidate BPM processor
RF-ADCs/DACs, programmable logic, CPU in a single package
Evaluation board “ZCU” acquired for proof-of-concept tests
Xilinx Zynq XCZU28DR (gen. 1)
Source: Zynq UltraScale+ RFSoC Data Sheet: Overview (DS889)
RF Features |  | 
Max. RF input Frequency (GHz) |  | 4
Decimation / Interpolation |  | 1x, 2x, 4x, 8x
12-bit RF-ADC | # of ADCs | 8
 | Max Rate (GSPS) | 4.096
14-bit RF DAC | # of DACs | 8
 | Max Rate (GSPS) | 6.554
Notes: RF Features
Max. RF input Frequency (GHz)
4
Decimation / Interpolation
1x, 2x, 4x, 8x
12-bit RF-ADC
# of ADCs
8
Max Rate (GSPS)
4.096
14-bit RF DAC
# of DACs
8
Max Rate (GSPS)
6.554

## Slide 8: Experimental setup
Custom analogue front-end board by Selim Ozdogan
Software/firmware developed by Irene Degl’Innocenti
Board installed in region accessible during operation
BPMSY.4L5

## Slide 9: Data taking [MD6948]
Beam 1 test bunch
Beam 1 reference bunch
Beam 2 reference bunch
Beam 2 test bunch
Bunch crossing timing
WINDOW 1
WINDOW 2
WINDOW 3
100 ns
200 ns

## Slide 10: Results
Blue: apparent position Red: corrected position
Black: reference bunch
RF cogging
Example correction

## Slide 11: Scan results

## Slide 12: OUTLOOK
Work continues on the RFSoC software and firmware
  Use all eight channels, integration with LHC control system
Paper recently submitted to Phys. Rev. Accel. Beams

## Slide 13: Counter beam compensation concept
Primary challenge of the IR-BPM system is the simultaneous presence of signals from both BPMs
Studies using simulated waveforms suggest the power from a single beam can be estimated from the total power at both ends of a single stripline and some calibration parameters that vary with the bunch crossing timing
+
=
scale factor sets amplitude of each term according to intensity, position of beam
Each waveform can be expressed in the form:

## Slide 14: Counter beam compensation algorithm
Standard position calculation ()
Power compensation
, calculated in real time
, and calculated in advance from reference waveforms
“Power” of the waveform

## Slide 15: Reference waveforms
Reference waveforms sourced from model or (single-beam) observations
Voltage
Sample index

## Slide 16: Power compensation parameters
The parameters are functions of the elementwise sums of the below waveforms
coupled × coupled
isolated × isolated
coupled × isolated