## Slide 1: Characterisation of a 3D aluminium cavity for the QSHS experiments
Joshoua Condicion Esmenda1, Sam Furness1, Edward Laird1,
Paul Smith2, Ed Daw2 and Yuri A. Pashkin1
1 Lancaster University
2 The University of Sheffield

## Slide 2: Al cylindrical cavity drawing
Material: Al grade 6061 Chemical composition (%): Mg 0.8 – 1.2; Si 0.4 – 0.8; Cu 0.15 – 0.4

## Slide 3: Al cylindrical cavity with HTS coating
2.3 GHz, 1.6 L, EuBCO+APC Tapes (Fujikura)
TM010
TM011
Q ~ 3.5e6 @ 8 T
Courtesy of Dahno Ahn, CAPP IBS, Korea

## Slide 4: Cylindrical cavity resonant modes
Transverse Magnetic (TM)
Transverse Electric (TE)
and are the zeroes of the Bessel functions of the first and second kind
is the relative permittivity of the cavity filling
is the relative permeability of the cavity filling
Nodal indices:
 the number of nodal diameters in the circular degree of freedom
 the number of nodal circles in the circular degree of freedom
 the number of nodal planes along the length of the cavity
L
r

## Slide 5: Resonant modes
Fundamental Mode TM010
Second Mode TM011
Third Mode TM012
4th and 5th Modes TM110
6th and 7th Modes TM111
Mode TE111

## Slide 6: Equilibration time of superconducting aluminium
Assume L = 10-1 m
Just at Tc (Pobell’s book)
= 2 x 10-1 J/kg K
= 5 x 102 W/Km
= 2700 kgm-3
→ ~ 10-2 s
~
Plausible combination at 100 mK
= 10-6 J/kg K (Sahling, Abens 2001)
= 10-3 W/Km (Pobell’s book)
= 2700 kgm-3
→ ~ 3 x 10-2 s
The heat equation:
capacity

## Slide 7: Cavity holders
Eight 1 cm thick clamps consisting of two brackets each
1 mm gap in between the top and bottom brackets
121.5 mm
101.5 mm
30 mm
10 mm

## Slide 8: Al cavity attached to the MXC plate
Power In
Power Out
RF Cavity
Copper clamps
Brass screws
MXC Plate
Copper cables
All measurements performed with minimal antenna insertion

## Slide 9: Al cavity attached to the MXC plate
Power In
Power Out
Copper plate to fix the coax lines
All measurements performed with the minimal antenna insertion
Сollet mechanisms
to hold the ends of coax lines

## Slide 10: Wideband scan at room T
TM010
2.5012 GHz
TM011
2.8025 GHz
TM012
3.5412 GHz
TM110
3.9830 GHz
TM111
4.1738 GHz
2.5018 GHz 2.7917 GHz 3.5210 GHz 3.9862 GHz 4.1743 GHz
Resonance frequencies obtained from COMSOL modelling
Notes: Change to dB scale

## Slide 11: Before and after pumping at room T
f= 2.5007
Q = 16,958
f = 2.5016
Q = 16,719
Prediction: 1.00030
Input power = 10 dBm
pumping

## Slide 12: Fundamental mode at different T
//
//
Note: Baselines for each curve were adjusted to be equal for presentation
f = 2. 5108 GHz
Q = 52,737
f = 2. 5120 GHz
Q = 61,172
f = 2. 512096 GHz
Q = 25,044,537
f = 2.5016 GHz
Q = 33,915
Input power = 0 dBm
For T = 300 K  100 K
MHz

## Slide 13: Lowest three modes at base T
FM: Q = 25,044,537 2nd mode: Q = 20,353,042 3rd mode: Q = 21,063,194
f = 2.512 GHz
f = 2.815 GHz
f = 3.557 GHz

## Slide 14: Splitting of higher deg. modes
4th mode: Q = 11,803,640 at -15 dBm
5th mode: Q = 507,029 at -15 dBm
6th mode: Q = 3,739,055 at -15 dBm
7th mode: Q = 560,843 at -15 dBm
f = 4.000606 GHz
f = 4.002217 GHz
f = 4.192729 GHz
f = 4.193083 GHz

## Slide 15: Q vs T
Q is 16.575 x 106 at 16.5 mK
f0 = 2.512100429 GHz
WA Phillips (1987)
(arb. units)

## Slide 16: f vs T
37 kHz shift
which is 0.0015%

## Slide 17: Summary
Al cavity characterised in a wide temperature range, million
Sharp changes of in Q and 0 observed at about Tc
Q vs T dependence dominated by BSC and TLS losses
Frequency shift caused by changes of the total inductance of the cavity