## Slide 1: Commissioning a CTBT Coincidence Detector System Underground at SNOLAB
DEREK KONG | SNOLAB | SUMMER 2022

## Slide 2
The Comprehensive Nuclear-Test-Ban Treaty (CTBT)
UN General Assembly adopted treaty with basic obligations banning detonation of nuclear explosions
Global radionuclide stations set up to monitor potential nuclear events
Xenon radioisotopes are of interest, as possible products of nuclear detonations
Proposed identification by beta-gamma coincidence detection

## Slide 3
Cooper et al. 2019. Radioxenon net count calculations revisited
Radioxenon Isotopes
Beta gamma coincidence detection
Four different isotopes
Distinguishing features:
  Horizontal lines (broad beta spectrum)
  Diagonal lines (Compton scattering)
Beta Energy
Gamma Energy

## Slide 4: The CTBT System
Beta Detector (PIPS)
Gamma Detector(s) (HPGe)
Housing and auxiliary
Data
Preamplifier
Power
Network/Ethernet
Voltage
LN2 (Liquid Nitrogen)
Source
Lynx MCA

## Slide 5: Key Coincidence Patterns
Compton Scattering
Co-60
Full 1st Gamma, Partial 2nd Gamma Energy Deposit
Na-22
Full Beta, Full Gamma Energy Capture
Xe Isotopes
Beta Energy
Gamma Energy
Beta Energy
Gamma Energy
Beta Energy
Gamma Energy

## Slide 6: Gamma Detector (HPGe)
HPGe gamma-gamma system already in use underground
Optimizing Lynx MultiChannel Analyzer (MCA) pulse processing
  Low FWHM values constitute accuracy

## Slide 7
U238 609 keV FWHM Trend
Scatter Representation
Heatmap Representation
Conclusion: Rise Time set to 4.0μs
Flat Top set to 1.8μs

## Slide 8: Beta Detector (PIPS)
Provided by Health Canada following use in CTBT system development
PIPS sensitive to electronic noise, reconfiguration necessary
U238 Beta Spectrum
Rewiring/Isolation

## Slide 9: Rise Time Adjustments
Beta Energy (keV)
Na-22 coincidence @ 3.8μs Rise Time
Beta Energy (keV)
Gamma Energy (keV)
Data runs of ~45 minutes of Na-22
511, 1275, and 1786 keV gamma lines well defined on both
Take HC recommendation of “higher” rise time
Na-22 coincidence @ 2.8μs Rise Time
Beta Energy (keV)
Gamma Energy (keV)
Partial 2nd Gamma Deposit

## Slide 10: Gain
Recommended gain setting by HC: 11.32
Beta-gamma coincidence showed majority of beta counts at < 600 keV
Co-60 b-g coincidence @ 4 Gain
Beta Energy (keV)
Gamma Energy (keV)
Co-60 b-g coincidence @ 11.32 Gain
Beta Energy (keV)
Gamma Energy (keV)
Compton Scattering

## Slide 11
Rocks and soil surrounding SNOLAB have trace amounts of U-238 and Th-232
Decays lead to higher concentration of Rn gas underground
Interferes with background gamma counts, purged via LN2
Underground Radon Rates
Rn-220
t1⁄2 ~ 55s
Rn-222
t1⁄2 ~ 4d
U-238
Th-232
Parent
Daughter

## Slide 12: LN2 Purge Flow Rates
Chart: Gamma Line Counts vs LN2 Flow Rate
351 keV | 0.001587 | 0.001767 | 0.001633 | 0.001589 | 0.001003 | 0.0007639 | 0.0009595 | 0.0009539 | 0.0007982 | 0.0009076 | 0.0005744 | 8.302e-05
609 keV | 0.001192 | 0.001261 | 0.001202 | 0.001234 | 0.000895 | 0.0006671 | 0.0007129 | 0.0007041 | 0.0006866 | 0.0005736 | 0.000505 | 0.0003135
1120 keV | 0.0003441 | 0.0002204 | 0.0003153 | 0.0002964 | 0.0001752 | 0.0002425 | 0.0001797 | 0.0001861 | 0.0001314 | 0.0006511 | 0.0002684 | 0.0002721
1764 keV | 0.0003107 | 0.0004303 | 0.0003394 | 0.0003359 | 0.0002666 | 0.0002233 | 0.0002702 | 0.0005154 | 0.0002937 | 0.0002418 | 0.0002537 | 0.0001546
2 L/min
4 L/min
6 L/min
Three LN2 rates tested against U-238 gamma lines
Data ranging from three to seven days
Lower energy lines saw more drastic decreases in counts per second
Settled on 4 L/min

## Slide 13: The Future
Further optimization of beta detector settings
A second PIPS detector has been installed
Once up to standard, SNOLAB could become premier location for air sample analysis

## Slide 14: Thank You!

## Slide 15: Cooper et al. 2019. Radioxenon net count calculations revisited https://d-nb.info/1201957443/34
Image References