## Slide 1: SACLA's thermionic low-emittance gun and the future perspective
Kazuaki Togawa
RIKEN SPring-8 Center
October 8, 2020
Electron-Ion Collider Workshop

## Slide 2: Outline
Overview of the x-ray free-electron laser facility, SACLA
CeB6 thermionic gun
A problem and a solution of the SACLA’s CeB6 gun
Future perspective

## Slide 3: Overview of the x-ray free-electron laser facility SACLA

## Slide 4: X-ray free-electron laser facility, SACLA
SPring-8 storage ring (8 GeV)
SACLA (8 GeV)

## Slide 5: SACLA accelerator layout
Courtesy of T. Hara

## Slide 6: SACLA accelerator main components
Electron injector
C-band accelerator
In-vacuum undulator
Photon beamline

## Slide 7: SACLA beam specifications
 | BL3 / BL2 | BL1
Electron beam energy | 8.5 GeV (max.) | 800 MeV (max.)
Bunch charge | 0.2 – 0.3 nC | 0.2 – 0.3 nC
Peak current | >10 kA | >300 A
Pulse repetition rate | 60 Hz (max.) | 60 Hz (max.)
Photon energy | 4 – 15 keV (0.08 – 0.3 nm) | 20 – 150 eV (8 – 60 nm)
FEL pulse energy | ~0.5 mJ @10 keV | ~100 μJ @100 eV
Pulse width | <10 fs | <1 ps
Spectrum band width | 0.5 % | 3%

## Slide 8: SACLA beam performances
Electron beam profile
Emittance
15 mm
XFEL beam profile
0.75 mm
Output trend

## Slide 9: CeB6 thermionic gun

## Slide 10: Cerium hexaboride cathode
Electron microscope
Properties
Very flat surface by self evaporation
(surface roughness <1 mm)
Low work function (~2.4 eV)
Long lifetime (>10,000 hours)
Rapid recovery from contamination
Design Parameter (Ideal Case)
Thermal Emittance
Emission Density
= 0.4 mm mrad
> 42 A/cm2

## Slide 11: CeB6 cathode assembly
CeB6 Cathode (f3 mm)
Graphite Sleeve
Graphite Heater
SiN Ceramic Base
Molybdenum
Current Lead
Tantalum Thermal Shield

## Slide 12: 500-kV electron gun test stand (2003)

## Slide 13: Electron gun chamber
-500 kV
Ground
Insulation Oil
Vacuum
Pressure : <3x10-7 Pa
Baking is not needed.
High-voltage Wehnelt
Material : Super-clean SUS316L
Surface : Chemical etching+
Ultra-pure water rinsing

## Slide 14: 500-kV beam performance
Beam Waveform
I-V Curve
Beam Current
(CT Monitor)
Gun Voltage
(CVD Monitor)
Space Charge Limit
Schottky Effect
Droop of CT Monitor
~1 ns part is cut out from the flattop by a beam chopper.
The gun operates in temperature limited region to reduce emittance growth due to space charge.

## Slide 15: Emittance measurement
Transverse Beam Profile
Phase Space Distribution
Beam Energy : 500 keV
Peak Current : 1.0 A
Pulse Width : 3 ms (FWHM)
Normalized rms Emittance
: 1.1 mm mrad (full particles)
: 0.6 mm mrad (90% core)

## Slide 16: Clear beam at the SACLA injector
Hard edge and good uniformity is realized.
Low slice emittance ~0.4 mm mrad should be achieved.
CeB6 Cathode
Circular Collimator
(f4 mm typical)
Beam Chopper
Buncher
Screen Monitor
(Alumina Ceramics)
5 mm
Beam

## Slide 17: A problem and a solution of the SACLA’s CeB6 gun

## Slide 18: Cathode lifetime problem at SACLA
Cathode emission lifetime is restricted to 1-year at SACLA in spite of 2 or 3 years lifetime at the prototype accelerator SCSS.
Backward flow of the electrons from the L-band accelerator may be a source of the short lifetime.
Power (temperature)
explosion

## Slide 19: 1D particle simulation of the SACLA injector
Energy (MeV)
Z (m)
10 ns
20 ns
30 ns
L-band in
Backward flow
Reflection
Particle number：10001
Initial bunch length：1.8 ns
Gun exit (Collimator)
1-dimensional particle simulation without space charge.
High-energy electrons can flow backward and cause damage to the cathode.
Beam kicker installed in front of the L-band accelerator will remove the backward electrons.
10 MeV
20 MeV
30 MeV

## Slide 21: Impure bunch sweeper
An impure bunch sweeper was installed in front of the L-band injector linac to sweep out the unwanted bunch tail in this summer.
This was introduced to remove electron contaminations that mix in empty buckets in the storage ring.
The sweeper will work as a backward electron kicker.

## Slide 22: Future perspective

## Slide 23: Future perspective
Future XFELs require extremely small emittance ().
Emittance relation: ,
Lowering the operation temperature is not effective to achieve a lower emittance due to square root function.
Making the cathode size smaller is important.
In return for the small cathode size, extremely high-density of thermionic emission is required ().
CeB6 cathode has a possibility to produce such a high-density low-emittance beam, if a higher-gradient electric field is applied on the cathode surface.
Search for new-type thermionic cathodes is also important.

## Slide 24: Summary
CeB6 thermionic gun has been derivering low-emittance high-intensity electron beams for SACLA for 10 years.
Cathode lifetime problem may be solved by removing or reducing the backward high-energy electrons from the injector linac.
Reducing an emittance of the cathode is essential to increase an XFEL output. The CeB6 gun should have a possibility.