## Slide 1
Construction and Measurement of a Tuneable Permanent Magnet Quadrupole for Diamond Light Source
Alex Bainbridge
ASTeC, UKRI-STFC Daresbury Laboratory
IoP Particles and Beams Group Meeting
25-26 July 2022
Liverpool, UK

## Slide 2
1 Introduction to ZEPTO
Development and motivation of ZEPTO tuneable permanent magnet technology.
2 Magnet assembly
Challenges of assembling a tuneable permanent magnet with strong forces present.
3 Measurements to date
Gradients and compensation magnetic centre movements.
4 Conclusions
Current and planned developments.
Agenda

## Slide 3
Original project - CLIC
333 dipoles
333 dipoles
42000 quadrupoles
288 dipoles
288 dipoles
ZEPTO (Zero Power Tuneable Optics) project started as a collaboration between CLIC (Compact Linear Collider) and STFC Daresbury Laboratory.
Aim is to save power and costs (both financial and environmental) by switching from resistive electromagnets to permanent magnets.

## Slide 4
Original project - CLIC
333 dipoles
333 dipoles
42000 quadrupoles
288 dipoles
288 dipoles

## Slide 5
The ZEPTO project created 3 prototypes for CLIC.
All were designed, constructed and measured at STFC Daresbury Laboratory.
All use NdFeB magnet blocks attached to a motion system, with fixed steel poles defining the field shape.
Project 1: High strength quad
B gradient from 60 – 15 T/m
Shepherd et al, Journal of Instrumentation, Vol 9, T11006, 2014
Project 2: Low strength quad
B gradient from 43 – 3.5 T/m
Shepherd et al, IPAC’14, Dresden, Germany, paper TUPRO113.
Project 3: Wide range dipole
B adjustable from 1.1 T to 0.45 T
Bainbridge et al., IPAC’17, Copenhagen, Denmark,, paper THPIK105.
Zero Power Tuneable Optics (ZEPTO)
2017-2019
2012 - 2014

## Slide 6
New technology demonstrator
We want to show that the benefits of ZEPTO technology are not limited to CLIC, and can be used as an energy saving `drop-in' replacement for electromagnets on accelerators such as existing light sources.
We have developed a new ZEPTO quadrupole for the booster-to-storage ring section of Diamond Light Source.
0
0.2
0.4
1.2
0.8
1.0
0.6
1.4
0.2
0.4
1.2
0.8
1.0
0.6
1.4
Contours B (T)

## Slide 7
OPERA 3D model
SmCo blocks attached to vertical motion system with 90mm movement range and 1 µm precision
Fixed “inner yoke” defines the quadrupole field
Fixed “outer yoke” can steer flux away from the beam, allowing the gradient to be adjusted
No magnetic contact between yokes, the gap between is filled with aluminium for rigidity
Simulations predicted a gradient tuning range from 23 T/m down to 0.35 T/m
The outer yoke acts as a short circuit for the flux, allowing the magnet to be ‘turned off’.
New technology demonstrator
32 mm aperture

## Slide 8
Previous ZEPTO magnets used large NdFeB blocks to produce the magnetic flux, which are difficult to manufacture and handle.
New magnet uses a “carriage” with a number of smaller blocks set into an aluminium latticework.
Advantages over larger single blocks which include:
Blocks are easier and cheaper to manufacture.
Damaged single blocks may be replaced instead of the entire carriage.
Variation in block BH curves may be compensated.
Modular magnet "families" may be created by strategic removal of blocks.
The gaps between blocks do not affect field quality as the flux must pass through a steel pole before reaching the beam.
We use SmCo instead of NdFeB for improved radiation hardness and temperature stability.
No thermal shunts are needed
Permanent magnets
300 mm
33 mm
50 mm

## Slide 9
In previous ZEPTO quadrupoles field measurements revealed the magnetic centre moving as function of magnet block position, mostly vertically, due to unexpectedly magnetic components.
Expecting this, each magnet carriage now has a dedicated motor and associated driving ball screw to help compensate.
Look-up table can be applied to either carriage position to compensate and restore the magnetic centre to the geometric centre.
Dual motor system
Shepherd et al, IPAC’14, Dresden, Germany, paper TUPRO113.
Motor
PM Carriage
Ball screw
Carriage holder

## Slide 10
Split design
Electromagnetic quadrupoles can typically be split for installation around an existing vacuum vessel – previous ZEPTO quadrupoles could not.
The new version may be split about the vertical axis and reassembled around an existing vacuum system using a dedicated assembly frame.
This allows replacements or repair to be conducted without needing to keep the accelerator offline for the additional time needed to restore the quality of the vacuum.
Vacuum chamber
Temporary PM carriage holder
Pre-bolted inner and outer yokes
Yoke supports
Rail

## Slide 11
Assembly forces
Simulated
Estimated
Extrapolated
Forces between the PM blocks and steel yokes must be calculated in advance and managed during assembly!
A Hall probe was used to measure the flux density entering the steel yoke for as much of the approach as possible, and from this an estimate of the actual force was derived

## Slide 12
Operational forces
FX
FY
The simulated magnetic forces are a function of PM block position and are highly complex
Horizontal forces are balanced on each side of carriages but act on the steel yoke, causing a movement risk to components.
FX
FY

## Slide 13
Gradient vs gap
Some variation between simulated and measured gradients observed.
Gradient lower than predicted but still within requirements (19 to 0.5 T/m normal operation)
Actual range of sweep is 1-85 mm (instead of 0-90 mm) due to limit switch positions.
Tuneability 22.37 to 0.47 T/m
Predicted range was 22.72 T/m to 0.3 T/m

## Slide 14
Vertical centre position
Movements in Y centre observed as a function of carriage position.
Y centre slowly rises as gap enlarges and then suddenly starts to drop.
Cause of movement and direction reversal unknown – magnetic materials in pedestal? Uneven motion?

## Slide 15
Vertical centre position
By offsetting the bottom carriage position from the top carriage position we have found this is correctable!
Works with either carriage, but resulting look-up table is slightly different.

## Slide 16
Vertical centre position
By applying the look-up table we can effectively counter all vertical movement of the magnetic centre!
Dataset shown has 7 µm drop in centre position, with 5 µm drop expected from temperature variation alone (contraction of magnet pedestal).

## Slide 17
Horizontal centre position
Can replicate trend by offsetting one half of outer yoke – Assembly issue?

## Slide 18
Homogeneity – integrated
Integrated gradient homogeneity as function of X at max strength, from ZX map.
Follows similar pattern to sweep across x centre.
Integrated homogeneity is at worst 1.3%
Results almost identical at 40mm carriage position!
Chart: X / mm
| -6 | -5 | -4 | -3 | -2 | -1 | 0 | 1 | 2 | 3 | 4 | 5
gradient of integral / T.m/m | 0.0044944 | -0.00072284 | -0.0025713 | 0.007532 | 0.012773 | 0.0077356 | 0 | 0.0028782 | 0.0086501 | 0.0097304 | 0.0070499 | -0.00021314

## Slide 19
We have built a new ZEPTO quadrupole prototype with several evolutions over the original prototypes.
We have successfully demonstrated that it is possible to build a ZEPTO magnet that can be split and installed around an existing beam pipe.
We have shown that it is possible to keep the vertical magnetic centre position stable by individually moving the magnet blocks according to a look-up table.
More work is still needed to understand and prevent horizontal centre movement.
The tolerances on the outer short-circuit yoke are just as important as on the pole tips!
Conclusions

## Slide 20: Acknowledgements
The team behind ZEPTO DLS prototype:
Ben Shepherd
  + Thanks to Walter Tizzano & Ian Martin from Diamond Light Source
+ Thanks to Alex Headspith, Callum Otter, Ryan Cash, Mike Liptrot, Paul Hindley and Mike Lowe
Alex Bainbridge
Nick Krumpa
  This work was funded by the STFC Proof of Concept Fund
Alex Hinton
Group Leader & ZEPTO co-inventor
Design & Modelling
Mechanical engineering
Measurements

## Slide 21
Thank You
@STFC_matters
Science and Technology Facilities Council
Science and Technology Facilities Council

## Slide 22
1 How much did this prototype cost?
The budget was approximately £100,000 as this was a one-off. Large production runs would be significantly cheaper.
2 How does the cost of a ZEPTO magnet compare to an electromagnetic quadrupole?
The up-front cost of the magnet is higher due to the cost of materials and the labour-intensive assembly process; but money is saved on power supplies, infrastructure, cooling systems and long-term energy bills.
3 Are permanent magnets really a green alternative?
Its complicated. Mining rare earth materials is environmentally damaging and the carbon release from manufacturing a ZEPTO magnet is likely higher than for an electromagnetic alternative. However, the reduced power use can offset this within 1 year depending on the power grid of the country where they are used, and permanent magnet materials are now recyclable.
4 What about higher harmonics?
This design does break symmetry about the 45 degree lines, and so a small octupole harmonic will always be present. In addition, a sextupole harmonic has been measured suggesting the horizontal and vertical gaps between pole tips may not be equal.
Frequently Asked Questions
5 Is ZEPTO technology patented?
Yes, patent numbers WO-2012046036-A1 and PCT/GB2020/050714