## Slide 1: A new design of MPGD: Micro-Mesh Micro Pixel Chamber (M3-PIC)
A.Ochi*, Y.Homma, T.Dohmae, H.Kanoh, T.Keika, S.Kobayashi, Y.Kojima, S.Matsuda, K.Moriya, A.Tanabe, K.Yoshida
Kobe University

## Slide 2: Introduction
Introduction to m-PIC
Design of new M3-PIC
Advantages using micro mesh

## Slide 3: Introduction to m-PIC
M3-PIC is based on m-PIC
  m-PIC：micro pixel gas chamber
  Large area with PCB tech.
  pitch :400μm
  high gas gain
  small discharge damage
4%
15000
1700(with capillary)
Maximum gain
~35%
uniformity（σ）
400μm (300μm possible)
200μm
Pitch
30×30cm2
10×10cm2
Area
>30 days
Long time
7000
1000
Stable Gain
m-PIC
MSGC
Invented by A.Ochi and T.Tanimori ( NIMA 471 (2001) 264)
Application: X-ray imaging, Gamma camera, Medical RI tracing, etc.

## Slide 4: Introduction to m-PIC (cont’d)
Gaseous TPC with micro-electrodes
Scintillation Camera
Recoil Electron : 3D Sequential Track ->Direction, Energy
Less cost than semiconductor Scattered Gamma-ray : Energy, Direction
Tanimori, et al, 2004
α
Applications --- Micro TPC  Gamma ray camera (ETCC: Electron-Tracking Compton Camera)
３D Track
X, Y from m-PIC
＋
Z from timing

## Slide 5: Design of M3-PIC
Micro pixel chamber (m-PIC)
+
With micro mesh
  Higher gain in stable operation (~5x104)
  Low ion backflow (<1%)
400mm
100mm
70mm
230mm
165mm
Anode
Cathode
Support wire
Micro Mesh
1cm
Drift/detection area
(Filled by gas)
Drift plane

## Slide 6: Advantages using micro mesh
Higher gas gain will be attained safely (104-5)
  High electric field is formed larger area around the anode
  Without increase of e-field near cathode edge
    Electron emission from cathode edge is reduced
    Streamer from anode is quenched
Reduction of positive ion backflow
  m-PIC: ~30%
  M3-PIC: < 1%
Anode
Anode
Substrate
Mesh
Cathode
Anode
Substrate
Cathode
E-field strength on m-PIC
E-field strength on M3-PIC

## Slide 7: Setup and operation tests
Setup
Gas gain measurements
Ion backflow measurements

## Slide 8: Setup
0.5mm
Micro scope pictures
for same place
(different focus point)
Micro mesh mounted on m-PIC by hand.
Size of m-PIC = 3cm x 3cm.

## Slide 9: Gas gain measurements
Gain dependency on
    Anode voltage (=Va)
    Mesh voltage (=Vm)
    Drift field (=(Vd-Vm)/1cm)
165 mm
100mm
Drift Plane
Cathode
Anode
Signal
Vd
Va
Mesh
Vm
1cm

## Slide 10: Gain dependence of drift field
Higher drift field
  Lower electron collection efficiency on anode
  Gain decrease
  Energy resolution worse
    No escape peak found in 2kV/cm
Maximum gain
  100V/cm < E_d < 500V/cm
  E_d below 100V/cm
     Ion-electron recombination
E_drift = 300V/cm
E_drift = 1kV/cm
E_drift = 2kV/cm

## Slide 11: Electron drifts toward anode (simulation)
Electrons are absorbed in the mesh or cathodes when electric field in drift region is higer.
E_drift = 300V/cm
E_drift = 1kV/cm
anode
anode
cathode
cathode
mesh
mesh

## Slide 12: Gain dependence on Vm and Va
E_drift = 300V/cm
Maximum gain : 5 x 104
Gap of mesh: 165mm
Mesh thickness: 5mm
Gas: Ar:C2H6 = 50:50

## Slide 13: Ion backflow (IBF)
Ion backflow: The fraction of total avalanche-generated ions reaching to drift region.
Serious problem for TPC readout
m-PIC (no mesh)
M3-PIC
Low IBF
Simulation
Simulation
Ion drift lines from anodes

## Slide 14: Setup for ion backflow measurements
Pico-ammeter is inserted in Drift and Anode line
IBF = ( I_drift/I_anode)
Wireless data taking for keeping insulation
100mm
Drift Plane
Cathode
Anode
Vd
(-100V~-1kV)
Va
(~500V)
Mesh
(5mm or 20mm)
Vm
(-150V~350V)
A
A
Pico ammeter
Pico ammeter
Data collection by wireless
b (Sr90)

## Slide 15: IBF of M3-PIC
IBF dependence on drift field and mesh thickness
IBF dependence on mesh voltage
Gas: Ar 90% + C2H6 10%
Gain = 104 for these tests
Liner dependence of IBF on drift field
Small IBF for thicker mesh
Optimum point of mesh voltage
Minimum IBF = 0.5%
Vm = -250V
E_drift = 100V/cm
Mesh 20 micron

## Slide 16: Summary
New MPGD design: M3-PIC was developed
  Combined with m-PIC and micro mesh
  Ideal electrical fields are formed around anodes for gas avalanche
Prototype was manufactured and tested
Maximum gain of 5 x 104 has been attained at present
  A few time larger than the gain of simple m-PIC
Minimum IBF of 0.5% has been attained at present
Future Prospects
Optimization of structure parameters (mesh gap, mesh thickness … etc.) and operation parameters (HV, gas etc.)
Combination with existence large area m-PIC
  Testing imaging and tracking capabilities
  Long term operation test