## Slide 1
Development of a camera for imaging of prompt gamma rays in measurements of proton beam range
P.Busca1,2, C.Fiorini1,2, R.Peloso1,2, M.Basilavecchia3, T.Frizzi3, D. Prieels4,
J. Smeets5*, F. Stichelbaut4, A. Benilov4, F. Roellinghoff4
1Politecnico di Milano and INFN, Milano, Italy
2XGLab, Milano, Italy
4 Ion Beam Applications S.A., Louvain-la-Neuve, Belgium
5 Department of Nuclear Metrology, Université Libre de Bruxelles, Brussels, Belgium
Work supported in part by the European Union Seventh Framework Program (FP7/2007-2013) under grant agreements nos. 241851 (ENVISION) and 264552 (ENTERVISION).
HICAM gamma detector developed within the EC contract n.LSHC-CT-2006-037737.
*supported by the Belgian FNRS (aspirant)
Notes: B

## Slide 2
Proton therapy
relative dose
depth
g
p
Bragg-peak
particle (proton) therapy has a growing role in cancer treatment.
possibility to release the maximum of the dose in the target site, limiting the dose to normal tissue
the measurement of the proton beam range in the target is very important: real range of proton beams in patients may contain uncertainties of up to 10-15 mm (uncertainties on tissue composition, density, organ motions, patient positioning, etc).

## Slide 3
One method to measure proton beam range is based on the measurement of prompt gamma rays (energies up to 10MeV) emitted by excited nuclei during proton irradiation.
(F. Stichelbaut, Y. Jongen, 39th Meet. of the Particle Therapy Co-Operative Group, San Francisco, October 2003)
Monte Carlo
simulation
p beam
target
p beam
prompt
gs

## Slide 4
Practical concept of a prompt gamma camera which allows checking in real-time the range of a single pencil beam with a ‘mm’ accuracy
(D. Prieels, J. Smeets, F. Stichelbaut, A. Benilov, J.C. Dehaes, A. Dubus, F. Roellinghoff, 50th Meet. of the Particle Therapy Co-Operative Group (Philadelphia), May 2011

## Slide 5: HICAM:HIgh resolution gamma CAMera
- CsI(Tl) scintillator
- SDDs photodetectors
- 5x6cm2 and 10x12cm2 formats
intrinsic resolution~1mm
compactness
Anger camera principle
scintillation entering side
drift
Silicon Drift Detectors
Research project from European Community
(www.hi-cam.org)
Notes: 24 Novembre 2008
5
B

## Slide 6
The HICAM gamma camera development
(R.Peloso, et al., “The HICAM Gamma Camera”, Nuclear Science Symposium Conference Record (NSS/MIC), 2010 IEEE)

## Slide 7
Imaging performances with 99Tc (140keV)
LEUHR parallel hole collimator
 intrinsic resolution= 0.8mm
10,8 cm
8,7 cm
corrected for linearity and uniformity
Hole Ø=1mm
Pitch=3mm
Grid
99Tc
Notes: A
B

## Slide 8
Applications of the HICAM gamma camera
(P.Busca, et al., “Applications of the HICAM Gamma Camera”, Nuclear Science Symposium Conference Record (NSS/MIC), 2010 IEEE)
Notes: The hicam camera was tested in differents reserch fields and the obtained result will be presented on Thursday

## Slide 9
1cm thick LYSO scintillator
(h~37%@1MeV, ~22%@5MeV)
light collection purposely limited (i.e. all absorbing surfaces) to match readout ASIC dynamic range (designed for 200keV)
( ~ 5e-/keV, vs. 30ph./keV from LYSO)
Modification of the 5x6cm2
HICAM gamma camera
for high-energy (1-10MeV)
imaging

## Slide 10
Camera FOV
source
lead block
lead block thickness=3cm
Source: 60Co (1.17MeV, 1.33MeV)
Area considered for profile calculation
1D profile calculation along X:
integration on the Y direction
(1mm bin)
Laboratory characterization of the camera
Y
X

## Slide 12
slit collimator
opening: 6mm
Proton range measurements:
the experimental set-up
p beam

## Slide 13
PMMA target
proton
beam
slit collimator
HICAM
camera
g-rays
electronics

## Slide 14
Measured energy spectrum
LYSO
background
3-7MeV
range
used in
reconstruction

## Slide 15
Beam energy: 160MeV
slits closed
(neutrons + uncorrelated photons)
central profile
is fitted and
monitored
while moving the
target (and the Bragg-peak)
camera
position
- 4cm
camera
centered
camera
position
+ 4cm
Measured profiles
subtraction
simulation

## Slide 16
100MeV protons
each acquisition:
5min.
7·1010 protons
-10mm
+10mm
(preliminary analysis)

## Slide 17
Shift accuracy
X10
X0
X-10
reference value

## Slide 18
Precision
reference position
reference value
fitting error

## Slide 20
7·1010 protons
Precision vs. statistics

## Slide 21
160MeV protons
each acquisition:
2min.
1.7·1010 protons
(preliminary analysis)

## Slide 22
1.7·1010 protons

## Slide 23: Conclusions
Measurements of prompt gammas correlated to proton range were successfully made using the HICAM gamma camera
Although not optimized for 1-10MeV energy range (low efficiency, low scintillator light collection), satisfactory accuracy and precision (in the ‘mm’ range) were first measured
Future work
Further data analysis also to correlate results to dose expected in patient treatment.
Measurements to be repeated also with other protons energies (230MeV).
Gamma camera to be modified/redesigned to improve efficiency and speed