## Slide 1: Track based Alignment of the ATLAS Inner Detector Tracking System
Jochen Schieck
Ludwig-Maximilians-Universität München und Excellence Cluster Universe
for the ATLAS Collaboration
9th International Conference on Position Sensitive Detectors Aberystwyth 12th - 16th, September 2011

## Slide 2: The ATLAS Detector
ATLAS Inner Detector (ID) main tracking device of ATLAS
  consists of Pixel, Silicon strip (SCT) and drift tube (TRT) detectors
  single hit resolution between 10 μm (Pixel) and 130 μm (TRT)
You are here
Notes: Note:
The performance in the Endcaps the resolution increases by ~20% and multiple scattering contributes double

## Slide 3: Requirements from Tracking
construction and installation precision ~100 μm
track parameter resolution should not be degraded by more than 20%
  7 μm precision in RΦ for Pixel
  12 μm precision in RΦ for SCT
ultimate precision can be achieved with track based alignment procedure only
hardware based system (FSI) available in the SCT to monitor fast movements of support structures
momentum
ATLAS
ID TDR
impact parameter
Notes: up to 1um precision for W mass measurement

## Slide 4: How to align a Detector
residual
# of hits
residual = distance between
hit and the fitted track

## Slide 5: How to align Detector
residual
# of hits
imagine the real detector
looks like this:
 minimize residuals with respect to the module’s positions
but we think it
is like this:
if we had the
correct knowledge:

## Slide 6: How to align a Detector
residual
# of hits
residual = distance between
hit and the fitted track

## Slide 7: How to align Detector
residual
# of hits
imagine the real detector is misaligned, but we think it is perfect
 minimize residuals with respect to the module’s positions

## Slide 8: How to align Detector
residual
# of hits
imagine the real detector
looks like this:
 minimize residuals with respect to the module’s positions
but we think it
is like this:
if we had the
correct knowledge:

## Slide 9: Track based alignment procedure
ai
collect residual distributions for each module and minimize with respect to the alignment parameters
module-to-module correlations are taken into account by nested dependence on track parameters:
increased convergence speed
derivative of the residual w.r.t. the alignment parameters
derivative of the residual w.r.t. track
derivative of the track parameters w.r.t. alignment parameters

## Slide 10: Track based alignment procedure
start from survey information (Pixel)
perform hierarchically alignment following detector structure
align up to 6 degrees of freedom (d.o.f.) per structure
  35k d.o.f for Pixel and SCT (6 d.o.f. per module)
  700k d.o.f for TRT (2 d.o.f. per straw)
iterative procedure
  10-20 steps for production of complete alignment set
test2

## Slide 11: Input to the Alignment Algorithm
high pT–tracks to minimize multiple scattering effects
  tracks from collision events
    enough tracks during collisions produced –statistics no problem
    large time consumption due to repeating reconstruction of tracks in iterative alignment
    production of reduced data stream with selected high pT-tracks containing ID and track related information only
  tracks from cosmic ray events
    different track topology
    tracks connect top and bottom of ID
    hardly any illumination of endcap

## Slide 12: Residual Distribution
alignment minimizes the Χ2 all measurement residuals
  alignment returns residual distribution which is almost consistent with MC residual distribution with perfect geometry
  this is a necessary but not sufficient condition
  resulting Χ2-value may end up at a local Χ2-minimum (or valley – weak modes of the solution)

## Slide 13: ID Alignment – Quality Assurance
estimate alignment performance using control plots not based on residual distributions
  main focus on the improvement of the momentum scale
ATLAS Simulation
ATLAS-PHYS-PUB-2009-080

## Slide 14: Overall Position of Detector
alignment procedure returns internally consistent set of alignment constants
  6 d.o.f. (= overall position of detector) cannot be determined from alignment procedure
  initial position determined by requiring minimal module-by-module deviation from the perfect geometry
keep average position consistent for alignment updates (keep beam spot position constant)
  ID may be misaligned with respect to the solenoidal magnetic field (relative tilt)

## Slide 15: Solenoid Field Tilt
measure the invariant mass variation with respect to PDG-value (ΔM) as function of Φ
  tilt of the magnetic field is reflected in Φ-dependent J/ψ-mass
  rotation around the x-axis by 0.55 mrad reduces Φ-dependence significantly
Endcap A
Endcap C
x
z
y
B-Field

## Slide 16: E/p Momentum Constraint
identical calorimeter response for e+ and e-
E/p measurement sensitive to charge asymmetric bias on p-measurement (like curl or sagitta distortions)
measure E/p (η,Φ) and use offset between e+ and e- to constrain track momentum in alignment

## Slide 17: Improved Mass resolution using E/p constraint
mass resolution of Z significantly reduced with E/p constraint tracks used in the alignment procedure
data performance nearly reproduces MC predictions

## Slide 18: Improved Mass resolution using E/p constraint
E/p constraint removed Φ-dependence of Z-mass

## Slide 19: Effect on Impact Parameter Measurement
input tracks for alignment constraint to originate from primary vertex (d0= 0)
impact parameter measurements as a function of η and Φ indicate residual misalignments
  Δd0 up to 10 μm
impact on impact parameter based measurements expected

## Slide 20: Impact Parameter based Measurements
exclusive Bd lifetime measurement
τBd=1.51±0.04+0.04 ps (PDG: 1.519±0.007 ps)
(0.04 ps ~ 20 um for 10 GeV B-Meson)
ATLAS-CONF-2011-102
ATLAS-CONF-2011-092
impact parameter based b-flavour tagging
negative part of signed impact parameter (=resolution) returns up to 10% difference to Monte Carlo

## Slide 21: Some Final Remarks
production of alignment constants is time and computing intense
  several iterations with track (finding and) reconstruction required
  reduced data format (e.g. data stream with hits of high pT-tracks only) improves turnaround and speeds up validation
cosmic tracks are highly valuable for alignment performance
  consider trigger of cosmic tracks during data taking
minimizing residuals is the easy part
  focus early on measurements sensitive to global distortions

## Slide 22: Summary
alignment of ATLAS ID uses track based alignment
  minimization of residuals with respect to alignment parameters
residual distribution almost identical to MC with perfect geometry
  residual minimization necessary but not sufficient
momentum measurement considerably improved by
  realignment of B-field with respect to the ID
  require identical E/p measurement for e+ and e-
measurement of average offset of impact parameter d0 better than 10 μm

## Slide 23: Bonus Material

## Slide 24: The ATLAS Inner Detector
subdetector | r (cm) | element size | resolution ( X * Y) | hits/ track (Barrel) | channels
Pixel (Silicon) | 5-12.5 | 50mm * 400mm | 10mm * 115mm | 3 | 80x106
SCT (Silicon Strip) | 30-52 | 80mm * 12cm (stereo) | 17mm * 580mm | 4 | 6x106
TRT (straw tubes) | 56-107 | 4mm * 74cm | 130mm | 30 | 0.4x106

## Slide 25: Time Dependence of Alignment
single set of alignment constants used for 2011 data
small (<10 μm) changes visible
  changes related to mechanical interventions (ie. power cut)
pixel stable during data taking period
Notes: 179939 & 180149 there was a cooling system failure
180481 & 180614 there was a power cut
180710 & 182284 there was a LHC technical stop where the cooling & power was turned off
182519 & 182726 a fire alarm went off and as a result the cooling system was required to be shut down
182787 & 183003 the toroidal magnetic field was dumped

## Slide 26: Residual Distribution
mean value of residual distribution consistent with expectations from Simulation
significant improvement achieved by alignment using individual module DoFs