## Slide 1: Development of a gamma-ray imager using a large area monolithic 4x4 MPPC array for a future PET scanner
Takeshi Nakamori
T. Kato, J. Kataoka, T. Miura, H. Matsuda (Waseda U),
K. Sato, Y. Ishikawa, K. Yamamura, N. Kawabata (Hamamatsu),
H. Ikeda, G. Sato (ISAS/JAXA), K. Kamada (Furukawa Co., Ltd.)
12 September 2011
Position Sensitive Detectors 9 @ Aberystwyth

## Slide 2: Contents
PET and Semiconductor sensors
Performance of the MPPC array
Charge division readout
Other applications
Summary
2

## Slide 3: PET and our approach
Semiconductor photosensors are promising & successful
Positron Emission Tomography : e+e- → gg imaging
Compactness, low power, mass productivity, easy handling…
Insensitive to magnetic fields
APD-PET project
→ Realized large number of channels, fine/complex configurations
→ “DOI-PET”
→ “MRI-PET”
→ dedicated LSI (Koizumi+10)
→ sub-mm resolution (Kataoka+10)
→ >a few ns time res. (Matsuda)
Kataoka+10
16x16 APDs
MPPC can overcome !!
PET module
22Na image
TOF
DOI
3
TOF & DOI information
improve position accuracy

## Slide 4: Muti-Pixel Photon Counter
2D-array of APDs operated in Geiger mode
Charges proportional to the number of fired APDs
low bias voltage (<100V)
high gain (105-6)
Insensitive to magnetic field
4

## Slide 5: Characteristics summary
 | PMT | PD | APD | MPPC
Gain | 105-6 | 1 | 50-100 | 105-6
Q.E. (PDE) | >25 | >80 |  | >25
Volume | large | small |  | 
Interfered by B | Yes | No |  | 
Structure | complex | simple |  | 
Power Consumption | high | low |  | 
better for PET
MPPC vs APD
higher gain, doesn’t need CSAs
  much better S/N
  much better timing resolution (suit for TOF-PET, next slide)
less photo-detection efficiency
  worse energy resolution (see Poster [20] by Miura )
narrower dynamic range due to the limited number of pixels
  need linearity correction
5

## Slide 6: TOF Timing resolution
CFD
CFD
delay
TAC
MCA
22Na
MPPC or
APD
MPPC (G=9x105):
624ps (FWHM)
APD (G=50):
5300ps (FWHM)
LYSO
LYSO
APD : Charge Sensitive Amp.
MPPC: Current Amp.
MPPC
APD
Kato
APDs always require CSA
that limits time resolution.
6

## Slide 7: The Monolithic Array
4x4 array with 3x3 mm2 pixel
0.2 mm gap
50 um type (3600 APDs/pixel)
16 anodes, common cathode
A bit high dark count rate ~2Mcps @ 20 oC
(this was the first prototype: ~400 kHz in recent products)
Gain (105)
Bias Voltage (V)
4
10
70.8
72.0
Gain map (71.9V, 0 oC )
7.2%
ave. gain = 9.7x105
7
Kato+11, NIM A

## Slide 8: Scintillator array
Kamada+10
Pr:LuAG
Pr:LuAG+WLS
3x3x10 mm3 crystals
4x4 arrays
0.2 mm-thick BaSO4 reflector
LYSO(Ce)
LuAG(Pr)
LuAG(Pr) + WLS coating
Wavelength (nm)
Relative light yield
(in this work)
LYSO
LuAG
LuAG+WLS
310nm
420nm
(HPK; 1x1 mm2 type)
r = 7.1g/cc
t = 40ns
25 kph/MeV
r = 6.7g/cc
t = 20ns
20 kph/MeV
faster (better time res.)
8

## Slide 9: Energy spectra
Kato+11
LYSO
LuAG
LYSO :13.8%
LuAG :14.7%
LuAG+WLS :14.0%
LuAG+WLS
137Cs source, 0 oC, 71.9V
w/o current amplifier
Linearity corrected
Discrete readout with Q-ADC
Energy resolution for 662 keV:
9

## Slide 10: Charge spectra
Kato+11
WLS enhanced the light yields detected by ~30%
Still much less than LYSO
Yet we prefer LuAG for better timing resolution
“UV-enhanced MPPC” could be a solution
Yoshino+11
channel A-1
10
Q.E. of UV-enhanced APD
310nm

## Slide 11: Charge division readout
Can reduce the number of read out
Often applied for MAPMT
0 oC, 71.9V, LYSO array, 137Cs
Interaction positions are nicely resolved
Spectra from each pixel extracted
X position (arb unit)
Y position (arb. unit)
11
LYSO
ave. DE/E ~ 9.9 %(FWHM)
ave. FHWM ~ x:0.274, y:0.263 mm
Kato

## Slide 12: Optimization of R-chain
Minimize averaged FWHM (mm) from X- and Y-projection :
Too many degrees of freedom
Just tried 3 criteria with Rx and Ry
12
Rx
Rx
Ry
(Rx=Ry)
(Rx=39W)
(Rx/Ry=1/2)
averaged FWHM (mm)
0.6
0.2
Rx (W)
Ry (W)
Rx (W)
0
200
102
103
0
100
r
y
x
r
y
x
r
y
x
(Rx, Ry)= (51W, 100W) is the best here,
but there could be better ones…
80 W
51 W
sr = sqrt(sx2+sy2)

## Slide 13: Our efforts
S
Gamma-ray measurement
PET application
Better pos. res.
TOF
Low threshold
Kato+11 (in prep)
Miura+
4x4 + FPC
dedicated “fast” LSI
Coincidence technique
Waveform-DAQ
BG surpression
Phoswitch counter
MPPCs
development
Matsuda+11 (in prep)
Sub-mm LYSO array
Poster [20] by Miura
DOI
Miura+
13
Buttable 8x8
Waveform-DAQ
signal precision

## Slide 14: Waveform acquisition
14
Domino Ring Sampler 4 : Fast analog memory LSI
developed for particle experiments (MEG, MAGIC,..)
  Low cost, low power : 140 mW/8ch
  fast sampling : Max 5 GHz, 1 V/12 bit
Time (ns)
Voltage (mV)
Suit for large number of channels
Digital filtering & noise reduction
Capable of pulse shape discrimination
etc… lots of potential
Applicable to (DOI-)PET !
Demonstration
3x3mm2 MPPC + LYSO
w/ current amp.
20 oC, G=7.5e5
Spectra obtained
Noise level reduced
137Cs
137Cs
109Cd
Energy (keV)
Energy (keV)
Q-ADC
waveform
(S. Ritt+)
662keV
88 keV
22 keV
(digital filtered)

## Slide 15: Summary
MPPC is a promising photosensor, especially for TOF-PET scanner
We developed a monolithic 4x4 MPPC array to be applied for PET
We showed the performance of the array as a gamma-ray detector
LYSO(Ce) is better than LuAG(Pr) at this moment, even with the wavelength shifter
We also demonstrated the charge division readout which works well.
Lots of wonderful results will be published soon !
15