## Slide 1: A proton Computed Tomography based medical imaging system
M. Scaringella1, M. Bruzzi2,3, M. Bucciolini3,4,10, M. Carpinelli8,9, G. A. P. Cirrone5, C. Civinini3, G. Cuttone5, D. Lo Presti6,7, S. Pallotta3,4,10, C. Pugliatti6,7, N. Randazzo6, F. Romano5, V. Sipala8,9, C. Stancampiano5, C. Talamonti3,4,10, E. Vanzi11, M. Zani3,4
1Information Engineering Department, University of Florence, Florence, Italy 2Physics and Astronomy Department, University of Florence, Florence, Italy 3INFN - Florence Division, Florence, Italy 4Department of Biomedical, Experimental and Clinical Sciences, University of Florence, Florence, Italy 5INFN - Laboratori Nazionali del Sud, Catania, Italy 6INFN - Catania Division, Catania, Italy 7Physics and Astronomy Department, University of Catania, Catania, Italy 8Chemistry and Pharmacy Department, University of Sassari, Sassari, Italy 9INFN Cagliari Division, Cagliari, Italy
10 SOD Fisica Medica, Azienda Ospedaliero-Universitaria Careggi, Firenze, Italy
11 Fisica Sanitaria, Azienda Ospedaliero-Universitaria Senese, Siena, Italy
PSD10 – 7-12 September – University of Surrey

## Slide 2: proton radiotherapy
Protons release a high dose in a well defined region at the end of their range (Bragg peak) → sparing of healthy tissue
Range can be modulated by varying the proton energy
The released dose can be shaped to the target tumor through the weighted superposition of proton beams at different energies (Spread Out Bragg Peak - SOBP)
ADVANTAGES
PRESENT LIMITATIONS
Uncenrtanties in the measurement of the Stopping Power distribution inside the tissue

## Slide 3: motivations for a proton imaging system
Treatment planning:
Presently defined using X-CT
but protons and photons interact differently with
matter
pCT Direct measurement of the stopping power maps with the same particle used to irradiate
Patient positioning:
Presently this is done using conventional X ray tomographies (X-CT) taken before the proton treatment session and in a potentially different setup:
pCT Precision improvement if positioning and treatment could be done in one go

## Slide 4: Tracks with multiple scattering
L
Measurements:
entry position and angle
Proton true trajectory
L  straight line with
confidence limits
Measurements: entry and
exit position and angle, exit position
L’
L’  straight line with
confidence limits
Measurements: entry and
Exit position and angle +
Most Likely Path (MLP)
calculation
L’’
L’’ curved trajectory with
Narrower confidence limits

## Slide 5: Most likely path error envelope
MLP example with 200MeV
kinetic energy protons in
20cm of water:
Entry: Y(0) = 0.2cm
Y’(0) = -10mrad
Exit: Y(20) = -0.1cm
Y’(20) = +10mrad
Silicon microstrip detectors:
320mm thick
200mm strip pitch
[cm]
[cm]
200MeV in
90MeV out

## Slide 6: pCT apparatus
Single particle proton tracking: silicon strip detectors → MLP
Residual energy measurement: crystal calorimeter → energy loss
The single event information can be processed by reconstruction algorithms to produce the tomographic image
P1
P2
P3
P4
z
x
y
PARAMETER | VALUE
Proton beam kinetic energy | 250 -270 MeV
Proton beam rate | 1 MHz
Spatial resolution | < 1 mm
Electronic density resolution | <1%
Detector radiation hardness | >1000 Gy
Dose per scan | < 5 cGy

## Slide 7: INFN-PRIMA collaboration: pCT apparatus
Four x-y silicon microstrip based
tracking planes
Yag:Ce calorimeter
Proton entry and exit positions and directions
Proton residual energy
First test at INFN-LNS: May 2011
CATANA beam line:
62 MeV protons

## Slide 8: Tracker module
Parallel strip read-out
Local data storing during measurement
Ethernet data download at measurement completion
Front-end board
Digital board

## Slide 9: Si Sensor and Front-end ASIC
Binary read-out:
6.6 x 1.6 mm2
32 inputs - 32 outputs
670 mW power consumption
Vcc=+3.3 V
p on n
single sided
<100>
200mm thick
200mm strip pitch

## Slide 10: Tracker performance
62 MeV proton
minimum released charge
in 200mm of silicon
180 MeV proton
minimum released charge
in 200mm of silicon
Calibration
Efficiency
Signal duration
Noise occupancies
Data LNS+TSL

## Slide 11: Calorimeter and DAQ
4 YAG:Ce
scintillating
crystals
30 x 30 mm2 x 100mm
each
4 Photodiodes
18x 18 mm2

## Slide 12: INFN-Prima: pCT image
PMMA phantom
36 projection steps: 0°  360°
An average of 950000 events per projection
E0=62MeV INFN-LNS
Filtered Back Projection algorithm
Tomographic equation
(Wang, Med.Phys. 37(8), 2010: 4138)
«projection»
Unknown stopping power distribution (at E0)
Evaluation of the “projection” term (through numerical integration starting from NIST tables and using the measured Eres)
Eres
Wang projection
6 mm
4 mm
20 mm
20 mm

## Slide 13: INFN-Prima: pCT image
No selection cuts has been applied to the data sample
E. Vanzi et al. The PRIMA collaboration: preliminary results in FBP reconstruction of pCT data –Nucl. Instr. Meth. A, 730 (2013), p. 184
Chart: Resolution (no cut)
| 2 | 4 | 6 | 8 | 10 | 12
20/128 | 0.78762 | 0.88283 | 0.93064 | 0.96498 | 0.99158 | 1.0128
noise: 6.3%
noise: 2.4%
Butterworth filter: order 4, cut-off 20/128 of the Nyquist freq. (Nyf=2.5mm-1 )

## Slide 14: INFN pCT upgrade
A system similar to the one already tested
  Microstrip tracker
  YAG:Ce calorimeter
But with a 50 x 200 mm2 field of view
On-line data aquisition
1 MHz capability
Rectangular aspect ratio to perform tomographies in slices
Beam pipe
Tracker planes
Phantom
Calorimeter
Silicon sensors

## Slide 15: Silicon microstrip detectors
Chart: Bias Voltage (V)
| 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 45 | 50 | 55 | 60 | 65 | 70 | 75 | 80 | 85 | 90 | 95 | 100 | 105 | 110 | 115 | 120 | 125 | 130 | 135 | 140 | 145 | 150 | 155 | 160 | 165 | 170 | 175 | 180 | 185 | 190 | 195 | 200
 | 16.679 | 22.76 | 28.8 | 34.8 | 38.78 | 41.4 | 43.65 | 45.63 | 47.42 | 49.06 | 50.51 | 51.89 | 53.12 | 54.27 | 55.3 | 56.25 | 57.13 | 57.96 | 58.76 | 59.54 | 60.36 | 61.05 | 61.78 | 62.5 | 63.17 | 63.86 | 64.53 | 65.15 | 65.8 | 66.41 | 67.02 | 67.58 | 68.18 | 68.72 | 69.26 | 69.82 | 70.34 | 70.87 | 71.4 | 71.89
 | 19.48 | 27.89 | 35.74 | 43.89 | 48.79 | 51.89 | 54.51 | 56.81 | 58.89 | 60.76 | 62.43 | 64.03 | 65.37 | 66.62 | 67.77 | 68.97 | 69.74 | 70.65 | 71.5 | 72.33 | 73.4 | 73.94 | 74.72 | 75.47 | 76.17 | 77.1 | 77.54 | 78.21 | 78.88 | 79.51 | 80.17 | 80.78 | 81.41 | 81.99 | 82.57 | 83.14 | 83.7 | 84.25 | 84.81 | 85.32
 | 15.89 | 21.96 | 27.65 | 33.58 | 37.29 | 39.75 | 41.79 | 43.57 | 45.21 | 46.66 | 48 | 49.25 | 50.36 | 51.39 | 52.36 | 53.25 | 54.07 | 54.86 | 55.64 | 56.36 | 57.13 | 57.82 | 58.51 | 59.19 | 59.81 | 60.45 | 61.07 | 61.69 | 62.31 | 62.89 | 63.47 | 64.02 | 64.58 | 65.14 | 65.64 | 66.18 | 66.66 | 67.19 | 67.7 | 68.21
 | 20.94 | 30.21 | 38.5 | 47.42 | 52.67 | 55.92 | 58.67 | 61.15 | 63.36 | 65.3 | 67.1 | 68.7 | 70.12 | 71.43 | 72.57 | 73.6 | 74.58 | 75.44 | 76.3 | 77.1 | 77.95 | 78.66 | 79.42 | 80.15 | 80.86 | 81.57 | 82.22 | 82.89 | 83.58 | 84.18 | 84.86 | 85.44 | 86.05 | 86.68 | 87.24 | 87.83 | 88.38 | 88.94 | 89.5 | 90.02
 | 16.047 | 22.3 | 28.18 | 34.3 | 38.27 | 40.74 | 42.91 | 44.9 | 46.61 | 48.14 | 49.81 | 50.88 | 52.1 | 53.16 | 54.21 | 55.12 | 55.99 | 56.82 | 57.64 | 58.38 | 59.21 | 59.9 | 60.61 | 61.33 | 61.95 | 62.62 | 63.25 | 63.88 | 64.52 | 65.11 | 65.7 | 66.26 | 66.87 | 67.4 | 67.93 | 68.5 | 69.02 | 69.53 | 70.07 | 70.56
 | 27.65 | 41.81 | 54.42 | 67.94 | 74.86 | 79.29 | 82.83 | 86.04 | 88.79 | 91.3 | 93.52 | 95.49 | 97.25 | 98.79 | 100.25 | 101.4 | 102.49 | 103.51 | 104.47 | 105.4 | 106.38 | 107.23 | 108.08 | 108.9 | 109.67 | 110.43 | 111.19 | 111.95 | 112.71 | 113.4 | 114.14 | 114.8 | 115.48 | 116.17 | 116.8 | 117.48 | 118.1 | 118.76 | 119.4 | 119.99
 | 21.09 | 30.7 | 39.32 | 48.63 | 53.67 | 57.18 | 60.02 | 62.52 | 64.76 | 66.74 | 68.54 | 70.16 | 71.63 | 72.94 | 74.11 | 75.16 | 76.09 | 77 | 77.87 | 79.01 | 79.58 | 80.51 | 81.11 | 81.86 | 82.56 | 83.29 | 83.97 | 84.69 | 85.37 | 86 | 86.65 | 87.24 | 87.87 | 88.47 | 89.05 | 89.65 | 90.21 | 90.9 | 91.4 | 91.93
 | 16 | 21.98 | 27.42 | 33.3 | 37.22 | 39.71 | 41.8 | 43.65 | 45.33 | 46.86 | 48.26 | 49.52 | 50.69 | 51.77 | 52.76 | 53.68 | 54.55 | 55.36 | 56.14 | 56.9 | 57.71 | 58.4 | 59.11 | 59.81 | 60.47 | 61.14 | 61.77 | 62.42 | 63.06 | 63.64 | 64.24 | 64.8 | 65.38 | 65.94 | 66.47 | 67.02 | 67.53 | 68.08 | 68.61 | 69.11
 | 16.895 | 23.53 | 30.16 | 36.92 | 41.42 | 44.28 | 46.72 | 48.82 | 50.74 | 52.46 | 54.06 | 55.49 | 56.81 | 58.07 | 59.16 | 60.2 | 61.13 | 62.06 | 62.92 | 63.75 | 64.64 | 65.35 | 66.12 | 66.86 | 67.54 | 68.26 | 68.89 | 69.59 | 70.25 | 70.87 | 71.52 | 72.12 | 72.74 | 73.35 | 73.91 | 74.51 | 75.05 | 75.63 | 76.22 | 76.75
 | 15.534 | 21 | 26.63 | 32.68 | 36.26 | 38.71 | 40.77 | 42.58 | 44.24 | 45.72 | 47.07 | 48.34 | 49.49 | 50.54 | 51.51 | 52.44 | 53.3 | 54.09 | 54.89 | 55.68 | 56.47 | 57.19 | 57.88 | 58.56 | 59.21 | 59.88 | 60.51 | 61.15 | 61.77 | 62.35 | 62.93 | 63.48 | 64.07 | 64.63 | 65.17 | 65.73 | 66.22 | 66.75 | 67.27 | 67.75
 | 16.125 | 21.7 | 27.2 | 32.99 | 36.91 | 39.42 | 41.8 | 43.42 | 45.11 | 46.62 | 48.03 | 49.3 | 50.47 | 51.57 | 52.57 | 53.48 | 54.35 | 55.19 | 55.96 | 56.73 | 57.55 | 58.24 | 58.95 | 59.66 | 60.3 | 60.97 | 61.59 | 62.23 | 62.87 | 63.47 | 64.06 | 64.6 | 65.17 | 65.74 | 66.26 | 66.79 | 67.32 | 67.9 | 68.38 | 68.89
 | 18.68 | 26.44 | 34.12 | 42.99 | 46.97 | 50.1 | 52.67 | 55.01 | 57.04 | 58.9 | 60.6 | 62.12 | 63.52 | 65.13 | 65.98 | 67.06 | 68.06 | 69 | 69.89 | 71.02 | 71.7 | 72.46 | 73.25 | 74.03 | 74.74 | 75.48 | 76.17 | 76.87 | 77.55 | 78.21 | 78.88 | 79.51 | 80.15 | 80.77 | 81.39 | 82.01 | 82.59 | 83.19 | 83.77 | 84.32
36 p on n single-sided silicon microstrip detectors (HPK):
  51.2 x 51.2 mm2 active area
  <100> crystal type
  320 mm thickness
  200 mm pitch
  256 channels
  I(Vfd)~2-3nA/cm2
No bad strips for all 36 detectors (9216 strips)
Chart: 
| 75 | 80 | 85 | 90 | 95 | 100
Vfd | 0 | 3 | 28 | 1 | 4 | 0
Full Depletion Voltage (V)

## Slide 16: Microstrip sensors layout
207320mm
55750mm
207320mm
55750mm
Mounted on the two sides of a PCB which houses the front-end and readout electronics.
Sensors are overlapped to assure hermeticity
X-view
Y-view
Errors in the overlap regions
~ 2.75% of the total tracker active area

## Slide 17: Tracker DAQ architecture
Central acquisition
VIRTEX 6
CALORIMETER DAQ
Trigger
enable
Trigger
7
gen
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
6 readout links
200(x8) Mbs each
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
4 readout links
200(x8) Mbs
each
The front end of each detector will read-out by a Xilinx Spartan 6 FPGA (SS)
When a trigger occurs the corresponding SS containing at least one hit will send the data to a central FPGA (SM)
The SM will then send tha data to the central acquisition board
Trigger
gen
Trigger_en
7
Ethernet link

## Slide 18: Tracker front-end board

## Slide 19: New calorimeter
2x7 YAG:Ce Crystals Array
Size: 3x3x10cm3
CHASSIS NI PXIe-1071
RT Controller
NI PXI-8102
FlexRIO
NI PXIe-7962R
Ad.Mod.
NI-5751
14 Analog Channels
Tracker
7Dig I/O GEN
Dig. Trigger
Disable Trigger
Silicon Photodiodes 1.8x1.8cm2
Fast Charge Amplifier + Shaper
x14
Data Acquisition System
Parallel read-out
Sampling: 5MS/s
24 Samples x event

## Slide 20: Conclusions
A pCT device will be useful to increase the effectiveness of hadron therapy (patient positioning and treatment plan precision)
The ‘Prima’ collaboration has built a prototype (5x5cm2) capable of acquiring tomographic images
Tomographic images have been reconstructed using FBP
  0.9mm (FWHM) spatial resolution
  2.4% (r.m.s.) noise
An upgraded detector with an extended field of view necessary to perform pre-clinical studies has been designed and it is now under construction