## Slide 1: ProSPECTus: A Compton camera for medical Imaging
By Amina Patel

## Slide 2: Contents
Background
Compton imaging
Detectors
Importance of position resolution
Simulation results
Summary

## Slide 3: Motivation : Medical Imaging
Nuclear medicine techniques Functional images
ProSPECTus designed to overcome limitations in conventional Single Photon Emission Computed Tomography (SPECT) system.
Prevents co registry with Magnetic Resonance Imaging (MRI) scanners
High dose or Long scans
Multiple Scans
Main Gamma Camera Components

## Slide 4: Compton Camera: Principles
Gamma ray interacts with both detectors
Energy deposited used to find out scattering angle via Compton scatter formula
Conic represents all possible source locations
Position interactions create cone axis
Many events lead to overlap of conics revealing source location -
2θ
Source
Location

## Slide 5: ProSPECTus System : Scatter Detector
8mm
Front
Back
71mm
Material : Lithium drifted Silicon
Active area: 3500 mm2
13 contact strips on each face of width 5mm
Voxel size = position resolution ( 5 x 5 x 8 ) mm
Detector Performance :
Average intrinsic timing resolution at 511 keV : 16ns
Average Energy resolution @ 60keV : 1.5 keV

## Slide 6: ProSPECTus System: Absorber
60mm
60mm
20mm
Back Face
Front Face
Material : High purity Germanium
Active area: ( 60 x 60 x 20 ) mm3
12 contact strips on each face of width 5mm
Voxel size = position resolution ( 5 x 5 x 20 ) mm
Detector Performance :
Average intrinsic timing resolution at 511 keV : 22 ns
Average Energy resolution @ 60keV : 2 keV

## Slide 7: The ProSPECTus System
Prototype system
ProSPECTus

## Slide 8: Modeling the ProSPECTus System
Investigate and improve Image quality capabilities of ProSPECTus
Image resolution quantitative assessment of image quality
ProSPECTus system modeled in GAMOS (GEANT4-based Architecture for Medicine-Oriented Simulations)1
Creates output text file that contains required information in suitable format for image reconstruction code
1A Compton camera application for the GAMOS GEANT4-based framework , L..J. Harkness, NIMA

## Slide 9: Image Quality :Image Resolution
LEHR
LEHS

## Slide 10: Image Quality : Pulse Shape Analysis (PSA)
Principle : Manipulates signals generated from contact strips to determine position of interaction
Two types :
Parametric
Generate a database of pulse shapes for each interaction position

## Slide 11: Pulse Shape Analysis (PSA) : Parametric
Time (ns)
Normalised Magnitude
R. Cooper, Thesis
a) Depth: Charge Collection
Parametric PSA has two components :
Depth- Charge Collection
Precision in X and Y direction – Image charges
Z= depth
y
x

## Slide 12: Pulse Shape Analysis :Parametric
a) Depth: Charge Collection
10%
30 %
90 %

## Slide 13: PSA : Parametric
h
e
h
e
h
e
h
e
h
e
b) Precision in X and Y – Image Charges

## Slide 14: PSA :Pulse Shape Database
Absorber Detector
Z
X
Y
Coincident events:
137Cs – 662 keV
Absorber Detector : 374 keV
Scintillator detector : 288 keV

## Slide 15: Position Resolution Investigations
Exact Position on Scatter with normal segmentation on Absorber ( no physics effect) : 26 mm
Exact Position on absorber detector with normal segmentation on Scatter detector ( no physics effects): 19 mm
Position sensitivity most important in the absorber detector
Simulations used to investigate lateral and depth position precision independently

## Slide 16: Position Resolution Investigations
Depth Position Resolution Investigation: Absorber
(mm)
Including physics effect
Position resolution
Position Resolution in Depth (Z) mm

## Slide 17: Position Resolution Investigations
Depth Position Resolution Investigation: Absorber
(mm)
Including physics effect
Position resolution
Position Resolution in Depth (Z) mm

## Slide 18: Position Resolution Investigations
(mm)
Pixel dimensions in X and Y
Lateral Position Resolution Investigation: Scatter

## Slide 19: Position Resolution Investigation Conclusions
Improvement in position resolution will improve image resolution
Position resolution in depth has more impact than that in X and Y
Position resolution most important in absorber detector

## Slide 20: Image Quality : Phantom
Drilled holes for rods
50 mm
100 mm
120 mm
Perspex Cylinder
60 mm
10 cm
10 cm
Rod diameters: 6.0, 8.0, 10.0, 11.0, 12.4 mm
Rod diameters: 9.5, 11.0, 10.0, 16.4 mm
139Ce: 165 keV
Sam : s.colosimo@liverpool.ac.uk

## Slide 21: Image Quality : Phantom
Characteristic Position resolution
1mm3 Position resolution in absorber

## Slide 22: Image Quality : Phantom
Characteristic Position resolution
1mm3 Position resolution in absorber

## Slide 23: Summary/Future Work
Conclusion
Position resolution has significant influence on image quality
Pulse shape analysis techniques will be applied to the ProSPECTus absorber detector to improve image resolution
Phantom will be used to compare the ProSPECTus system to current SPECT systems
Current Work
Acquiring experimental images with point sources and phantom
Apply parametric PSA to data
Future Work
Generate a pulse shape database for the absorber detector
Develop PSA algorithms
Images with Technetium (141 keV) to be compared to SPECT images

## Slide 24: Collaborators
L.J. Harkness-Brennan1 , A.J. Boston1 , H.C. Boston1, S.Colosimo1,J.R. Cresswell1, J.Dormand1, T.Hughes1 D.S. Judson1, P.J. Nolan1, J.A. Sampson1, N. Clague2, J .Groves2, J .Headspith2, A .Hill2, I.H. Lazarus2, J. Simpson2, W.E. Bimson3, G.J. Kemp3, and D.Gould4
1 Department of Physics, The University of Liverpool, UK,
2 STFC Daresbury Laboratory, UK
3 MARIARC, The University of Liverpool, UK
4 Royal Liverpool University Hospital, Liverpool, UK

## Slide 25: Compton Efficiency: Previous Work
Factor of 3 difference between simulated and experimental data.
GAMOS2 Monte carlo simulations previously validated against Compton data and matched.
1. A. Sweeney. Compton Imaging for Homeland Security, Thesis 2014.
2. L.J. Harkness , A Compton camera application for the GAMOS GEANT4-based framework , NIMA 2012
1
Energy (keV)
Efficiency (%)
Absolute Compton Efficiency
0
200
400
600
800
1000
1200
1400
1600
0.00E+00
1.00E-03
2.00E-03
3.00E-03
4.00E-03
5.00E-03
6.00E-03

## Slide 26: Compton Efficiency : Thresholds
Threshold ~15keV
( L.J.Harkness, Thesis 2010)

## Slide 27: Compton Efficiency : Thresholds
( L.J.Harkness, Thesis 2010)

## Slide 28: Compton Efficiency : Thresholds
Threshold
~ 6 keV
( L.J.Harkness, Thesis 2010)

## Slide 29: Compton Efficiency: Coincidence Window
Example TAC spectrum from
FWHM= Time Resolution ~ 15ns
22NA
Done for energy range 80 keV – 511 KeeV
Notes: Redo TAC PEAK in MATLAB ( axis bigger)

## Slide 30: Coincidence Window Results

## Slide 31: Image Quality : Pulse Shape Analysis (PSA)
Principle : Manipulates signals generated from contact strips to determine position of interaction
Two types : 1) parametric 2) generate a database of pulse shaped for each interaction position
Notes: Add in Reference

## Slide 32: Image Quality : Pulse Shape Analysis (PSA)
Principle : Manipulates signals generated from contact strips to determine position of interaction
Two types : 1) parametric 2) generate a database of pulse shaped for each interaction position
Notes: Add in Reference

## Slide 33: Image Quality : Pulse Shape Analysis
PSA investigations
GAMOS Simulated data (141 keV)
Absorber position resolution dominant
Position resolution 1mm3
Plan
Develop pulse shape analysis algorithms
Develop pulse shape database for Germanium detector

## Slide 34: Motivation : Single Photon Emission Computed Tomography (SPECT)
Medical imaging
SPECT

## Slide 36: Image Quality : Pulse Shape Analysis (PSA)
Principle : Manipulates signals generated from contact strips to determine position of interaction
Two types : 1) parametric 2) generate a database of pulse shaped for each interaction position
3
4
5
6
7
Magnitude (keV)
Time (ns)
Time (ns)
Time (ns)
Time (ns)
Time (ns)
Depth: Charge
Collection
Lateral: Image Charges
R. Cooper, Thesis
Notes: Add in Reference

## Slide 37
Chart: Investigating Z thickness on Scatter detector
| 1.0 | 3.0 | 4.5 | 9.0
No DB and perfect energy resolution | 5.95 | 6.71 | 7.44 | 11.23
DB and Perfect Energy Resolution | 10.09 | 10.66 | 11.98 | 17.58
No DB and 1.5 keV energy resolution | 6.01 | 6.67 | 7.39 | 11.02
DB and 1.5 keV energy Resolution | 10.7 | 10.74 | 11.99 | 18.5
no Psa on scatter including DB and energy resolution | 22.91

## Slide 38
Chart: Investigation of Absorber Thickness
| 1.0 | 2.0 | 4.0 | 5.0 | 10.0 | 20.0
No Doppler broadening and perfect energy resolution | 5.87 | 6.21 | 7.01 | 8.01 | 11.81 | 25.9
Doppler broadening and perfect energy resolution | 10.06 | 10.31 | 12.17 | 12.95 | 17.67 | 30.41
Dopple broadening and 1.5 keV energy resolution | 5.83 | 6.23 | 6.98 | 8.1 | 11.41 | 26.18
Doppler broadening and 1.5keV energy resolution | 10.74 | 11.04 | 12.03 | 12.94 | 17.61 | 33.04
No PSA on absorber detector including DB and 1.5 keV energy resolution | 34.8

## Slide 39
Chart: Scatter Detector
| 4.0 | 3.0 | 2.0 | 1.0
No Doppler broadening and perfect energy resolution | 17.97 | 16.24 | 13.48 | 9.7
Dopple broadening and 1.5 keV energy resoution | 21.71 | 18.59 | 15.99 | 15.5

## Slide 40
Chart: Absorber detector
| 1.0 | 2.0 | 3.0 | 4.0 | 5.0
No Doppler broadening and perfect energy resolution | 25.5 | 28.51 | 26.19 | 30.46
Doppler broadening and 1.5 keV energy resolution | 31.81 | 33.27 | 30.79 | 31.45