## Slide 1: Imaging of Ra-223 with a small-pixel CdTe detector: potential for improved image quantification for radionuclide dosimetry
James Scuffham1,2 Silvia Pani2, Paul Seller3, Paul J Sellin2, Matthew C Veale3, Matthew D Wilson3, Robert J Cernik4
1. Department of Nuclear Medicine, Royal Surrey County Hospital, Guildford, UK.
2. Department of Physics, University of Surrey, Guildford, UK
3. STFC Rutherford Appleton Laboratory, Didcot, UK
4. School of Materials, University of Manchester, UK

## Slide 2: Prostate Cancer
Most prevalant cancer in men (>40,000 diagnoses per year; 13% of all cancers; 25% of new cases in men)i
Highest incidence in the 75-79 age groupi
81.4% survive more the 5 yearsi
Commonly spreads to the bones in advanced disease (~90% of patients with castration-resistant, prostate cancer (CRPC) have bone metastases)ii
i – ONS Cancer Statistics Registrations, June 2013
ii - Scher H, et al. N Engl J Med. 2012;367:1187-1197, supplemental appendix.

## Slide 3: Bone metastases
Metastases are common in the ribs, spine, pelvis and hips
Can cause debilitating pain and severely impact quality of life
Treatments are palliative

## Slide 4: Xofigo R
New radiopharmaceutical (223RaCl2) by Bayer
Bone-seeking alpha-emitter
Licensed in the UK in Feb 2014 after promising clinical trial results:
20
0
40
60
80
100
Survival %
Months Since Randomisation
0
9
15
24
30
39
36
33
27
21
18
12
6
3
Increased median overall survival of 3.6 months
Increased time to serious skeletal event (5.8 months)
Increased time to Alkaline Phosphatase progression (3.6 months)

## Slide 5: The Alpha Advantage
Marrow
Bone
Range of α-particle: (short range – 2 to 10 cell diameters). High LET, increased chance of double-strand DNA breaks
a emitter
Range of β-particle: (long range – 10 to 1000 cell diameters)
Bone
β-emitter
Marrow

## Slide 6: The Disadvantages
223Ra’s emissions are not optimal for imaging with conventional gamma cameras
Injected activities are generally around 6MBq  very low external dose rate (good for patients’ family, not good for imaging!).

## Slide 7: Why do we need imaging?
Currently very simplistic administration regimen based on patient weight (50kBq/kg)
BUT…patient weight does not necessarily correlate to tumour burden, radiopharmaceutical uptake, or even the patient’s bone marrow reserve
Therapy could be tailored to the individual patient by adjusting administered activity based on uptake and clearance in the bone
Need to image the biodistribution and clearance of radiopharmaceutical, as this varies greatly between patients
 “Personalised medicine”

## Slide 8: Imaging with gamma cameras
Hindorf et al, NMC 2012
Typical Energy Spectra:
Typical Patient Images (MEGP):
“Low-energy” Collimators
“Medium Energy” Collimators

## Slide 9: The HEXITEC Detector
1mm thick CdTe
80 x 80 pixels, 250mm pitch
5-200keV spectroscopy per pixel
10000 frames per second
-500V detector bias
Peltier cooling, room temperature operation
Corrections for charge sharing
Pixelwise energy and uniformity calibrations
Energy resolution 0.8% at 140keV (vs 9% for conventional gamma camera)

## Slide 10: HEXITEC Ra-223 Spectrum
Rn Ka2 – Cd Ka Escape 57.9keV
Rn Ka1 – Cd Ka Escape 60.6keV
Rn Ka2 81.1keV
Rn Ka1 83.8keV
Rn Kb1 94.9keV
Rn Kb2 97.9keV
Ra-223 g 122.3keV
Ra-223 g 144.3keV
Ra-223 g 154.2keV
Rn Lb1 14.3keV
Rn Lg 16.8keV
Glass vial containing ~6MBq Ra-223
3mm lead pinhole collimator
Spectrum collected over 15 hours

## Slide 11: HEXITEC Imaging with Ra-223
“Body phantom” containing 6 spheres (10, 13, 17, 22, 28, 37mm)
Spheres: 434kBq/ml Cylinder: 2.4kBq/ml Cold background
3mm lead pinhole collimator at 28cm from spheres
Planar acquisition, “tiled” with 4 stops
50mins “live time” per stop
Bias refresh (60s on, 4s off, 6s settle)
Energy calibration applied, form images with 85±5keV energy window
Stitch 4 tiles together

## Slide 12: Imaging Results
4 (5?) spheres visible
Low counts; noisy image
Detector edge effects due to “stitching”.
Efficiency limited by pinhole collimation and distance from phantom

## Slide 13: Comparison with gamma camera
GE Infinia SPECT/CT
“Medium Energy” parallel-hole collimators
50mins, 89.5±13.4%keV
5 (6?) spheres visible?
Better statistics
Central compartment visible

## Slide 14: HEXITEC Ra-223 conclusions
Spectroscopy is excellent  potential for improved quantitative imaging due to improved scatter rejection
Sensitivity with lead pinhole collimators too poor for clinical imaging
Need higher sensitivity collimators
Larger area detector would also be needed for clinical imaging

## Slide 15: Ongoing work – RAL and Surrey
Simulation work to optimise parallel-hole collimator design for increased sensitivity for both Tc-99m and Ra-223
Planning to test thicker CZT detector material shortly
Large-area detectors under development (Seller et al, this morning at PSD)