## Slide 1
Characterisation of GaAs:Cr Detector Systems for High Flux X-Ray Imaging
M. C. Veale*, B. Cline, J. Coughlan, M. Hart, T. Nicholls,
A. Schneider, P. Seller, G. Carini, P. Hart, I. Horswell,
E. Gimenez-Navarro, I. Pape, K. Sawhney, N. Tartoni,
A. D. Lozinskaya, V. A. Novikov, O. P. Tolbanov, A. Tyazhev and A. N. Zarubin

## Slide 3
Synchrotron Light Sources
Pilatus
Medipix
HEXITEC

## Slide 4
Free Electon Lasers
CSPAD
AGIPD
LPD

## Slide 5
Free Electon Lasers
CSPAD
AGIPD
LPD
Where Next?

## Slide 6
LCLS II < 30 keV
MaRIE > 40 keV
Matter-Radiation Interaction in Extremes
FEL Upgrades

## Slide 7
Laser-Driven Radiation Sources
ELI-NP (Romania)
g-rays 0.1 -20 MeV
CLF (UK) & Others
Exotic Beams (g, n, e-, p+…)

## Slide 8
Laser-Driven Radiation Sources
ELI-NP (Romania)
g-rays 0.1 -20 MeV
CLF (UK) & Others
Exotic Beams (g, n, e-, p+…)
What About Detectors??

## Slide 9
The Large Pixel Detector

## Slide 10
10%
Transmission

## Slide 11
75%
Transmission

## Slide 12
GaAs:Cr Testing @ STFC
GaAs:Cr HEXITEC.
E < 40 keV FWHM is similar to CdTe.

## Slide 13
GaAs:Cr Testing @ STFC
GaAs:Cr HEXITEC.
E < 40 keV FWHM is similar to CdTe.
STFC Assembled Medipix detectors.
Fine-pitch indium bonding @ STFC.
Good imaging performance.

## Slide 14
GaAs:Cr Testing @ STFC
GaAs:Cr HEXITEC.
E < 40 keV FWHM is similar to CdTe.
STFC Assembled Medipix detectors.
Fine-pitch indium bonding @ STFC.
Good imaging performance.
High Flux Performance?

## Slide 15
GaAs:Cr LPD Modules
Fabrication of GaAs:Cr LPD sensors by TSU.
Direct attach to LPD ASIC.
X Pixel Pitch = 400 mm
Y Pixel Pitch = 250 mm
500 mm thick.
-1000 V (electron readout)
Readout using an LPD super-module DAQ.

## Slide 16
GaAs:Cr sensors tested at the LCLS FEL (USA)
Experiment on MFX Dec ‘16.
500mm thick Si & GaAs:Cr LPD sensors.
The FEL Setup
Compared to 500mm CSPAD.
First known test of high-Z detectors at a FEL.

## Slide 17
GaAs:Cr LPD

## Slide 18
Detector Uniformity
Beam energy 9.5 keV.
Repetition Rate = 30 Hz
Cu (8 keV) flat field at 90o.
Gain = x100 5pF (Highest)
LPD Frame Rate = 4.5 MHz
The challenge: no two pulses are alike! (+/- 100%)

## Slide 19
Si: Intensity Histograms

## Slide 20
Si LPD Calibration = 0.425 keV ADU-1
Average Energy Per Pixel = 303 keV
Average Flux per Pulse = 1.2 MeV mm-2
Si Flux Calculation

## Slide 21
GaAs:Cr Flux Calculation
GaAs:Cr LPD Calibration = 0.507 keV ADU-1
Average Energy Per Pixel = 140 keV
Average Flux per Pulse = 1.4 MeV mm-2
1.2 MeV mm-2 (Si)

## Slide 22
Detector Material | Flux (keV mm-2) | FWHM (%)
Si | 1221 | 49
GaAs:Cr | 1405 | 64
GaAs:Cr greater pixel to pixel variation.

## Slide 23
Comparison with CSPAD

## Slide 25
n x n | Area (mm) | r2
1 x 1 | 0.25 x 0.40 | 0.15
2 x 2 | 0.50 x 0.80 | 0.58
4 x 4 | 1.00 x 1.60 | 0.80

## Slide 26
n x n | Area (mm) | r2
1 x 1 | 0.25 x 0.40 | 0.15
2 x 2 | 0.50 x 0.80 | 0.58
4 x 4 | 1.00 x 1.60 | 0.80
GaAs:Cr Medipix

## Slide 27
Direct Irradiation
9.5 keV X-Rays
GaAs:Cr
Attenuator
Flat field measurements still at low flux ~ 101 ph.
Direct irradiation of the sensor.
Flux per pulse ~ 1 x 1011.
Energy per pulse ~ 1 PeV (0.2 mJ).
Attenuation of beam to avoid ablation of sensors!

## Slide 28
Y (mm)
Intensity (ADU)
~ 1.5 mm
Low gain mode (5pF x1)
Clear intensity variations.
4.5 MHz Imaging

## Slide 29
Highest intensity signals show “positive” signals.
Modification and collapse of electric field.
Dark image correction fails.

## Slide 30
Cause = ~ 1 ns hole lifetime

## Slide 31
What Flux???
CdZnTe
GaAs:Cr

## Slide 32
Direct Irradiation of Si Sensors

## Slide 33
Low gain mode.
~ 350 keV ADU-1
550 – 750 MeV per pixel.
~ 3 GeV mm-2 per pulse.
~ 250 MeV per GaAs:Cr pixel.
Direct Irradiation of Si Sensors

## Slide 34
Conclusions
Development of extreme flux and extreme energy facilities.
These next generation facilities are pushing detector limits!
Technologies beyond Si needed…

## Slide 35
Development of extreme flux and extreme energy facilities.
These next generation facilities are pushing detector limits!
Technologies beyond Si needed…
GaAs:Cr is far from perfect
Pixel to pixel variations. Polarisation @ GeV mm-2
High-Flux CdZnTe an alternative?
B. Thomas Poster [102]
Conclusions

## Slide 36
FEL “Micro-Structuring” of GaAs:Cr
*1 PeV pulses directly on to detector

## Slide 37
Any Questions?