## Slide 1: First results of the MYTHEN-III strip detector prototype
Roberto Dinapoli :: SLS Detector Group :: Paul Scherrer Institut
PSD11 :: 07.09.2017

## Slide 2: Outline
Why strip detectors
MYTHEN
  Description
  Performance
  Experimental issues
MYTHEN III
  Design specifications
  MYTHEN III 0.1 prototype
  Preliminary measurements
Perspectives
SLS
(Swiss Light Source)

## Slide 3: Outline
Why strip detectors
MYTHEN
  Description
  Performance
  Experimental issues
MYTHEN III
  Design specifications
  MYTHEN III 0.1 prototype
  Preliminary measurements
Perspectives
SLS
(Swiss Light Source)
SMS
(Swiss Milk Source)

## Slide 4: Why strips?
1D Applications don’t require pixels
X-Ray powder diffraction
(XRPD)
X-Ray emission spectrometers

## Slide 5: Why strips? Chip designer’s view
1D arrangement makes ASIC design less challenging
Smaller ASIC required (easier, cheaper, faster)
Simpler control/readout logic
Less channels per area
    Faster frame rate
    Smaller data throughput
    Resources much less critical
Big system much simpler than 2D detectors
Wire- instead of bump- bonding
    Cheaper
    Smaller pitches possible
    Simpler
    More input capacitance (noise)

## Slide 6: MYTHEN
Direct conversion in the sensor
Standard sensor:
  320 µm thick silicon sensor
  50 µm pitch
  8 mm long strips
  4 mdeg at 76 cm distance
Sensor material and characteristics can be changed
Single photon counting architecture
  “Noiseless”
  Large dynamic range
  Fluorescence suppression
  Pump-probe experiments
Microstrip sYstem for Time rEsolved experimeNts
Amplifier and
Shaper
From detector
+1
threshold
Comparator
24 bit counter
Sensor
ROCs
X-rays

## Slide 7: MYTHEN I
At the Material Science beamline (2001 until 2007)
Covering 60 degrees (15k channels)
DMILL technology
XRPD pattern acquisition faster than 1 minute instead of 1 hour as with the high resolution scanning detector
  Counter giving random counts and other minor issues
  By redundant measurements and data correction it could be used at the beamline for fast and time resolved measurements
B Schmitt,et al,
Nucl. Instr. . Meth. .Phys. Res. A 501 (2003), 267-272,
doi: 10.1016/S0168-9002(02)02045-4

## Slide 8: MYTHEN II
A. Bergamaschi al., J. Synchrotron Rad. 17, 653-668 (2010)
http://dx.doi.org/10.1107/S0909049510026051
Read Out chip
  0.25um UMC
  128 channels, 50µm pitch
  24 bit counter (configurable 16-8-4-1 bit)
  < 100us read out time
  200-300 e- noise
  100-700 ns “shaping” (dead) time
Modules
  1280 channels, 6.4 cm
  Needs controller board
Systems
  Multi module boards
  Embedded Linux system and FPGAs
  100 Mbs Ethernet, TCP/IP data transfer
  Continuous frame rate up to 20 Hz - 1 kHz
  Depending on dynamic and angular range
  Burst mode up to 10 kHz, 32 frames

## Slide 9: Why MYTHEN III?
It’s getting old!

## Slide 10: Why MYTHEN III?
It’s getting old!
Acquisition system outdated
Electronics components not available any longer

## Slide 11: Why MYTHEN III?
It’s growing!

## Slide 12: Why MYTHEN III?
It’s growing!
Take advantage of technological progress
Improve data quality
More flexibility for more applications
MYTHEN group
SLS detectors group

## Slide 13: Why MYTHEN III?
It’s growing!
Take advantage of technological advances
Improve data quality
More flexibility for more applications
Get rid of pixels for a while…
MYTHEN group
SLS detectors group

## Slide 14: Fluorescence suppression
Fluorescence is not flat field correctable and must be suppressed
Distance between beam, threshold and fluorescence line should
be 3-5 x noise
Difficult energy/threshold choice for complex samples
Sample containing iron
Reduce the noise -> 100-200 electrons
Threshold
6 keV
9 keV
counts
Channel number

## Slide 15: Flat field acquisition
Flat field is necessary in XRPD to get accuracy better than 1%
Annoying and time consuming flat field acquisition for each pair of X-ray and threshold energy
  Large statistics, must be retaken periodically
Reduce the threshold dispersion -> <10 electrons

## Slide 16: High intensity peaks (high photon flux)
Pileup can be corrected, but within limits!
  Much less for high gain settings
For better data quality reduce the flux and avoid rate corrections
Increase the analog speed-> Dead time 40-200 ns
PILEUP EFFECT

## Slide 17: MYTHEN III: the chip
110 nm-Al UMC technology, with enclosed layout
Dual selectable polarity
>40 kHz frame rate
Multiple comparators with separate thresholds, 6 trimbits/comparator, independently gatable
Multiple 24bit counters with possible partial readout
Dead time free operation
Several operation modes

## Slide 18: The channel
TBs2
6
POLARITY
TBs3
6
TBs1
6
Counting
logic
Counter 3
Counter 2
Counter 1
Vth1
Vth2
Vth3
ChSel1
ChSel2
Chsel3
Analog Pulsing
Status register (3x6TB,3x1Mask,3 dummy)
7
7
7
21
Control and readout logic
EN1,EN2,EN3
PUMPPROBE
DIGPULSING
serialIN,
Control signals
serialOUT

## Slide 19: Modes of operation
Chip modes of operation
The chip can work in different modes of operation:
Normal counting mode
  Standard
  With multi-threshold rate correction
  With energy windowing
Pump-probe mode
Time over threshold (with external reference clock)
Analog pulsing mode
Digital pulsing mode
Trim bits load mode

## Slide 20: Normal mode configuration
TBs2
6
POLARITY
Counter 2
Vth2
EN
Analog Pulsing
Status register (3x6TB,3x1Mask,3 dummy)
7
21
Control and readout logic
serialIN,
Control signals
serialOUT

## Slide 21: Energy windowing configuration
TBs2
6
POLARITY
TBs3
6
TBs1
6
Counter 3
Counter 2
Counter 1
Vth1
Vth2
Vth3
EN
EN
EN
Analog Pulsing
Status register (3x6TB,3x1Mask,3 dummy)
7
7
7
21
Control and readout logic
serialIN,
Control signals
serialOUT
One comparator per counter

## Slide 22: Energy windowing configuration
One comparator per counter (same gate, different thresholds)
ESRF
  Monochromator-free operation at undulator beamlines
  Selection of the
  characteristic line
  Higher harmonic suppression

## Slide 23: Energy windowing configuration
One comparator per counter (same gate, different thresholds)
ESRF
  Monochromator-free operation at undulator beamlines
  Selection of the
  characteristic line
  Higher harmonic suppression
Well, let’s be more realistic…

## Slide 24: Rate correction using multiple thresholds
We are exploring possibilities in simulations, Mythen III enables us to compare simulation results with real data
Same configuration as for energy windowing (one comparator per counter, every comparator with independent threshold)
24
Threshold =0.5 E
Threshold =1.5 E
Threshold =2.5 E
Multiple thresholds set at (0.5+n)*E

## Slide 25: Pump-Probe configuration
TBs2
6
POLARITY
Counter 3
Counter 2
Counter 1
Vth2
EN1
EN2
EN3
Analog Pulsing
Status register (3x6TB,3x1Mask,3 dummy)
7
21
Control and readout logic
serialIN,
Control signals
serialOUT
One comparator, several counters with different enable windows

## Slide 26: Pump-probe operation
  Stroboscopic measurements with multiple counting slots
  Piezo crystal with hysteresis
V
V
One comparator, several counters with different enable windows
  Pumped-unpumped measurements
  with isolated/sliced bunch
time
Isolated Bunch (probe)
Gate1
Laser (pump)
Gate 2
Gate 3
Sample

## Slide 27: MYTHEN 3 DAQ system
GOTTHARD/JUNGFRAU based readout board
Use same DAQ architecture as all other SLS detectors
  Common control software to facilitate integration at beamlines
    TCP/IP controls
  UDP data streaming
    ~10kByte/frame (worst case)
    1 or 10 Gbit/s data throughput / module
Fully parallel controls and data streaming
  Detector size independent frame rate

## Slide 28: MYTHEN III 0.1 prototype
Two groups of 32 channels with different preamplifier architectures
  Characterize performance
  Test logic and RO
  Debugging for submission of the final chip
ASIC
SENSOR
ASIC, LAYOUT VIEW

## Slide 29: MYTHEN III 0.1 performances
Prototype fully functional
Tested in the X-ray box and at the beamline (SLS and ELETTRA)
Readout/control logic tested up to 100MHz clk
Dead time free operation working
CHIPTEST BOARD

## Slide 30: Calibration
gain
x
x
x
x
x
x
x
1. Take threshold scans at different photon energies
Extract the inflection points
Calculate the gain [DACunits/keV]
Repeat procedure for all Vrf-settings
Vrf
Vrs
x

## Slide 31: Noise performance
  New chiptest board needs to be further debugged!
  Simulations suggest improvements still possible
Noise
Inflection point Eγ
Charge sharing Cs
N0
PRELIMINARY

## Slide 32: Threshold dispersion
Find inflection points for all channels
Calculate and apply trim bits to reduce dispersion
Cu
Untrimmed: 121 e-
Trimmed: 6.5e-
Mo
Untrimmed: 142 e-
Trimmed: 7.2 e-

## Slide 33: Rate capability
35ns
40ns
49ns
60ns
87ns
158ns
1. Plot the measured rate vs the reference rate I0
Estimate the ideal theoretical rate with a linear fit
Fit the ratio ε of the measured and theoretical rate Φ to find the dead time τd (paralyzable counter): ε=exp(-τdΦ)

## Slide 34: Noise vs dead time
PRELIMINARY

## Slide 35: Perspectives and conclusions
Development of MYTHEN3 just started
  First prototype under test
  Final chip end 2017
  Modules end 2018
  Detector system at SLS in 2019
With respect to previous generation (MYTHEN II):
Lower noise
  Easier choice of energy/threshold
Improved threshold dispersion
  Smaller flat field corrections
Faster readout (>40kHz)
Higher count rates
Three independent counters
  Dead time free
  Pump-probe with multiple gates
  Energy windowing
…and more ideas

## Slide 36
Back: Marie Ruat, Bernd Schmitt, Sophie Redford, Aldo Mozzanica, Erik Fröjdh.
Middle: Jiaguo Zhang, Carlos Lopez, Marie Andrae, Rebecca Barten, Martin Brueckner, Christian Ruder,
Dominic Greiffenberg, Seraphin Vetter.
Front: Xintian Shi, Dhanya Thattil, Gemma Tinti, Anna Bergamaschi, Marco Ramilli, Roberto Dinapoli, Davide Mezza.
Acknowledgements
Nicola
Casati
Antonio Cervellino
MS
beamline

## Slide 37: Thank you!
OPEN POSITIONS: https://www.psi.ch/pa/job-opportunities/1525