## Slide 1: SPECTROMETERS
Teresa Kurtukian Nieto
kurtukia@lp2ib.in2p3.fr
LP2i Bordeaux
CNRS/IN2P3 Université de Bordeaux
1st Lecture: 24/01/2023, 12:00 - 13:00
Definitions; Formalism; Main ion-optical elements
2nd Lecture: 25/01/2023, 10:30 - 11:30
Higher Orders ; Exemples

## Slide 2
Literature
• Optics of Charged Particles, Hermann Wollnik, Academic Press, Orlando, 1987.
• The Optics of Charged Particle Beams, David.C Carey, Harwood Academic Publishers, New York 1987.
• A First- and Second-Order Matrix Theory for the Design of Beam Transport Systems and Charged Particle Spectrometers. Karl L. Brown. SLAC Report-75. June 1982
https://cds.cern.ch/record/283218/files/SLAC-75.pdf
Computing Codes:
• COSY INFINITY 10.0 Beam Physics Manual https://www.bmtdynamics.org/cosy/manual/COSYBeamMan100.pdf
• GICOSY Manual https://web-docs.gsi.de/~weick/gicosy/
• TRANSPORT http://aea.web.psi.ch/Urs_Rohrer/MyWeb/trans.htm
Thanks B. Jacquot and Franck Varenne/ GANIL for some material used in this course

## Slide 3
What is a Spectrometer?
In the broadest sense a spectrometer is any instrument that is used to measure the variation of a physical characteristic over a given range.
A dipole magnet is the simplest electromagnetic spectrometer to scan on mass-to-charge ratio (m/q)
Focal plane

## Slide 4
Dispersion

## Slide 5
GSI FRS Germany

## Slide 6
BIGRIPS RIKEN Japan

## Slide 7
LISE (wien filter)
GANIL, France

## Slide 8
VAMOS
GANIL, France

## Slide 10
1. Lorentz force: A charged particle moving in an electromagnetic field experiences a force.
Why it works?
Thanks to the Lorentz force F and Newton’s second law
Electic Magnetic
Force Force
This force causes a centripetal acceleration and consequently a circular motion of the particle in the medium based on the equations described below.
Radius r → ρ
2. Newton’s second law

## Slide 11
Electrostatic selection :
Difficult to bend energetic particles with
raisonnable E field due to sparking
Most used for low energy particles keV
F= q E
FElectic = Fcentripetal
E

## Slide 12
Aston Nobel price (1919) : E+ B selection with a « mass spectrograph »
  identification Stable isotopes : 20-22Ne; 35-37Cl & mass measurement

## Slide 13
Magnetic Separation:
FMagnetic = Fcentripetal
Fmagnetic = q v B
, → Magnetic Rigidity
Beam rigidity quantifies how difficult it is to bend the beam and is given by the total momentum divided by the total charge
Wien Filter: Felectic = Fmagnetic
v = E/B with E ⊥ Β
m/

## Slide 14
The simplest m/q magnetic spectrometer : 1 dipole magnet
40’s: Manhattan project U-235/U-238 enrichment (B selection)
Dipole →mass dispersion

## Slide 15
Bending through Dipole Field

## Slide 16
The dipole elements also have focusing/defocusing properties.
With edges perpendicular to the optical axis (edge angle 0◦) focuses the beam in the
bending plane (x). There is no focusing action in the y direction.
If the magnet edge angles deviate from 90◦, the focusing power in the x direction can be
adjusted. If the edge angle is made positive (as shown), there is weaker focusing
in the x direction. If the angle is negative, there is stronger focusing in the x direction.

## Slide 17
Changing the edge angle also has an important effect in the y direction:
if the angles are positive, the fringing field of the magnet will focus the beam in the y direction
Overall, this means that the focusing in the x direction can be traded for y focusing.
The focal length from the edge focusing is given by.

## Slide 19
Focal plane

## Slide 20
Focal plane

## Slide 21
Focusing in both planes : doublets, triplets

## Slide 22
Resolving power
The term resolving power is the ability of a spectrometer to resolve adjacent peaks in a mass spectrum and is often used interchangeably with resolution. The separation of peaks for singly charged ions can be expressed as a mass difference δm

## Slide 23
Resolution = ΔxFWHM / dx/dm, Resolving power = 1/Resolution
Mass dispersion usually expressed in meters (m) (SI):
cm/% (centimeters per 100%) ;
mm/‰
Notation :
Dm
dx/dm, physical meaning
Matricial notation (see later)
(x|δ) Wollnik
R16, T16, M16
Resolving power R =

## Slide 24
Exercise 1:
Imagine a spectrometer with a dispersion of 30 cm/%
and beam width of 1 mm FWHM on the focal plan detector.
What is the resolving power R ?
30
30000
1500

## Slide 25
Beam optics (basics)
Already seen:
Dispersion and focalisation with dipoles
Focalisation with quadrupoles
Resolution
Next concepts:
Particles coordinates
Beam emittance
Optical Matrices following Taylor expansion
Angular Acceptance
Bρ Acceptance

## Slide 26
The Coordinates
Notations in the Literature is not consistent!
Wollnik GICOSY | Brown | TRANSPORT | COSY | Meaning
x | x | x | r1 = x | the horizontal displacement of the arbitrary ray with respect to the assumed central trajectory.
a | x’ | θ | r2 = a = px/p0 | the angle this ray makes in the horizontal plane with respect to the assumed central trajectory.
y | y | y | r3 = y | the vertical displacement of the ray with respect to the assumed central trajectory
b | y’ | φ | r4 = b = py/p0 | the vertical angle of the ray with respect to the assumed central trajectory
l | l |  | r5 = l = −(t − t0)v0γ/(1 + γ) | the path length difference between the arbitrary ray and the central trajectory.
δ | δ | dp/p= |  | fractioned momentum deviation of the ray from the assumed central trajectory
δU |  |  | r6 = δK = (K − K0)/K0 | energy difference ray with respect to the reference energy
δm |  |  | r7 = δm = (m − m0)/m0 | mass difference ray with respect to the reference energy
δe |  |  | r8 = δz = (z − z0)/z0 | charge difference ray with respect to the reference energy
Wollnik GICOSY | Brown | TRANSPORT | COSY | Meaning
x | x | x | r1 = x | the horizontal displacement of the arbitrary ray with respect to the assumed central trajectory.
a | x’ | θ | r2 = a = px/p0 | the angle this ray makes in the horizontal plane with respect to the assumed central trajectory.
y | y | y | r3 = y | the vertical displacement of the ray with respect to the assumed central trajectory
b | y’ | φ | r4 = b = py/p0 | the vertical angle of the ray with respect to the assumed central trajectory
l | l |  | r5 = l = −(t − t0)v0γ/(1 + γ) | the path length difference between the arbitrary ray and the central trajectory.
δ | δ |  |  | fractioned momentum deviation of the ray from the assumed central trajectory
δU |  |  | r6 = δK = (K − K0)/K0 | energy difference ray with respect to the reference energy
δm |  |  | r7 = δm = (m − m0)/m0 | mass difference ray with respect to the reference energy
δe |  |  | r8 = δz = (z − z0)/z0 | charge difference ray with respect to the reference energy

## Slide 27
Beam emittance
The emittance is defined as the six-dimensional volume limited by a contour
of constant particle density in the (x, px, y, py, z, pz) phase space. This volume obeys the Liouville theorem and is constant in conservative fields
optical system
The area of the particle distribution is conserved but the area of the elliptical envelope increases.

## Slide 28
Beam emittance
ε is the two-dimensional transverse emittance, and α, β and γ are known
as the Twiss parameters

## Slide 29
The beam size : important for the design

## Slide 30
percentage = (1 - exp (-numberOfSigmas2/2)) · 100%,
(x/σx)2 + (y/σy)2 = numberOfSigmas2.
The percentage of bivariate normally distributed data covered by an ellipse whose axes have a length of numberOfSigmas · σ can be obtained by integration of the probability distribution function over an elliptical area.
where the numberOfSigmas is the radius of the "ellipse":
the numberOfSigmas =1 ellipse covers 39.3% of the data,
the numberOfSigmas =2 ellipse 86.5%,
and the numberOfSigmas =3 ellipse 98.9%.
From the formula above we can show that if we want to cover p percent of the data,
we have to chose numberOfSigmas as
This results in the following equation,
numberOfSigmas = √(-2 ln(1-p/100)).
For covering 95% of the data we calculate numberOfSigmas = 2.45.
Resolving power (95%) =

## Slide 32
Modelling of ion optical transport lines
1. Trajectories : exact equations
integrate the particle equation of motion using mesh based maps for E and B fields [field map 3D]
Examples of codes : ZGOUBY
But generally we can do simpler : Matrix approach

## Slide 33
Modelling of ion optical transport lines
Taylor expansion in x, a, y, b and δ
+
+
=
First order

## Slide 34
6x6 Matrix
representing first optic element
(usually a Drift)
Ray at final
Location n
Ray at initial
Location 0
(e.g. a target
or slits)
Complete system is represented in first order by one Matrix Rsystem = Tn Tn-1 … T0
Transfer matrix formalism

## Slide 35
Transfer matrix formalism
Most crucial parameters :

## Slide 37
Exercise 2:
For covering 95% of the beam ellipse data which value of sigma in ΔX we should use for calculating the resolving power?
1 𝛔
2.35 𝛔 (FWHM)
2.45 𝛔

## Slide 38
Supplemental slides

## Slide 39
Following Taylor expansion the trajectory component Xi after propagation through an ion optical element can be calculated from
where Yi are the components of the trajectory before the ion optical element, and
(Xi | Yj), (Xi | YjYk), (Xi | YjYkYl), . . . are the first-order, second-order, third-order, . . . transfer coefficients
This can be described as matrix–vector multiplication with :
6 × 6 matrix in first order
6 × 62 matrix in second order,
6 × 63 matrix in third order, etc.
Transfer matrix formalism

## Slide 40
Transfer matrix formalism