## Slide 1: First results of the high-rate μRWELL prototypes made by new techniques：
You Lv1, Yi Zhou1, R. De Oliveira3, Bertrand Mehl3, Antonio Teixeira3, Olivier Pizzirusso3, Simon Williams3, Lunlin Shang4, Jianbei Liu1, Ming Shao1, Zhiyong Zhang1
State Key Laboratory of Particle Detection and Electronics, USTC
CERN
Lanzhou Institute of Chemical Physics Chinese Academy of Sciences
RD51 Collaboration Meeting
6/10/2020
PEDF: Patterning-Etching-Drilling-Filling
PEDP: Patterning-Etching-Drilling-Plating
DEF: Drilling-Etching-Filling
DEP: Drilling-Etching-Plating

## Slide 2: Contents
Motivation
Novel solutions to high-rate μRWELL
Characterization of the detectors with X-rays
Summary

## Slide 3
ATLAS high-η muon tagger: identify muons that penetrate the endcap calorimeter by reconstructing track segments in the tagger.
Radial dimensions: 25cm-90cm
: Between the endcap calorimeter cryostat and the JD shielding of NSW.
ATLAS High-η muon tagger

## Slide 4
High rate capability: up to 10 MHz/cm2 Spatial resolution: a few Hundreds of micrometers
High granularity: a few mm2 (occupancy of 3% with a drift time of 100ns)
Total thickness: 5 cm
Detector requirements
R (ƞ) | 250 (4) | 340 (3.7) | 450 (3.4) | 610 (3.1) | 830 (2.8)
Rate (MHz) | 8.8 | 2.3 | 0.84 | 0.23 | 0.076
Hit rate vs. Radius
Fake muon vs angle cut
cut1: Polar angle cut2: Azimuth angle
Performance requirements:
Standalone simulation: detector requirements (rate capability & position resolution)

## Slide 5
μRWELL detector
Copper
Kapton
DLC
Epoxy glue
Pre-preg
Insulator layer
Readout
μRWELL PCB
Rate capability of μRWELL
DLC resistive electrode:
1. A drop of the amplifying voltage on the DLC resistive layer caused by radiation, reducing the capability to stand high particle fluxes.
2. Higher the radiation rate, larger current drawn through the resistive layer, larger the voltage and gain drop.
For larger area detector, the gain drop caused by the radiation is more seriously.
μRWELL: Novel MPGD with resistive electrode and a single stage of well-type gas amplification.
μRWELL is one of the candidate detectors of high-ƞ muon tagger.
High-rate: key performance to the application on high-ƞ muon tagger.
In order to stand high particle fluxes, the fast-grounding technique is presented.

## Slide 6
Fast-grounding μRWELL (SG2++ & SBU )
1. SG2++ type (Cu grid): The copper clad on the DLC is etched to conductive grounding lines by photo-lithography.
2. SBU type (Sequential Build Up): Current evacuation achieved by two stacked DLC layers.
Matrix of conductive vias manufactured with SBU technology are used to connect DLC layer to grounding.
SG2++ type
SBU type
For SG2++, after gluing the APICAL substrate onto the readout PCB, it is impossible to see the fast grounding lines, so it is impossible to align the mask for APICAL etching.
For SBU, still have same alignment problems when making the conductive vias.
Ideal
Reality

## Slide 7
Motivation
Novel solutions to high rate μRWELL
Characterization of the detectors with X-rays
Summary

## Slide 8
Novel idea of high-rate μRWELL (PEDF)
PEDF: Patterning , Etching , Drilling & Filling
No copper-coated DLC needed, better resistivity control;
No alignment problems even goes to large area;
Larger contact area between DLC and silver glue, improving the connection.
Advantages:
Apical
DLC
Cu
Prepreg
PCB
Step1: Copper & APICAL etching, to make a big hole, with DLC on bottom.
Step2: Drill a small hole, the copper of the readout pad expose to air.
Step4: Make μRWELL structure and remove the copper around silver glue.
Step3: Use silver glue to connect the DLC to readout pad.

## Slide 9
PEDP: Patterning , Etching , Drilling & Plating
DEF: Drilling , Etching & Filling
DEP: Drilling , Etching & Plating
PEDP,DEF,DEP μRWELL
Different options of high-rate μRWELL.
Four different types (PEDF , DEF , PDEF , DEP) of μRWELL PCB have been produced at CERN and transferred to USTC .

## Slide 10
High-rate μRWELL prototypes
Pad parameters
Size: 0.85mm @X & 2.85mm @Y
Pitch: 1mm @X @ 3mm @Y
Channel: 48(X) by 16(Y) = 768
Fast-grounding hole: 1mm diameter, total 80 holes with a geometry acceptance of 97.5 %.
These μRWELL PCBs (PEDF,PEDP,DEF,DEP) are assembled to μRWELL prototypes.
The performance of these detectors are tested with X-rays.
Small-pad readout
μRWELL PCB
Fast-grounding holes
Top-Copper
DLC
DLC
Sensitive area
5cm X 5cm
Small-pad readout applied to high-rate μRWELL detectors in order to meet the requirements of high-η muon tagger.
The parameters of readout pads almost same as small-pad readout MM.
C. Di Donato., et al. Small-pads resistive Micromegas prototype. NIM A. Volume 958, 2020.
5mm
6mm

## Slide 11
Motivation
Novel solutions to high rate μRWELL
Characterization of the detectors with X-rays
Summary

## Slide 12
Test setup
Work gas: Ar/iC4H10 = 95/5
Hirose adapter: 128 readout pad → 1 channel → grounded by 50Ω terminator or Keithley 6482 Picoammeter.
Hirose adapter
To 6482
DLC resistive electrode: grounded via a 50 Ω terminator or 6482 Picoammeter.
Induced signal: picked up from the top-copper or readout-pad.
Current signal: picked up from readout-pad monitored by Keithley 6482.

## Slide 13
PEDF:
8keV X-rays
FW=16.4%
Gain vs drift field
PEDF
Normalized gain, fixed the avalanche electric field.
PEDF&DEF: Maximum gain (~ 3kV/cm).
PEDP&DEP: Maximum gain (~ 1kV/cm).
PEDP
DEP
DEF
The thickness of copper cladded on the APICAL is 5 μm (PEDF&DEF).
The plating process make the thickness of copper up to 15 μm (PEDP&DEP). This is the reason to the different behavior between PEDF and PEDP.

## Slide 14
A simulation was carried out to understand the influence on the gain caused by the thickness of copper electrode.
Simulation of the gain vs. drift field
Detector geometry: ANSYS Electron avalanche and the efficient gain: Garfield++
Cu: 5 μm
Cu: 15 μm
The work gas set to Ar/CO2 (70/30) in this simulation, the Ar/iC4H10 (95/5) gas mixture will be simulated in the next step.
The simulation result is qualitative consistent with the test result.

## Slide 15
Definition of gas gain
1. Current gain: absolute gain, current signal exported from readout pad and recorded by Keithley 6482 Picoammeter.
2. Induced signal gain: induced signal exported from readout pad or copper electrode, it was amplified by a pre-amplifier followed by a main amplifier, the signal finally recorded by MCA.
I: current R: rate E: average ionization energy
C: capacitance of pre-AMP U: effective voltage
A pulse generate a signal , and it was input to the test port of Pre-Amp to calibrate the test system (pre-amplifier, main amplifier, MCA).
Two different definition of gas gain presented: Current gain and Induced signal gain
The induced signal gain is the effective gain, it is influenced by the weighting field.
For DAQ system, it always receive induced signal, the induced signal gain is more accurate than current gain.

## Slide 16
PEDP
PEDF
The current gain is about 3 times bigger than induced signal gain.
It is due to that the effective weighting field smaller than 1.
Gain vs avalanche electric field
Current gain:
~30000 @420V
Induced signal gain:
~10000 @420V
DEF
DEP

## Slide 17
Weighting field in μRWELL
1. Ramo Theorem: induced signal :
2. Weighting field En(x): setting the voltage of interested electrode to 1, and other electrode to 0.
μRWELL detector:
The second electron only drift in the well-hole.
70 μm insulating material (50um Prepreg+12um Kapton+10um epoxy glue) between readout pad and DLC electrode, reduce the effective weighting field in well-hole.
Weighting potential: about 0.42 (0.48 to 0.9) in the well-hole.
Induced signal gain is smaller than current gain.
Induced charge: proportional to weighting potential.
Well-hole
Insulating material

## Slide 18
Rate capability (PEDP)
PEDP
Current Gain: 14000
PEDP
Current Gain: 6000
PEDP
Current Gain: 14000
PEDP
Current Gain: 6000
Same current with different gain, same gain drop. The ohm effect result in the gain drop.
Same current with different collimator diameters, different gain drop. The radiation area changes, the effective resistivity changes, the gain drop changes.
Rate: (1MHz/cm2)，
Gain (6000): 1 @1mm, 1 @ 3mm, 0.9 @6mm, 0.85 @ 8mm.
Gain (14000): 0.97 @1mm, 0.95 @ 3mm, 0.85 @6mm, 0.68 @8mm.
Different collimator diameters: 8 mm, 6 mm, 3 mm, 1 mm are used.
For a m.i.p. the primary ionization is 10 times smaller than 8 keV X-rays with a drift gas gap of 3mm.
The log of the relative gain drops linearly with current
Rate capability: assessed by detector gain as a function of counting rate per unit area.
Gain vs. current: Study the ohm effect of DLC resistive electrode in μRWELL.

## Slide 19
Rate capability (PEDF & DEF)
Rate: (1MHz/cm2)，
PEDF gain: 1 @1mm, 0.95 @ 3mm, 0.9 @6mm, 0.8 @ 8mm.
DEF gain: 0.95 @1mm, 0.85 @ 3mm, 0.8 @6mm, 0.7 @8mm.
PEDF
Current gain: 6500
PEDF
Current gain: 6500
DEF
Current Gain: 5000
DEF
Current Gain: 5000
The DEF detector don’t have pattering during the etching process. It may make the hole irregularly, and a poor performance.
For PEDF and DEF: The filling process is not a standard process. The fast-grounding hole would be damaged when removing the excess silver glue.
The rate capability of PEDP is better than PEDF.

## Slide 20
Problems of DEP μRWELL
DEP
(Before flushed)
1. When voltage applied to 425V, discharge occurs (~ 1 μA), then short circuit happened between the DLC and copper.
2. The PCB was flushed by high pressure water gun, then baked in oven with 70 degree.
3. The DEP detector work well in low rate.
4. A ripple signal occurs in high rate. The current exported from readout-pad shows larger fluctuation that make it impossible to have a rate capability result.
When test the highest gain, the PEDF and PEDP also occurs discharge (~ 1μA), but PEDF and PEDP can back to the good work condition.
DEP
(After flushed)

## Slide 21
Resistivity measurement
Measuring the resistivity between DLC and fast-grounding holes.
Measurement
PEDF
PEDP
DEF
DEP
Connected
Out of range (>500 MΩ)
PEDP: 6 holes out of range.
Resistivity: 90 to 110 MΩ
PEDF: 32 holes connected.
Resistivity: 350 to 450 MΩ
DEP: All holes disconnected.
DEF: 1 holes connected.
Resistivity: 380 MΩ
DLC
Grounding hole
For DEP and DEF, it is uncontrollable when etching the grounding holes due to that no pattering during etching process. Bad connection between DLC and grounding holes.
For PEDF, some of the grounding holes would be damaged when removing the excess silver glue.
The PEDP shows the best connection between DLC and grounding holes.

## Slide 22
Next test plan
The rate capability with a larger radiation area (cm2) will be tested in order to study the relationship of the rate capability and radiation area. The effective resistivity with different radiation area will be studied in the next step.
A cosmic tracker system (GEM or MM) will be built in the next step to test the efficiency/position resolution/dead area of small-pad μRWELL. The APV25 readout chip will be used.

## Slide 23
Large-area high-rate μRWELL detector
The layout of the μRWELL-NT-50cmX50cm
The active area is divided into 20 sectors
A 50cmX50cm μRWELL have been designed, the PEDP technique will be used for the large area high-rate μRWELL.

## Slide 24
Summary
1. A standalone simulation was carried out to estimate detector performance of high-ƞ muon tagger.
Rate capability: up to 10 MHz/cm2 Position resolution: A few hundreds of micrometers
2. Four different technique of high-rate μRWELL was presented. The current gain can be larger than 30000, with a induced signal gain larger than 10000. The rate capability of these detectors was tested, and the best performance is achieved with PEDP technique.
3. The performance (rate capability with larger irradiation area, efficiency, position resolution, dead area) will be tested in next step. A larger area μRWELL with PEDP technique have been designed.
Thanks to Giovanni Bencivenni, Eraldo Oliveri and CERN GDD for helpful suggestions and discussions.

## Slide 25
Thank You

## Slide 26
BACKUP

## Slide 27
ATLAS geometry
High-ƞ muon tagger
Materials: 2mm PCB + 3mm gas + 2mm PCB
Total ten detectors with a distance of 10mm each.
1. ATLAS geometry
Pythia8 generate p-p collide events.
Physical process: SoftQCD:nonDiffractive (minimum-bias component)
2. High-ƞ muon tagger

## Slide 29
ORTEC 142AH刻度
Test input: A fast rise time (less than 40 ns) followed by a slow exponential decay (200 to 400 μs).
Energy Output: The charge-sensitive loop is essentially an operational amplifier with a 1-pF capacitive feedback.
While test pulses are being furnished to the Test input, connect either the detector (with bias applied).
信号产生器的刻度信号以及响应
Leading: 10 ns
Trailing: 240 μs
1 KHz
Duty=15%