## Slide 1: GEM production scheme for the sphenix tpc
Klaus Dehmelt
RD51 Collaboration Meeting
October 09, 2020

## Slide 2: Sphenix and the time projection chamber
Oct-09-2020
sPHENIX @ the Relativistic Heavy Ion Collider RHIC in 2023

## Slide 3: sphenix tpc
A next-generation TPC operated in continuous readout mode using Gas-Electron Multiplier (GEM) avalanche w/ low Ion Back Flow (IBF).
Front End Electronics (FEE) uses SAMPA chip (developed by ALICE).
Data Aggregation Module (DAM) uses the FELIX board (ATLAS exp)
High Momentum Resolution:
Large Lever Arm (Maximize Active Area)
High Precision (Good Single Point Resolution)
High Accuracy (Low Distortion)
Design choices reflect these considerations.

## Slide 4: sphenix tpc

## Slide 5: Sphenix time projection chamber
ALICE Quadruple GEM schematics

sPHENIX Quadruple GEM schematics
V | sPHENIX |  |  | 
Setting | 1 | 2 | 3 | 4
G1 Top | 4208 | 4658 | 5124 | 5118
G1 Bot | 3951 | 4401 | 4851 | 4861
G2 Top | 3051 | 3351 | 3651 | 3661
G2 Bot | 2721 | 3021 | 3321 | 3342
G3 Top | 1821 | 1971 | 2121 | 2142
G3 Bot | 1409 | 1559 | 1709 | 1709
G4 Top | 1379 | 1529 | 1679 | 1679
G4 Bot | 900 | 1050 | 1200 | 1200
IBF for these configurations determined with X-ray
    0.44%, 0.39%, 0.33%, 0.31%
All configurations tested in test-beam

## Slide 6: sphenix tpc: readout modules

## Slide 7: sphenix tpc: readout modules
R2 GEM structure
1206 Resistor Placeholder

## Slide 8: sphenix tpc: readout modules

## Slide 9: sphenix tpc gem prototyping
Starting with R2  various iterations
  Version 1  segmentation 16 sectors
  Version 1  HV-cutout

## Slide 10: sphenix tpc gem prototyping
Prototype test beam 2018

## Slide 11: sphenix tpc gem prototyping
Prototype test beam 2018 results
  Promising w.r.t. physics goals
  Need for revisiting design
Version 2a
  Like version 1  segmentation 16 sectors
  New: HV lead from GEM
  Prototype test beam in 2019

## Slide 12: sphenix tpc gem prototyping
Prototype test beam 2018 results
  Promising w.r.t. physics goals
  Need for revisiting design
Version 2a
  Like version 1  segmentation 16 sectors
  New: HV lead from GEM
  Prototype test beam in 2019
Version 2b
  Identical to version 2a but different length of HV lead + some clean-ups

## Slide 13: sphenix tpc Final gem Production
Submitted final design to CERN MPT-Workshop (Rui, Alexis, Bertrand)
  DXF-file format

## Slide 14: sphenix tpc Final gem Production
Submitted final design to CERN MPT-Workshop (Rui, Alexis, Bertrand)
  DXF-file format
  Arrangement of GEM on template
  Type of GEM
My request:

## Slide 15: sphenix tpc Final gem Production
Submitted final design to CERN MPT-Workshop (Rui, Alexis, Bertrand)
  DXF-file format
  Arrangement of GEM on template
  Type of GEM
Their response:

## Slide 16: sphenix tpc Final gem Production
Submitted final design to CERN MPT-Workshop (Rui, Alexis, Bertrand)
  DXF-file format
  Arrangement of GEM on template
  Type of GEM
My request:

## Slide 17: sphenix tpc Final gem Production
Submitted final design to CERN MPT-Workshop (Rui, Alexis, Bertrand)
  DXF-file format
  Arrangement of GEM on template
  Type of GEM
Their response

## Slide 18: sphenix tpc Final gem Production
Submitted final design to CERN MPT-Workshop (Rui, Alexis, Bertrand)
  DXF-file format
  Arrangement of GEM on template
  Type of GEM
Received first article GEMs
Their response

## Slide 19: sphenix tpc Final gem Production
Submitted final design to CERN MPT-Workshop (Rui, Alexis, Bertrand)
  DXF-file format
  Arrangement of GEM on template
  Type of GEM
Received first article GEMs
Verified design and submitted final production order

## Slide 20: GEM production scheme for the sphenix tpc
Twenty years of experience with Rui’s workshop
Outstanding professionalism and quality
Very flexible procedures
Very cooperative dealing
COMPASS – PHENIX/HBD – EIC-prototyping – sPHENIX
  Started off in a cleanroom next to Rui’s workshop
  Now: ~ 4000 mi (~ 6500 km) distance
  still same encounter
Conclusion